## Slide 3: What will change?
Old detector
Read-out noise: 2 ph/pix
  100 ADU/pixel
1.15s read-out, ~0.3 Hz
  23 min data set @ 30 MGy
  rastering impractical
16 bits
0.7% inactive area
315 mm face; 80 μm PSF
  1312 orders across face
No max count rate
3% SHSSS error
  7235 eV optimal
New detector
No read-out noise
  no more strategy
2.03 ms read-out, 25 Hz
  200 s data set @ 30 MGy
  10 min/mm2 raster 1% lifedose
20 bits
8.5% inactive area
430 mm face, 172 μm PSF
  833 orders across face
Max 107 ph/s/pix
< 1% SHSSS error
  S-SAD, 5 keV?

## Slide 4: What will change?
Old detector
Read-out noise: 2 ph/pix
  100 ADU/pixel
1.15s read-out, ~0.3 Hz
  23 min data set @ 30 MGy
  rastering impractical
16 bits
0.7% inactive area
315 mm face; 80 μm PSF
  1312 orders across face
No max count rate
3% SHSSS error
  7235 eV optimal
New detector
No read-out noise
  no more strategy
2.03 ms read-out, 25 Hz
  200 s data set @ 30 MGy
  10 min/mm2 raster 1% lifedose
20 bits
8.5% inactive area
430 mm face, 172 μm PSF
  833 orders across face
Max 107 ph/s/pix
< 1% SHSSS error
  S-SAD, 5 keV?

## Slide 5: Simulated diffraction image MLFSOM

## Slide 6: Simulated diffraction image MLFSOM
simulated
real

## Slide 7: Simulated diffraction image (2008)
simulated
real

## Slide 8
Simulated | statistic | Real
5.4% | Rmerge | 6.2%
18.9 | I/sd | 16.2
1.9 | I/sd (1.4 Ǻ) | 1.6
1.7 4.0 0.020 | SDCORR | 1.0 2.2 0.065
36.8 | PADFPH | 31.46
3.871 | mlphare f” | 3.476
0.178 | FOM | 0.192
0.545 | FOMDM | 0.664
0.6270 | CC(1H87) | 0.6090
0.170 | Rcryst | 0.184
0.205 | Rfree | 0.231

## Slide 9: a perfect detector
“normal”
“perfect”

## Slide 10
normal | statistic | perfect
4.0% | Rmerge | 3.9%
24.3 | I/sd | 28.1
3.7 | I/sd (1.4 Ǻ) | 7.3
1.37 3 0.02 | SDCORR | 0.98 11 0.05
38.01 | PADFPH | 35.16
4.104 | mlphare f” | 3.973
0.224 | FOM | 0.272
0.579 | FOMDM | 0.563
0.7113 | CC(1H87) | 0.7051
15.72 | Rcryst | 16.64
18.30 | Rfree | 19.38

## Slide 11: Optimal exposure time (faint spots)
thr Optimal exposure time for data set (s)
tref exposure time of reference image (s)
bgref background level near weak spots on
reference image (ADU)
bg0 ADC offset of detector (ADU)
bghr optimal background level (via thr)
σ0 rms read-out noise (ADU)
gain ADU/photon
m multiplicity of data set (including partials)
adjust exposure
so this is ~15

## Slide 12: What will change?
Old detector
Read-out noise: 2 ph/pix
  100 ADU/pixel
1.15s read-out, ~0.3 Hz
  23 min data set @ 30 MGy
  rastering impractical
16 bits
0.7% inactive area
315 mm face; 80 μm PSF
  1312 orders across face
No max count rate
3% SHSSS error
  7235 eV optimal
New detector
No read-out noise
  no more strategy
2.03 ms read-out, 25 Hz
  200 s data set @ 30 MGy
  10 min/mm2 raster 1% lifedose
20 bits
8.5% inactive area
430 mm face, 172 μm PSF
  833 orders across face
Max 107 ph/s/pix
< 1% SHSSS error
  S-SAD, 5 keV?

## Slide 13
Damage is done
by dose (MGy)
proportional to photons/area
not time
not heat
Sliz P, Harrison SC & Rosenbaum G (2003). Structure 11, 13-19.
Garman EF & McSweeney SM (2006). J. Sync. Rad. 14, 1-3.
Owen RL, Rudino-Pinera E & Garman EF (2006). PNAS 103, 4912-4917.
Leiros et al. (2006). Acta Cryst. D 62, 125-132.
Holton JM (2007). J. Synch Rad. 14, 51-72.

## Slide 14
The amount of data you get
from any given crystal
is independent
of data collection time
1 um3 = 106 photons

## Slide 15: 10 MGy/Å
what the is a MGy?
http://bl831.als.lbl.gov/
damage_rates.pdf
Holton J. M. (2009) J. Synchrotron Rad. 16 133-42

## Slide 16
How long will my crystal last?
synch line type flux beamsize flux density dose max xtal min site
ph/s μm ph/μm2/s rate lifetime lifetime
ALS 4.2.2 MAD 2.2e11 75x80 3.7e+07 27.3 kGy/s 18 m 73 s
ALS 5.0.1 mono 1.6e11 100 2.0e+07 10.4 kGy/s 48 m 3.2 m
ALS 5.0.2 MAD 8e11 100 1.0e+08 51.8 kGy/s 9.6 m 39 s
ALS 5.0.3 mono 1.7e11 100 2.2e+07 11 kGy/s 45 m 3 m
ALS 8.2.1 MAD 1.8e11 100 2.3e+07 11.7 kGy/s 43 m 2.9 m
ALS 8.2.2 MAD 2.3e11 100 2.9e+07 14.9 kGy/s 34 m 2.2 m
ALS 8.3.1 MAD 9e11 70 2.3e+08 119 kGy/s 4.2 m 17 s
ALS 8.3.1 typical 6e11 70 1.6e+08 115 kGy/s 4.3 m 17 s
ALS 12.3.1 MAD 1.8e11 100 2.3e+07 11.7 kGy/s 43 m 2.9 m
ALS 12.3.1 ML 4.0e12 100 5.1e+08 513 kGy/s 58 s 3.9 s
APS 8-BM MAD 1e11 200 2.5e+06 1.27 kGy/s 6.6 h 26 m
APS 14-BM-C mono 5.8e10 200 1.4e+06 738 Gy/s 11 h 45 m
APS 14-BM-D MAD 3.3e9 200 8.2e+04 42 Gy/s 8.3 d 13 h
APS 14-ID-B MAD 6.0e10 200 1.5e+06 763 Gy/s 11 h 44 m
APS 17-BM MAD 1.1e11 200 2.8e+06 1.4 kGy/s 6 h 24 m
APS 17-ID MAD 2.3e11 200 5.8e+06 2.93 kGy/s 2.8 h 11 m
APS 19-BM MAD 2.0e11 70x60 4.8e+07 24.2 kGy/s 21 m 83 s
APS 19-ID MAD 1.3e13 80x40 4.1e+09 2.07 MGy/s 15 s 0.97 s
APS 19-ID typical 5.5e11 100x100 5.5e+07 28 kGy/s 18 m 71 s
APS 22-BM MAD 7e12 80x40 2.2e+09 1.23 MGy/s 24 s 1.6 s
APS 22-ID MAD 7e12 80x40 2.2e+09 1.23 MGy/s 24 s 1.6 s
APS 22-ID typical 1.5e12 80 2.3e+08 119 kGy/s 4.2 m 17 s
APS 23-ID-B MAD 1e13 75x25 5.3e+09 3.01 MGy/s 10 s 0.66 s
APS 23-ID typical 1.5e12 80 2.3e+08 119 kGy/s 4.2 m 17 s
APS 24-ID-C MAD 1.3e13 20x60 1.1e+10 5.23 MGy/s 5.7 s 0.38 s
APS 24-ID-E MAD 0.5e13 20x100 2.5e+09 1.19 MGy/s 25 s 1.7 s
APS 31-ID MAD 2e12 70 4.1e+08 194 kGy/s 2.6 m 10 s

## Slide 18
Holton & Frankel (2010) Acta D 66 393-408.

## Slide 19
Where:
IDL - average damage-limited intensity (photons/hkl) at a given resolution
105 - converting R from μm to m, re from m to Å, ρ from g/cm3 to kg/m3 and MGy to Gy
re - classical electron radius (2.818 x 10-15 m/electron)
h - Planck’s constant (6.626 x 10-34 J∙s)
c - speed of light (299792458 m/s)
fdecayed - fractional progress toward completely faded spots at end of data set
ρ - density of crystal (~1.2 g/cm3)
R - radius of the spherical crystal (μm)
λ - X-ray wavelength (Å)
fNH - the Nave & Hill (2005) dose capture fraction (1 for large crystals)
nASU - number of proteins in the asymmetric unit
Mr - molecular weight of the protein (Daltons or g/mol)
VM - Matthews’s coefficient (~2.4 Å3/Dalton)
H - Howells’s criterion (10 MGy/Å)
θ - Bragg angle
a2 - number-averaged squared structure factor per protein atom (electron2)
Ma - number-averaged atomic weight of a protein atom (~7.1 Daltons)
B - average (Wilson) temperature factor (Å2)
μ - attenuation coefficient of sphere material (m-1)
μen - mass energy-absorption coefficient of sphere material (m-1)
Self-calibrated damage limit
Holton & Frankel (2010) Acta D 66 393-408.
No flux
No symmetry

## Slide 20
Crystal diameter (micron)
Resolution (Å)
B = 0
B = 20
B = 60
B = 170
B = 330
B = 560
B = 850
Bigger is better, but not by much

## Slide 21: B factor from image analysis
B = 500

## Slide 22: B factor from image analysis
B = 20

## Slide 23: lysozyme: real and reciprocal

## Slide 24: lysozyme: thermal motion

## Slide 25: Muybridge’s galloping horse (1878)
Real space
reciprocal
“Time-resolved” diffraction

## Slide 26: Muybridge’s galloping horse (1878)
Real space
reciprocal
Average electron density

## Slide 27
Home Lab Resolution (Å)
Synchrotron Resolution (Å)
Predicting resolution limits

## Slide 28
Synchrotron Resolution (Å)
XFEL Resolution (Å)
Predicting resolution limits

## Slide 29
X-ray beam size (μm)
Resolution (Å)
too big
too small
just right

## Slide 30: “8-shooter” adjustable beam size
rotary
pinholes:
300
100
75
50
30
25
10
“small”

## Slide 31: How to get time at LCLS
Can’t be done at synchrotron
Showcase LCLS
Won’t waste valuable beam time ($20/s)
Work there
Hot biological project
Tons of small crystals

## Slide 32: Advanced Light Source

## Slide 33: Advanced Light Source
1.9 GeV | 500 mA
2 x 0.04 nm·rad | 
251 x 8 μm | 
63 m dia. | 12 sectors
175 staff | 43 beamlines

## Slide 34: Advanced Light Source
5.0.2
Macromolecular
Crystallography
Facility
ALS MX ca 1996

## Slide 35: Advanced Light Source
5.0.2
Macromolecular
Crystallography
Facility
ALS MX ca 1996

## Slide 36: Advanced Light Source
5.0.2
5.0.1
5.0.3
Macromolecular
Crystallography
Facility
ALS MX ca 2000

## Slide 37: Advanced Light Source
5.0.2
5.0.1
5.0.3
Macromolecular
Crystallography
Facility
ALS MX ca 2001

## Slide 38: Advanced Light Source
5.0.2
5.0.1
5.0.3
8.3.1
Macromolecular
Crystallography
Facility
ALS MX ca 2001
UCB/UCSF

## Slide 39: ALS 8.3.1 optics
ADSC Quantum 315r
Superbend
Plane
Parabolic
mirror
Torroidal
mirror
Si(111)
monochromator
Protein Crystal
pinhole
Scatter
guard
divergence
slits
MacDowell et. al. (2004) J. Sync. Rad. 11 447-455
2:1 demagnification cancels spherical aberrations
Pt/Rh mirrors 1.3 x 108 ph/s/μm2 ≈ 300 s/xtal
total power dissipation: 100 W
Notes: Most notable acheivements are high brightness with very little “wasted” light. The discovery made by Howard Padmore? At ALS showed that a 2:1 demagnification cancels nearly all of the spherical abberations in the parabolic/torroid system. 30-50% of the total x-ray intensity passes through the 100um pinhole. The flux through the 100um pinhole at the superbend beamlines is comparable to that of the 5.0.2 wiggler or an APS bending magnet, but only 150 Watts of power must be dissipated by the beamline optics. In contrast, wiggler beamlines must disipate kilowatts. The optics here are much simpler and not exotic in any way. The system is stable and robust. There is never a need to re-focus the beam at a new wavelength. Only a 2-minute automatic tuneup is recommended for very large energy changes (more than 1-2 keV).

## Slide 40: ALS 8.3.1 optics
Dectris Pilatus3 6M
Superbend
Plane
Parabolic
mirror
Torroidal
mirror
Si(111)
monochromator
Protein Crystal
pinhole
Scatter
guard
divergence
slits
MacDowell et. al. (2004) J. Sync. Rad. 11 447-455
2:1 demagnification cancels spherical aberrations
Pt/Rh mirrors 1.3 x 108 ph/s/μm2 ≈ 300 s/xtal
total power dissipation: 100 W
Notes: Most notable acheivements are high brightness with very little “wasted” light. The discovery made by Howard Padmore? At ALS showed that a 2:1 demagnification cancels nearly all of the spherical abberations in the parabolic/torroid system. 30-50% of the total x-ray intensity passes through the 100um pinhole. The flux through the 100um pinhole at the superbend beamlines is comparable to that of the 5.0.2 wiggler or an APS bending magnet, but only 150 Watts of power must be dissipated by the beamline optics. In contrast, wiggler beamlines must disipate kilowatts. The optics here are much simpler and not exotic in any way. The system is stable and robust. There is never a need to re-focus the beam at a new wavelength. Only a 2-minute automatic tuneup is recommended for very large energy changes (more than 1-2 keV).

## Slide 41: ALS 5.0.2 optics
16-pole wiggler
Plane
Parabolic
mirror
Torroidal
mirror
Si(111)
monochromator
Protein Crystal
divergence
slits
Dectris Pilatus3 6M
Zwart et. al. (2015) SRN 28 22-27
Sample
slits
1 x 108 ph/s/μm2 ≈ 300 s/xtal
Notes: Most notable acheivements are high brightness with very little “wasted” light. The discovery made by Howard Padmore? At ALS showed that a 2:1 demagnification cancels nearly all of the spherical abberations in the parabolic/torroid system. 30-50% of the total x-ray intensity passes through the 100um pinhole. The flux through the 100um pinhole at the superbend beamlines is comparable to that of the 5.0.2 wiggler or an APS bending magnet, but only 150 Watts of power must be dissipated by the beamline optics. In contrast, wiggler beamlines must disipate kilowatts. The optics here are much simpler and not exotic in any way. The system is stable and robust. There is never a need to re-focus the beam at a new wavelength. Only a 2-minute automatic tuneup is recommended for very large energy changes (more than 1-2 keV).

## Slide 42: Advanced Light Source
5.0.2
5.0.1
5.0.3
8.3.1
Macromolecular
Crystallography
Facility
ALS MX ca 2001
UCB/UCSF

## Slide 43: Advanced Light Source
5.0.2
5.0.1
5.0.3
8.2.2
8.3.1
8.2.1
Macromolecular
Crystallography
Facility
ALS MX ca 2002
UCB/UCSF

## Slide 44: Advanced Light Source
5.0.2
5.0.1
5.0.3
8.2.2
8.3.1
8.2.1
ALS MX ca 2002
Berkeley
Center for
Structural
Biology
UCB/UCSF

## Slide 45: Advanced Light Source
5.0.2
5.0.1
5.0.3
MBC 4.2.2
8.2.2
8.3.1
8.2.1
Berkeley
Center for
Structural
Biology
ALS MX ca 2005
UCB/UCSF

## Slide 46: ALS 4.2.2 optics
Research Detectors CMOS 8M
Superbend
Plane
Parabolic
mirror
Plane
parabolic
mirror
sagital
monochromator
Protein Crystal
Sample
slits
divergence
slits
1 x 108 ph/s/μm2 ≈ 300 s/xtal
Notes: Most notable acheivements are high brightness with very little “wasted” light. The discovery made by Howard Padmore? At ALS showed that a 2:1 demagnification cancels nearly all of the spherical abberations in the parabolic/torroid system. 30-50% of the total x-ray intensity passes through the 100um pinhole. The flux through the 100um pinhole at the superbend beamlines is comparable to that of the 5.0.2 wiggler or an APS bending magnet, but only 150 Watts of power must be dissipated by the beamline optics. In contrast, wiggler beamlines must disipate kilowatts. The optics here are much simpler and not exotic in any way. The system is stable and robust. There is never a need to re-focus the beam at a new wavelength. Only a 2-minute automatic tuneup is recommended for very large energy changes (more than 1-2 keV).

## Slide 47: Advanced Light Source
5.0.2
5.0.1
5.0.3
MBC 4.2.2
8.2.2
8.3.1
8.2.1
Berkeley
Center for
Structural
Biology
ALS MX ca 2005
UCB/UCSF

## Slide 48: Advanced Light Source
5.0.2
5.0.1
5.0.3
MBC 4.2.2
8.2.2
8.3.1
8.2.1
Berkeley
Center for
Structural
Biology
ALS MX ca 2006
UCB/UCSF
12.3.1 SIBYLS
MX & SAXS

## Slide 49: ALS 12.3.1 optics
ADSC Quantum 315r
Superbend
Plane
Parabolic
mirror
Torroidal
mirror
Si(111)
monochromator
Protein Crystal
divergence
slits
Sample
slits
4 x 107 ph/s/μm2 ≈ 1600 s/xtal
Notes: Most notable acheivements are high brightness with very little “wasted” light. The discovery made by Howard Padmore? At ALS showed that a 2:1 demagnification cancels nearly all of the spherical abberations in the parabolic/torroid system. 30-50% of the total x-ray intensity passes through the 100um pinhole. The flux through the 100um pinhole at the superbend beamlines is comparable to that of the 5.0.2 wiggler or an APS bending magnet, but only 150 Watts of power must be dissipated by the beamline optics. In contrast, wiggler beamlines must disipate kilowatts. The optics here are much simpler and not exotic in any way. The system is stable and robust. There is never a need to re-focus the beam at a new wavelength. Only a 2-minute automatic tuneup is recommended for very large energy changes (more than 1-2 keV).

## Slide 50: ALS 12.3.1 optics
ADSC Quantum 315r
Superbend
Plane
Parabolic
mirror
Torroidal
mirror
1% ML
monochromator
Protein Crystal
divergence
slits
Sample
slits
1.5 x 109 ph/s/μm2 ≈ 40 s/xtal
Notes: Most notable acheivements are high brightness with very little “wasted” light. The discovery made by Howard Padmore? At ALS showed that a 2:1 demagnification cancels nearly all of the spherical abberations in the parabolic/torroid system. 30-50% of the total x-ray intensity passes through the 100um pinhole. The flux through the 100um pinhole at the superbend beamlines is comparable to that of the 5.0.2 wiggler or an APS bending magnet, but only 150 Watts of power must be dissipated by the beamline optics. In contrast, wiggler beamlines must disipate kilowatts. The optics here are much simpler and not exotic in any way. The system is stable and robust. There is never a need to re-focus the beam at a new wavelength. Only a 2-minute automatic tuneup is recommended for very large energy changes (more than 1-2 keV).

## Slide 51: ALS 12.3.1 double endstation

## Slide 52: Advanced Light Source
5.0.2
5.0.1
5.0.3
MBC 4.2.2
8.2.2
8.3.1
8.2.1
Berkeley
Center for
Structural
Biology
ALS MX ca 2006
UCB/UCSF
12.3.1 SIBYLS
MX & SAXS

## Slide 53: Advanced Light Source
5.0.2
5.0.1
5.0.3
MBC 4.2.2
8.2.2
8.3.1
8.2.1
12.3.1 SIBYLS
MX & SAXS
Berkeley
Center for
Structural
Biology
ALS ca 2015
UC

## Slide 54: Advanced Light Source
5.0.2
5.0.1
5.0.3
MBC 4.2.2
8.2.2
8.3.1
8.2.1
12.3.1 SIBYLS
MX & SAXS
Berkeley
Center for
Structural
Biology
ALS ca 2018
UC
Gemini 2.0.1
μfocus 2.0.2

## Slide 55: ALS robots
5.0.2
5.0.1
5.0.3
MBC 4.2.2
ACTOR
8.2.2
8.3.1
Cool Hand Luke
8.2.1
12.3.1 SIBYLS
ACTOR
Berkeley
Auto-
Mounter
ALS MX robots
Stanford
Auto-
Mounter

## Slide 56: Pin compatibility
SLAC
ALS
SPINE

## Slide 57: SuperTong

## Slide 58: What if we didn’t have to choose?
Caltech
SSRL
Yale
Oxford
SPring-8
SPINE
MiTeGen RT
Syrrx
H. magnetic
H. hybrid
H. copper
H. screw cap
H. colored
Hampton screw cap
Hampton magnetic

## Slide 59: Hampton Pin
SuperTong

## Slide 60: Syrrx Pin
SuperTong

## Slide 61: plastic Pin
SuperTong

## Slide 62: Yale Pin
SuperTong

## Slide 63: SuperTong
SuperTong

## Slide 64: Robix ongonio

## Slide 65: CHL idlepos

## Slide 66: Carousel open

## Slide 67: “infinite capacity” sample carousel

## Slide 68: Global “success rate” of structure determination
100 s/dataset
200 days/year
~150 beamlines
~26,000,000 datasets/year
9640 PDBs in 2014
0.04% “success”

## Slide 69: Measured success rate: ALS 8.3.1
images x 106
1250
1000
750
500
250
0
PDB entries
Notes: Of the 1275 PDB deposits involving ALS 8.3.1, 1038 of them mention a data collection date. The timeline between collection and deposition for each PDB entry is shown in blue. For example, in April 2004 only 46 PDBs had been deposited, but the data for 252 had already been collected. The red line indicates the cumulative number of images collected at a given date. The kink in the red line in mid-2003 is coincident with an optics upgrade that tripled the flux. In 2014 another flux upgrade brought another factor of 4.5.

## Slide 70: Turning data into models
Number | Description | Percent
446028 | Images (~7 TB) | 33%
2346 | Data sets | 47%
449 | MAD/SAD (1:2) | 19%
120 | Published | 5.1%
8.3.1 in 2004

## Slide 71: Operational Efficiency
Seconds | Description | Percent
104490 | Assigned and available | 100%
42093 | Shutter open | 40%
52684 | Collecting (3026 images) | 50%
247s  45 | Mounting | 11%
229s  37 | Centering | 8%
 | Thinking | 31%

## Slide 72: The 100 structure/year club
PDBs | Sync | BL | Flux
2829 | APS | 19-ID | 1.3E+13
2110 | NSLS | X29A | 2.9E+11
1830 | APS | 22-ID | 7.0E+12
1651 | SLS | X06SA | 4.0E+12
1517 | SLS | X10SA | 2.5E+12
1475 | ESRF | ID29 | 1.0E+13
1447 | SSRF | BL17U | 4.0E+12
1384 | ESRF | ID14-4 | 5.0E+12
PDBs | Sync | BL | Flux
1384 | ESRF | ID14-4 | 5.0E+12
1362 | ESRF | ID23-1 | 4.0E+12
1285 | SSRL | BL11-1 | 4.0E+11
1282 | APS | 24-ID-C | 1.3E+13
1153 | ESRF | ID14-1 | 1.0E+11
1137 | SSRL | BL9-2 | 7.0E+11
1122 | ALS | 8.3.1 | 1.8E+11
1042 | APS | 17-ID | 7.0E+11
1026 | NSLS | X25 | 2.4E+11
USA
Europe
Decommissioned
8% of all membrane
protein structures!

## Slide 73: “The flattest thing we have ever seen”
Old mirror: 800 nRad
New mirror: < 200 nRad
Dec 17 2014
Flux at
460%
- LBNL Metrology Group

## Slide 75: What will change?
Old detector
Read-out noise: 2 ph/pix
  100 ADU/pixel
1.15s read-out, ~0.3 Hz
  23 min data set @ 30 MGy
  rastering impractical
16 bits
0.7% inactive area
315 mm face; 80 μm PSF
  1312 orders across face
No max count rate
3% SHSSS error
  7235 eV optimal
New detector
No read-out noise
  no more strategy
2.03 ms read-out, 25 Hz
  200 s data set @ 30 MGy
  10 min/mm2 raster 1% lifedose
20 bits
8.5% inactive area
430 mm face, 172 μm PSF
  833 orders across face
Max 107 ph/s/pix
< 1% SHSSS error
  S-SAD, 5 keV?

## Slide 76: MLFSOM: now available in “Pilatus”

## Slide 77: Pilatus pile-up for RT MX? same photons, different speeds
1 shot with
no attenuation

## Slide 78: Pilatus pile-up for RT MX? same photons, different speeds
Sum of 193 shots with
193-fold attenuation

## Slide 79: Pilatus pile-up for RT MX: same photons, different speeds
Collaboroators:
Aina Cohen
Ana Gonzalez

## Slide 80: Pilatus pile-up for RT MX: same photons, different speeds
Collaboroators:
Aina Cohen
Ana Gonzalez

## Slide 81: Pilatus pile-up for RT MX: same photons, different speeds
Collaboroators:
Aina Cohen
Ana Gonzalez

## Slide 82: What will change?
Old detector
Read-out noise: 2 ph/pix
  100 ADU/pixel
1.15s read-out, ~0.3 Hz
  23 min data set @ 30 MGy
  rastering impractical
16 bits
0.7% inactive area
315 mm face; 80 μm PSF
  1312 orders across face
No max count rate
3% SHSSS error
  7235 eV optimal
New detector
No read-out noise
  no more strategy
2.03 ms read-out, 25 Hz
  200 s data set @ 30 MGy
  10 min/mm2 raster 1% lifedose
20 bits
8.5% inactive area
430 mm face, 172 μm PSF
  833 orders across face
Max 107 ph/s/pix
< 1% SHSSS error
  S-SAD, 5 keV?

## Slide 83: Averaging with systematic error

## Slide 84
Gd edge-contrast map of ADSC Q315r detector
10% high
average
10% low

## Slide 85: Gadox calibration vs energy
photon energy (keV)
Relative absorption depth
same = good!
bad!

## Slide 86
Pick-up tool mark
Bragg
glitch
oxygen
inclusions
“tree rings”
1% high
average
1% low
~3x105 photon/pixel → 0.3% error
Pilatus: subtract smooth baseline
Gollwitzer & Krumrey (2014)
J. Appl. Cryst. 47, 378

## Slide 87: Detector pixels are 3D
51 μm
31 μm thick
17%
loss
Typical CCD
Mean Penetration depth 12 keV:
17.5 μm Gd2O2S 237 μm Si
8%
side
loss
172 μm
26% loss
320 μm thick
Pilatus
53%
side
loss
photon
75 μm
15%
loss
450 μm thick
Eiger
76%
side
loss
photon
Arrival angle of 2 Å spot: 29°

## Slide 88
N
N
N
θ
λ
detection
event
incoming
photon
PAD: need more than flood-field

## Slide 89
thickness
width
N
N
N
Bragg
glitch
oxygen
inclusion
θ
λ
detection
event
incoming
photon
PAD: need more than flood-field

## Slide 90
3.5
3.0
2.5
2.0
1.5
1.0
0.5
SHSSS
Systematic component of Rsseparate (%)
distance between spots (mm)
0.1 1 10 100
CCD detector
anomalous mates are always on
different modules
different
modules
Pilatus
SN113
Pilatus
SN001
Notes: Comparision of SHSSS measured by the detector-shift experiment on three different detectors:
The Q315R in ALS 8.3.1
The first Pilatus detector at the Swiss Light Soruce
A newer Pilatus detector at Stanford
Improvements in calibration technology have finally made
Pilatus competitive with calibration on CCDs

## Slide 91: Beam Flicker
time (seconds)
normalized flux through pinhole
pinhole
removed
0.15%/√Hz

## Slide 92
Source of error | realistic simulation | No SHSSS | Perfect detector
Photon counting | + | + | +
Shutter jitter | + | + | -
Beam flicker | + | + | -
Sample absorption | + | + | -
Radiation damage | + | + | -
Imperfect spindle | + | + | -
vignette | + | + | -
Corner correction | + | + | -
SHSSS | + | - | -
Rmeas (∞-10 Å) | 2.8% | 0.7% | 0.7%
I/σ asymptotic | 26.8 | 74.2 | 81.0
Threshold of a revolution in phasing
Holton et al (2014) "R-factor gap", FEBS Journal 281, 4046-4060.

## Slide 93
Required signal-to-noise (I/σ)
Solve-able proteins (%)
CCD
& early PAD
new PAD
(uncalibrated)
Goal

## Slide 94: Can you count to 1,000,000 ?
= 0.1%
sqrt(1,000,000)
1,000,000
= 3%
sqrt(1,000)
1,000
> 1000 is a waste!
photon
spot
Theoretically:
In reality:
ISa ~ 33
Rmeas ≈ 0.1% ?
ISa = 1000
Rmeas = ≈ 3%

## Slide 95: 140-fold multiplicity
SUBSET OF INTENSITY DATA WITH SIGNAL/NOISE >= -3.0 AS FUNCTION OF RESOLUTION
RESOLUTION NUMBER OF REFLECTIONS COMPLETENESS R-FACTOR R-FACTOR COMPARED I/SIGMA R-meas CC(1/2) Anomal SigAno Nano
LIMIT OBSERVED UNIQUE POSSIBLE OF DATA observed expected Corr
9.17 188919 1061 1071 99.1% 3.9% 4.8% 188919 257.47 3.9% 100.0* 91* 5.024 450
6.49 356827 1933 1933 100.0% 5.2% 5.5% 356827 214.33 5.2% 100.0* 86* 3.836 882
5.30 466503 2515 2515 100.0% 7.2% 7.0% 466503 165.13 7.2% 100.0* 76* 3.257 1175
4.59 548405 2974 2974 100.0% 7.2% 6.6% 548405 175.42 7.3% 100.0* 67* 2.589 1403
4.10 615117 3353 3355 99.9% 7.7% 6.7% 615117 174.13 7.7% 100.0* 59* 2.264 1594
3.74 684947 3734 3737 99.9% 9.4% 8.3% 684947 143.09 9.4% 100.0* 49* 1.953 1783
3.47 743557 4051 4051 100.0% 11.2% 10.1% 743557 120.17 11.2% 100.0* 39* 1.696 1942
3.24 799722 4366 4366 100.0% 14.1% 13.9% 799722 91.14 14.1% 100.0* 30* 1.333 2103
3.06 839252 4598 4603 99.9% 19.5% 20.2% 839252 65.79 19.5% 100.0* 23* 1.117 2214
2.90 891398 4904 4908 99.9% 29.0% 31.7% 891398 44.85 29.1% 99.9* 17* 1.008 2369
2.77 931622 5148 5151 99.9% 40.5% 44.8% 931622 32.58 40.6% 99.8* 11* 0.901 2493
2.65 969629 5384 5387 99.9% 52.8% 58.8% 969629 25.16 52.9% 99.8* 10* 0.866 2605
2.54 997213 5574 5574 100.0% 67.4% 76.0% 997213 19.47 67.6% 99.6* 2 0.804 2705
2.45 1031222 5889 5889 100.0% 88.9% 101.2% 1031222 14.58 89.2% 99.2* 4 0.831 2859
2.37 788617 6008 6008 100.0% 109.3% 125.5% 788617 9.97 109.7% 98.1* 5 0.829 2925
2.29 614145 6252 6252 100.0% 138.2% 161.4% 614145 6.87 138.9% 96.1* 1 0.760 3037
2.22 470283 6481 6481 100.0% 197.1% 231.7% 470283 4.03 198.6% 83.5* -1 0.721 3159
2.16 259836 6629 6631 100.0% 227.3% 268.7% 259831 2.41 230.8% 46.9* -1 0.677 3224
2.10 36259 4169 6807 61.2% 154.4% 169.4% 36229 1.28 163.6% 47.0* -2 0.660 1999
2.05 13567 3372 7037 47.9% 170.1% 187.0% 13503 0.68 196.5% 25.7* 3 0.629 1578
total 12247040 88395 94730 93.3% 15.7% 16.4% 12246941 54.30 15.8% 100.0* 12* 1.217 42499
16 crystals, 360° each, < 1 MGy, inverse beam
7235 eV ADSC Q210r Australian Synchrotron MX1
I/SIGMA
257.47

## Slide 96: 140-fold multiplicity
18 σ
Phased anomalous difference Fourier
data Courtesy of Tom Peat & Janet Newman
16 σ
Not
Lysozyme!

## Slide 97: 140-fold multiplicity
15σ = PO4
Phased anomalous difference Fourier
data Courtesy of Tom Peat & Janet Newman
Not
Lysozyme!

## Slide 98: 140-fold multiplicity
~2σ = Mg?
Phased anomalous difference Fourier
data Courtesy of Tom Peat & Janet Newman
Not
Lysozyme!

## Slide 99: 140-fold multiplicity
8.2σ = Na
DELFAN residual anomalous difference
data Courtesy of Tom Peat & Janet Newman
10 σ is enough for phasing!
Bunkoczi et al. (2015)."weak, single-wavelength anomalous", Nat. Meth. 12, 127-130.
Not
Lysozyme!
f” = 0.15

## Slide 100
ADSC
area
detector
beam center
3%
2%
1%
0
crystal in loop
Cu foil
X-ray excited
fluorescent
X-ray source
x-ray beam

## Slide 101: Cone-Beam Online X-ray Absorption Radiograph

## Slide 102: What will change?
Old detector
Read-out noise: 2 ph/pix
  100 ADU/pixel
1.15s read-out, ~0.3 Hz
  23 min data set @ 30 MGy
  rastering impractical
16 bits
0.7% inactive area
315 mm face; 80 μm PSF
  1312 orders across face
No max count rate
3% SHSSS error
  7235 eV optimal
New detector
No read-out noise
  no more strategy
2.03 ms read-out, 25 Hz
  200 s data set @ 30 MGy
  10 min/mm2 raster 1% lifedose
20 bits
8.5% inactive area
430 mm face, 172 μm PSF
  833 orders across face
Max 107 ph/s/pix
< 1% SHSSS error
  S-SAD, 5 keV?