## Slide 1: BL8.3.1 mono cryo upgrade Motivation
ALS-U $650M to make beam smaller
With current optics: no benefit from ALS-U
High-profile programs at ALS MX beamlines
  i.e. 2020 Nobel Prize
30x brightness gain is possible with ALS-U (5x at ALS)
$800k joint LDRD project: MBIB and ALS to match hard X-ray optics to ALS-U source properties
Must spearhead optics upgrades for ALS-U success
Cryo-cooling monochromators required

## Slide 2: Planned work overview
Cryo-cooling upgrade to BL8.3.1 monochromator
Keep original tank & all lead (Pb) sheets
Similar to 8.2.1, 5.0.2 monos
  Less radical: contents are Kohzu, not Oxford
  $110k vs $700k
Same type FMB Oxford cryo-cooler
Simpler, longer Si(111) crystals
  No more Z2 motor
Safety systems: EPS, seismic, burst discs
Document & draw Kohzu mono contents
focus

## Slide 3: 1st crystal
old
Si(111)
new
70L x 40W x 40T mm
Si(111)
50L x 60W x 38T mm
Thermal load: 60 W

## Slide 4: 2nd crystal
old
new
165L x 40W x 30T mm
90L x 51W x 40T mm
Thermal load: << 1 W

## Slide 5: No more Z2 motor

## Slide 6
Kohzu’s upgrade kit

## Slide 7
Kohzu’s upgrade kit

## Slide 10
64
152
250
70
91
115
42
36
227
25 x 88
35
600 dia

## Slide 11: Current BL8.3.1 Mono Contents
30mm SS
3 mm Lead (Pb)
(Cu) Compton Shield
“hanging”
Pb
panel

## Slide 12
Current situation

## Slide 13
Current
After modification
227
35
227
35

## Slide 14
Current
After modification

## Slide 15
After modification compared a current shield (pink)

## Slide 17
Design compromise:
Cooling line vibration
vs continuous Pb

## Slide 18: Kohzu’s proposed install schedule
JULY-31 (Wed): Start working at ALS AUG-1 (Thu): Working AUG-2 (Fri): Finish & start pumping down AUG-3 (Sat): Holiday AUG-4 (Sun): Holiday AUG-5 (Mon): Leak check, Tidying-up, Done
Takashi Suzuoki <suzuoki@kohzu.co.jp> Daisuke Shindo <shindo@kohzu.co.jp>

## Slide 19: Don’t need help with:
Remove/replace “hanging” Pb plate
Mono internals install – Kohzu techs
Connect and strain-relief cryo lines
Condensate drip management

## Slide 20
Condensate drip management
Current setup on BL8.2.1

## Slide 21: Condensate drip prevention
Will work with BCSB on fill lines
Best solutions
  Vacuum insulate lines and joints
  Passive/active heated Al cladding
  clear enclosure, purged with N2
N2

## Slide 22: Cryogen, overpressure and vacuum safety
Cryostat is entirely vacuum insulated
  all external surfaces at approximately room temperature
Boiloff gas discharge heated and directed down
All closed (or potentially closed) volumes have burst disks and/or pressure relief protection
Cryocooler has independent monitoring system with limits, audible klaxon alarms and emergency staff paging capability
Mono cooling circuit includes direct vacuum/cryogen seals hence:-
  Tank will be fitted with overpressure burst disk
  In-vacuum electronics protected by pressure interlock covering low conductivity pressure regime
  Ring vacuum is protected by front-end double Be window
  In-vacuum leaks of dry LN preferable to water

## Slide 23: Could use help with…
EPS – integrate Cryostat “ok” relay output
Seismic – same as other cryostats
  Replace eight wheels with 2” floor anchors
  < 100 lbs/anchor
Spec burst discs
Electrical power feed
Cryo-cooler placement

## Slide 24
Cryo cooler placement

## Slide 25: Back-out plan and crazy ideas
Re-install old mono parts
  not expected, last resort
Modify Pb to fit cryo rig - preferable
Modify cryo rig to fit Pb - possible
Run water through new mono parts
  In event of total cryo show-stopper
Daisy-chain 8.2.1 and 8.3.1 monos
  May be valuable solution in future
Alternate cooler: N2 gas at 10 L/s
  Very much lower cost

## Slide 26: Expected schedule
Jun 25 – open mono, measure Pb plate
July 1 – completed power, EPS
July 30 – burst discs, cooler placement
Aug 5 – Kohzu upgrade complete
Sep 15 – Cryo cooler install complete
Oct 13 – BL Commissioning

## Slide 27
Cu
Compton
Shield
“hanging”
Pb
sheet