## Slide 1: ALS 8.3.1 Upcoming Upgrades
Unattended screening
Kappa goniostat
Adjustable beam size 300 μm – 10 μm
Multilayers for rastering
ALS-U in 2022

## Slide 2: “rastering” to find the crystal
Membrane
protein crystal
X-ray beam
ice

## Slide 3: Why not just “raster” everything?
Rastering Needs:
Small beam
Very high flux
fast detector
Rapid image analysis
Small search area

## Slide 4: Rastering is robust, but slow
Damage & diffraction = flux / area
Gy/s = ph/s /2000 / um2
Example: flux= 1012 ph/s, 1 mm2 beam = 500 Gy/s
100% sample life → 30 MGy / 500 Gy/s = 17 hours
10% sample life → 3 MGy / 500 Gy/s = 1.7 hours
1% sample life → 300 kGy / 500 Gy/s = 10 min
Holton (2009)."beginner's guide rad dam", J. Synch. Rad. 16, 133-142.

## Slide 5: Rastering is robust, but slow
Damage & diffraction = flux / area
Gy/s = ph/s /2000 / um2
Example: 10% life, 1 mm2 sample, flux= 1012 ph/s
10x10 μm beam: 0.6 s
100x100 μm beam: 60s x 10x10 shots =
1% life & 1013 ph/s → 1 min
Holton (2009)."beginner's guide rad dam", J. Synch. Rad. 16, 133-142.
1% life → 10 min
1.7 hours
x 100x100 shots =
1.7 hours

## Slide 6: Rastering is robust, but slow
Damage & diffraction = flux / area
time/mm2 is fixed by flux
Example: ALS 8.3.1 – 200 kGy/s = 15 s for 10% lifedose
50 μm hole: 20x20 shots = 210 s/mm2
30 μm hole: 33x33 shots = 320 s/mm2
10 μm hole: 100x100 shots = 2730 s/mm2
10 μm hole + multilayer: 70 s/mm2
Holton (2009)."beginner's guide rad dam", J. Synch. Rad. 16, 133-142.

## Slide 7: Rastering is robust, but slow
Damage & diffraction = flux / area
time/mm2 is fixed by flux
Example: ALS 8.3.1 – 200 kGy/s ; 1% lifedose
50 μm hole: 20x20 shots = 21 s/mm2
30 μm hole: 33x33 shots = 32 s/mm2
10 μm hole: 100x100 shots = 273 s/mm2
10 μm hole + multilayer: 7 s/mm2
Holton (2009)."beginner's guide rad dam", J. Synch. Rad. 16, 133-142.

## Slide 8: Rastering is robust, but slow
Damage & diffraction = flux / area
time/mm2 is fixed by flux
Example: ALS 8.3.1 – 200 kGy/s ; 1% lifedose
50 μm hole: 20x20 shots = 21 s/mm2 20 Hz
30 μm hole: 33x33 shots = 32 s/mm2 35 Hz
10 μm hole: 100x100 shots = 273 s/mm2 37 Hz
10 μm hole + multilayer: 7 s/mm2 2 kHz
Holton (2009)."beginner's guide rad dam", J. Synch. Rad. 16, 133-142.

## Slide 9: Low cost, portable rastering solution
Sample crystal
In cryo-loop
(not to scale)
New
CMOS detector
($70,000)
20 frames/s
photon/noise = 2
Existing MX detector
0.4 frames/s
Photon/noise = 0.4
Existing
Retractable
Front cover
http://researchdetectors.com/cmos_1m.html
Collaboration: 8.3.1 and MBC beamlines

## Slide 11
MX
area
detector
beam center
3%
2%
1%
0
crystal in loop
Cu foil
X-ray excited
fluorescent
X-ray source
x-ray beam
Cone-Beam Online X-ray
Absorption Radiography
(CBOXAR)

## Slide 12: Cone-Beam Online X-ray Absorption Radiograph (CBOXAR)

## Slide 13
Cryo-jet
nozzle
Goniometer
magnet
cold gas
0.2% high
average
0.2% low
Contrast of 90% solvent
protein xtal in pure water: 0.2%

## Slide 14
pinhole
prism
microscope
Styrofoam™ backlight
backstop
Zero-parallax simplifies alignment

## Slide 15: Zero-parallax simplifies alignment
X-ray

## Slide 16: Zero-parallax simplifies alignment
visible light

## Slide 17: Automatic and Unattended Optical Re-Alignment (AUORA)
beamline
microscope
reference
image

## Slide 18: Automatic and Unattended Optical Re-Alignment (AUORA)
search

## Slide 19: Automatic and Unattended Optical Re-Alignment (AUORA)
search

## Slide 20: Automatic and Unattended Optical Re-Alignment (AUORA)
search

## Slide 21: Automatic and Unattended Optical Re-Alignment (AUORA)
search

## Slide 22: Automatic and Unattended Optical Re-Alignment (AUORA)
search

## Slide 23: Automatic and Unattended Optical Re-Alignment (AUORA)
match

## Slide 24: Automatic and Unattended Optical Re-Alignment (AUORA)
move

## Slide 25: Automatic and Unattended Optical Re-Alignment (AUORA)
aligned
4.5 s Acquire 360°
3.0 s Algorithm
5.0 s Motor moves

## Slide 48: Automatic and Unattended Optical Re-Alignment (AUORA)
~12 s mount→shutter Accuracy: ~2 pix ~1-2 °

## Slide 49: ALS 8.3.1 Upcoming Upgrades
Unattended screening
Kappa goniostat
Adjustable beam size 300 μm – 10 μm
Multilayers for rastering
ALS-U in 2022

## Slide 50
SmarAct’s
“SmarGon”
Commercialized
PRIGo

## Slide 51: ALS 8.3.1 Upcoming Upgrades
Unattended screening
Kappa goniostat
Adjustable beam size 300 μm – 10 μm
Multilayers for rastering
ALS-U in 2022

## Slide 52: “8-shooter” adjustable beam size
rotary
pinholes:
300
100
75
50
30
25
10
“small”

## Slide 53
X-ray beam size (μm)
Resolution (Å)
too big
too small
just right

## Slide 54
~0 MGy

## Slide 55
5 MGy

## Slide 56
11 MGy

## Slide 57
16 MGy

## Slide 58
22 MGy

## Slide 59
26 MGy

## Slide 60
De-focused beam intensity profile

## Slide 61: “The flattest thing we have ever seen”
Old mirror: 800 nRad
New mirror: < 200 nRad
- LBNL Metrology Group

## Slide 62: ALS 8.3.1 Upcoming Upgrades
Unattended screening
Kappa goniostat
Adjustable beam size 300 μm – 10 μm
Multilayers for rastering
ALS-U in 2022

## Slide 63: Fast “death ray” switching
Si(111)
Existing mono
Si(111)
ML
ML
0.014 % (5 min)

## Slide 64
Si(111)
Existing mono
ML
ML
1 % (~4 s)
Si(111)
Fast “death ray” switching

## Slide 65: Si(111) vs multilayers
0.014%
Si(111)
1%
multilayer
4 Å
1.9 Å
resolution
spectral
dispersion

## Slide 66: Wide-bandpass MX?
Mono | Si(111) | multilayer | Difference (%) | good
Energy (eV) | 11111 | 12027.5 | +8.2 | Same
Scale | 1.0000 | 1.0964 | +9.6 | Yes
Rmerge (%) | 7.5 | 8.4 | +12 | No
Ranom (%) | 5.6 | 5.5 | -1.8 | Same
I/sd | 13.2 | 12.3 | -6.8 | Same
I/sd (1.5Å) | 2.3 | 1.8 | -21.7 | No
Ano Patt | 5.96 | 4.75 | -20.3 | Same
Ano Phased | 41 | 39.33 | -4.1 | Same
FOM | 0.238 | 0.192 | -19.3 | Same
CC | 0.555 | 0.647 | -6.1 | Yes
Rcryst (%) | 26.70 | 26.48 | -0.2 | Same
Ligand peak | 8.213 | 8.250 | +0.4 | Same

## Slide 67: Wide bandpass reveals mosaicity from a still
Snell EH, Weisgerber S, Helliwell JR, Holzer K & Schroer K (1995)."Improvements in lysozyme protein crystal perfection through microgravity growth", Acta crystallographica. Section D, Biological crystallography 51, 1099-1102.
low mosaic
high mosaic
Δradial = 2ηD/cos22θ

## Slide 68: “rastering” to find the crystal
Membrane
protein crystal
X-ray beam
ice

## Slide 69
MX
area
detector
beam center
3%
2%
1%
0
crystal in loop
Cu foil
X-ray excited
fluorescent
X-ray source
x-ray beam
Cone-Beam Online X-ray
Absorption Radiography
(CBOXAR)

## Slide 70: Interleaved Scheduling
experiment queue
beamline
Minor 30s
Choe 120s
Alberta 60s
Alberta 60s
Choe 30s
Minor 30s

## Slide 71: ALS 8.3.1 Upcoming Upgrades
Unattended screening
Kappa goniostat
Adjustable beam size 300 μm – 10 μm
Multilayers for rastering
ALS-U in 2022

## Slide 72: ALS-U ring upgrade in 2022
Pros
Diffraction-limited source
Coherent up to 2 keV
Smaller source
6x flux onto 6 μm xtal
Undulator:
  5x flux of 8.3.1
  ~2x more flux after ALS-U
  0.7 μm beam vs 10 μm
Serial xtallography
Fine, fast rastering
Cons
1 year of dark
No coherence effect in MX
No flux increase (ph/s)
No change for xtals > 36 μm
$10M to keep superbend
Undulator = 5x 8.3.1 at ALS
Only 1 more undulator slot
Xtals die in 2.4 ms
Are your xtals 0.7 μm ?

## Slide 73: ALS-U: Small, soft beams

## Slide 74: ALS-U: Small, soft beams
Diamond
($100, reusable)

## Slide 75: ALS-U: Small, soft beams
Diamond
($100, reusable)
oil ?

## Slide 76: ALS-U: Small, soft beams
Diamond
($100, reusable)
oil ?

## Slide 77: ALS-U: Small, soft beams
oil ?
λ=2d sinθ
2.5 Å data
with 5 Å X-rays
Only analytic absorption correction:
International Tables for Crystallography, Vol. C, 2nd ed., chapter 6.3

## Slide 78
Photon Energy (eV)
Anomalous f” (electrons)
2472.7 2473.5
Bohic et al. Anal. Chem. 2008; 80(24):9557. doi: 10.1021/ac801817k

## Slide 79: S-MAD Sample size Dilemma

## Slide 80
Required signal-to-noise (I/σ)
Solve-able proteins (%)
Single-spot limit
CCD
& early PAD
Single-spot limit
latest PAD
(uncalibrated)
Single-spot limit
Perfect detector
Threshold of a revolution in phasing

## Slide 81: ALS 8.3.1 Upcoming Upgrades
Unattended screening
Kappa goniostat
Adjustable beam size 300 μm – 10 μm
Multilayers for rastering
ALS-U in 2022