## Slide 1: PowerPoint File available:
http://bl831.als.lbl.gov/
~jamesh/powerpoint/
CSHL_SvN_2015.pptx

## Slide 2: Acknowledgements
UCSF LBNL SLAC
ALS 8.3.1 creator: Tom Alber
Center for Structure of Membrane Proteins (CSMP)
Membrane Protein Expression Center II (MPEC)
Center for HIV Accessory and Regulatory Complexes (HARC)
UC Multicampus Research Programs and Initiatives (MRPI)
UCSF Program for Breakthrough Biomedical Research (PBBR) Integrated Diffraction Analysis Technologies (IDAT)
Plexxikon, Inc.
M D Anderson CRC
Synchrotron Radiation Structural Biology Resource (SLAC)
The Advanced Light Source is supported by the Director, Office of Science, Office of Basic Energy Sciences, Materials Sciences Division, of the US Department of Energy under contract No. DE-AC02-05CH11231 at Lawrence Berkeley National Laboratory.
Robert Stroud James Fraser John Spence
Chris Nielsen Clemens Schulze-Briese Aina Cohen Ana Gonzalez

## Slide 3: Snapshot from single virus particle
TEM
2 keV LCLS 200 fs Mimi virus single-shot.
200 nm
Reconstructed image
Resolution 20nm
Seibert, et al. (2011). Nature 470, 78-81.
Notes: what limits single-particle resolution ?
get 100 times more photons/shot if correctly focussed beam
get 100 times more with submicron beam focus on CXI chamber
Problem – scattering goes from N^2 to N (Wilson statistics).
So use a-priori information – modelling Axel Brummer
Make beam smaller than jet to elliminate streak
3

## Slide 4: The “success rate” of structure determination
100 s/dataset
200 days/year
~150 beamlines
~26,000,000 datasets/year
9640 PDBs in 2014

## Slide 5: signal vs noise
“If you don’t have good data, then you have no data at all.”
-Sung-Hou Kim

## Slide 6: signal vs noise
easy
hard
impossible

## Slide 7: signal vs noise
easy
hard
impossible
threshold of “solvability”

## Slide 8: signal vs noise
“If you don’t have good data, then you must learn statistics.”
-James Holton

## Slide 9: Adding noise

## Slide 10: Adding noise
12 + 12 = 1.42

## Slide 11: Adding noise
12 + 12 = 1.42
32 + 12 = 3.22
σtotal2 = σ12 + σ22

## Slide 12: Adding noise
12 + 12 = 1.42
32 + 12 = 3.22
σtotal2 = σ12 + σ22

## Slide 13: Adding noise
12 + 12 = 1.42
32 + 12 = 3.22
σtotal2 = σ12 + σ22

## Slide 14: Adding noise
12 + 12 = 1.42
32 + 12 = 3.22
102 + 12 = 10.052

## Slide 15: 100% SeMet incorporation
Trivial to solve

## Slide 16: 100% S -Met incorporation
Impossible to solve

## Slide 17: Phase Problem: Se vs S
fraction Sulfur
correlation coefficient

## Slide 18: Phase Problem: Se vs S
fraction Sulfur
correlation coefficient

## Slide 19: Phase Problem: Se vs S
fraction Sulfur
correlation coefficient

## Slide 20: Phase Problem: Se vs S
fraction Sulfur
correlation coefficient

## Slide 21: Phase Problem: Se vs S
fraction Sulfur
correlation coefficient

## Slide 22
29% Se 71% S

## Slide 23
28% Se 72% S

## Slide 24
21% Se 79% S

## Slide 25
11% Se 89% S

## Slide 26: Finding sites
fraction Sulfur
Number of correct sites

## Slide 27: Finding sites
fraction Sulfur
Number of correct sites

## Slide 28: Finding sites
fraction Sulfur
Number of correct sites

## Slide 29: Finding sites
fraction Sulfur
Number of correct sites
25
20
15
10
5
0
peak height ()

## Slide 30: Finding sites
fraction Sulfur
Number of correct sites

## Slide 31: MR simulation
Rmsd from perfect search model (Å)
corrupted model
Correlation coefficient to correct density

## Slide 32: The transitions are sharp!
How can we predict success/failure?
Know Thy Experiment

## Slide 33: Beamline optics … simplified
ADSC Quantum 315r
Protein Crystal
pinhole
Scatter
guard
X-ray
Source
Notes: Most notable acheivements are high brightness with very little “wasted” light. The discovery made by Howard Padmore? At ALS showed that a 2:1 demagnification cancels nearly all of the spherical abberations in the parabolic/torroid system. 30-50% of the total x-ray intensity passes through the 100um pinhole. The flux through the 100um pinhole at the superbend beamlines is comparable to that of the 5.0.2 wiggler or an APS bending magnet, but only 150 Watts of power must be dissipated by the beamline optics. In contrast, wiggler beamlines must disipate kilowatts. The optics here are much simpler and not exotic in any way. The system is stable and robust. There is never a need to re-focus the beam at a new wavelength. Only a 2-minute automatic tuneup is recommended for very large energy changes (more than 1-2 keV).

## Slide 34: X-ray data are 3D!

## Slide 35: Where do photons go?
beamstop
Transmitted (98%)
Protein
1A x-rays

## Slide 36: Elastic scattering

## Slide 37: Elastic scattering

## Slide 38: Where do photons go?
beamstop
elastic scattering (6%)
Transmitted (98%)
Protein
1A x-rays

## Slide 39: Elastic scattering

## Slide 40: Inelastic scattering

## Slide 41: Where do photons go?
beamstop
elastic scattering (6%)
Transmitted (98%)
Protein
1A x-rays

## Slide 42: Where do photons go?
beamstop
elastic scattering (6%)
Transmitted (98%)
inelastic scattering (7%)
Protein
1A x-rays

## Slide 43: Where do photons go?
beamstop
elastic scattering (6%)
Transmitted (98%)
inelastic scattering (7%)
Protein
1A x-rays
Re-emitted (99%)
Absorbed (~0%)

## Slide 44: Photoelectric absorption

## Slide 45: Photoelectric absorption
e-
+

## Slide 46: Fluorescence
+

## Slide 47: Fluorescence
e-
+

## Slide 48: Metal identification

## Slide 49: Auger emission
+

## Slide 50: Auger emission
++
e-

## Slide 51: Sample atoms

## Slide 52: Sample atoms

## Slide 53: Primary ionization
e-
+

## Slide 54: Secondary ionizations
e-
e-
e-
e-
e-
e-
+
+
+
+
+
+

## Slide 55: Ionization track
e-
e-
e-
e-
e-
e-
e-
e-
e-
+
+
+
+
+
+
+
+
+
e-
+

## Slide 58: Where do photons go?
beamstop
elastic scattering (6%)
Transmitted (98%)
inelastic scattering (7%)
Protein
1A x-rays
Re-emitted (99%)
Absorbed (~0%)

## Slide 59: Where do photons go?
beamstop
elastic scattering (6%)
Transmitted (98%)
inelastic scattering (7%)
Photoelectric (87%)
Protein
1A x-rays
Re-emitted (99%)
Absorbed (~0%)

## Slide 60: Where do photons go?
beamstop
elastic scattering (6%)
Transmitted (98%)
inelastic scattering (7%)
Photoelectric (87%)
Protein
1A x-rays
Re-emitted (~0%)
Absorbed (99%)
Re-emitted (99%)
Absorbed (~0%)

## Slide 61: Where do photons go?
beamstop
elastic scattering (6%)
Transmitted (98%)
useful/absorbed energy: 7.3%
inelastic scattering (7%)
Photoelectric (87%)
Protein
1A x-rays
Re-emitted (~0%)
Absorbed (99%)
Re-emitted (99%)
Absorbed (~0%)

## Slide 62
sample
detector
x-ray beam
scattering

## Slide 63
sample
detector
x-ray beam
scattering

## Slide 64
sample
detector
x-ray beam
scattering

## Slide 65
sample
detector
x-ray beam
scattering

## Slide 66
to source
to detector
d
d∙sin(θ)
θ
atom #1
atom #2
Bragg’s Law
nλ = 2d sin(θ)

## Slide 67
to source
to detector
d
d∙sin(θ)
θ
atom #1
atom #2
Bragg’s Law
nλ = 2d sin(θ)

## Slide 68
sample
x-ray beam
anomalous scattering
detector

## Slide 69
sample
detector
x-ray beam
anomalous scattering

## Slide 70
sample
x-ray beam
anomalous scattering
detector

## Slide 71
sample
detector
x-ray beam
anomalous scattering

## Slide 72
sample
detector
x-ray beam
anomalous scattering

## Slide 73
detector
sample
detector
x-ray beam
scattering

## Slide 76
scattering from a structure
sample
detector

## Slide 77
forward Fourier Transform
inverse Fourier Transform
no phase

## Slide 78
scattering from a structure
colored by phase
sample
detector
inverse Fourier Transform

## Slide 79
scattering from a lattice
colored by phase
sample
detector

## Slide 80
scattering from a lattice
colored by phase
sample
detector

## Slide 81
scattering from a lattice
colored by phase
sample
detector

## Slide 82
scattering from a crystal structure
colored by phase
sample
detector

## Slide 83: Resolution
http://bl831.als.lbl.gov/~jamesh/movies/resolution.mpeg

## Slide 84: Resolution: low-angle cutoff
http://bl831.als.lbl.gov/~jamesh/movies/lo_cut.mpeg

## Slide 85: R-factor
http://bl831.als.lbl.gov/~jamesh/movies/rfactor.mpeg

## Slide 86: Figure of Merit
http://bl831.als.lbl.gov/~jamesh/movies/dephase.mpeg

## Slide 87: Overloads
http://bl831.als.lbl.gov/~jamesh/movies/overloads.mpeg

## Slide 88: Completeness: missing wedge
http://bl831.als.lbl.gov/~jamesh/movies/osc.mpeg

## Slide 89: Completeness: random deletion
http://bl831.als.lbl.gov/~jamesh/movies/completeness.mpeg

## Slide 90: At the beamline…
Resolution
  problem: background
  solution: use as few pixels as possible
Phases
  problem: fractional errors
  solution: use as many pixels as possible

## Slide 91: beam size vs xtal size
  Put your crystal into the beam

## Slide 92: Get thee to a microbeam?
Evans et al. (2011)."Macromolecular microcrystallography", Crystallography Reviews 17, 105-142.

## Slide 93: beam size vs xtal size
  Put your crystal into the beam
  Shoot the whole crystal

## Slide 94: shoot the whole crystal

## Slide 95: shoot the whole crystal

## Slide 96: shoot the whole crystal

## Slide 97: beam size vs xtal size
  Put your crystal into the beam
  Shoot the whole crystal
  Shoot nothing but the crystal

## Slide 98: shoot nothing but the crystal

## Slide 99: shoot nothing but the crystal

## Slide 100: X-ray scattering “rules”:
1 μm crystal ≈ 1 μm water
≈ 1 μm plastic
≈ 0.1 μm glass
≈ 1000 μm air

## Slide 101
X-ray beam size (μm)
Resolution (Å)
too big
too small
just right

## Slide 102: beam size vs xtal size
  Put your crystal into the beam
  Shoot the whole crystal
  Shoot nothing but the crystal
  Back off!

## Slide 103: Background scattering

## Slide 104: Background scattering
$100,000.00
$100,000.00
$100,000.00
$100,000.00
$100,000.00
$100,000.00
$100,000.00
$100,000.00
real estate is
expensive
use it!

## Slide 105: The truth about x-ray beams
Term units significance
Flux photons/s duration of experiment
Beam Size μm match to crystal
Divergence mrad spot size vs distance
Wavelength Å resolution and absorption
Dispersion Δλ/λ spot size
Flux density ph/s/area scattering/damage rate
Fluence ph/area scattering/damage

## Slide 106: The truth about x-ray beams
Term units significance
Flux photons/s duration of experiment
Beam Size μm match to crystal
Divergence mrad spot size vs distance
Wavelength Å resolution and absorption
Dispersion Δλ/λ spot size
Flux density ph/s/area scattering/damage rate
Fluence ph/area scattering/damage

## Slide 107: The truth about x-ray beams
quantity | units | home source | APS 22-ID
flux | Photons/second | 1.7 x 109 | 7 x 1012
exposure | seconds | 20 | 1

## Slide 108: The truth about x-ray beams
quantity | units | home source | APS 22-ID
flux | Photons/second | 1.7 x 109 | 7 x 1012
exposure | seconds | 20 | 1
Dispersion | wavelength range / wavelength | 0.2% (Kα1 - Kα2) | 0.014% (Si111)
Divergence | milliRadian | 10 | 0.1
Beam size | microns | 70 | 40 (h) 80 (v)
Spectral brightness | Photons/s/mm2/ mR2/0.1%BW | 1.7 x 1010 | 1.5 x 1019

## Slide 109: The truth about x-ray beams
quantity | units | home source | APS 22-ID
flux | Photons/second | 1.2 x 108 | 7 x 1012
exposure | seconds | 286 | 1
Dispersion | wavelength range / wavelength | 0.014% (Si111) | 0.014% (Si111)
Divergence | milliRadian | 10 | 0.1
Beam size | microns | 70 | 40 (h) 80 (v)
Spectral brightness | Photons/s/mm2/ mR2/0.1%BW | 1.7 x 1010 | 1.5 x 1019

## Slide 110: The truth about x-ray beams
quantity | units | home source | APS 22-ID
flux | Photons/second | 1.2 x 104 | 7 x 1012
exposure | time | 33 days | 1 second
Dispersion | wavelength range / wavelength | 0.014% (Si111) | 0.014% (Si111)
Divergence | milliRadian | 0.1 | 0.1
Beam size | microns | 70 | 40 (h) 80 (v)
Spectral brightness | Photons/s/mm2/ mR2/0.1%BW | 1.7 x 1010 | 1.5 x 1019

## Slide 111: The truth about x-ray beams
quantity | units | home source | APS 22-ID
flux | Photons/second | 7.6 x 103 | 7 x 1012
exposure | time | 52 days | 1 second
Dispersion | wavelength range / wavelength | 0.014% (Si111) | 0.014% (Si111)
Divergence | milliRadian | 0.1 | 0.1
Beam size | microns | 56 | 40 (h) 80 (v)
Spectral brightness | Photons/s/mm2/ mR2/0.1%BW | 1.7 x 1010 | 1.5 x 1019

## Slide 112: Fine Slicing
Pflugrath, J. W. (1999)."The finer things in X-ray diffraction data collection", Acta Cryst. D 55, 1718-1725.
background
background

## Slide 113: Classes of error in MX
Dependence on signal
Time
none sqrt proportional
none
1/sqrt
1/prop.
CCD
Read-out
Photon
counting
Beam flicker
Shutter jitter
Sample vibration
Detector calibration
attenuation
partiality
Non-isomorphism
Radiation damage

## Slide 114: Optimal exposure time (faint spots)
thr Optimal exposure time for data set (s)
tref exposure time of reference image (s)
bgref background level near weak spots on
reference image (ADU)
bg0 ADC offset of detector (ADU)
bghr optimal background level (via thr)
σ0 rms read-out noise (ADU)
gain ADU/photon
m multiplicity of data set (including partials)
adjust exposure
so this is ~15

## Slide 115: beam size vs xtal size
  Put your crystal into the beam
  Shoot the whole crystal
  Shoot nothing but the crystal
  Back off!
  The crystal must rotate

## Slide 116: Get thee to a microbeam?
Evans et al. (2011)."Macromolecular microcrystallography", Crystallography Reviews 17, 105-142.

## Slide 117: At the beamline…
Resolution
  problem: background
  solution: use as few pixels as possible
Phases
  problem: fractional errors
  solution: use as many pixels as possible

## Slide 118
anomalous signal
Crick, F. H. C. & Magdoff, B. S. (1956) Acta Crystallogr. 9, 901-908.
Hendrickson, W. A. & Teeter, M. M. (1981) Nature 290, 107-113.
# sites
MW (Da)
ΔF
F
≈ 1.2 f”
√
World record!
ΔF/F = 0.5%
Wang, Dauter & Dauter (2006) Acta Cryst. D 62, 1475-1483.

## Slide 119: Fractional error
mult > (—)2
~3%
<ΔF/F>

## Slide 120: Can you count to 1,000,000 ?
= 0.1%
sqrt(1,000,000)
1,000,000
= 3%
sqrt(1,000)
1,000
> 1000 is a waste!
photon
spot
Theoretically:
In reality:
ISa ~ 33
Rmeas ≈ 0.1% ?
ISa = 1000
Rmeas = ≈ 3%

## Slide 121
Required signal-to-noise (I/σ)
Solve-able proteins (%)
Current
technology
Goal
Threshold of a revolution in phasing

## Slide 122
Source of error | realistic simulation | No SHSSS | Perfect detector
Photon counting | + | + | +
Shutter jitter | + | + | -
Beam flicker | + | + | -
Sample absorption | + | + | -
Radiation damage | + | + | -
Imperfect spindle | + | + | -
vignette | + | + | -
Corner correction | + | + | -
SHSSS | + | - | -
Rmeas (∞-10 Å) | 2.8% | 0.7% | 0.7%
I/σ asymptotic | 26.8 | 74.2 | 81.0
Threshold of a revolution in phasing
Holton et al (2014) "R-factor gap", FEBS Journal 281, 4046-4060.

## Slide 123
Spot centroid position (pixels)
Relative spot intensity
Spatial Heterogeneity in Sharp Spot Sensitivity

## Slide 124
Spot centroid position (pixels)
Relative spot intensity
Spatial Heterogeneity in Sharp Spot Sensitivity

## Slide 125: SHSSS! The dominant source of error for anomalous difference measurements
Spatial
Heterogeneity in
Sharp
Spot
Sensitivity
Flat
Field

## Slide 126: SHSSS! The dominant source of error for anomalous difference measurements
Spatial
Heterogeneity in
Sharp
Spot
Sensitivity
Flat
Field
90%
110%
x 1.11 =
x 0.91 =
100%
100%

## Slide 127: SHSSS! The dominant source of error for anomalous difference measurements
Spatial
Heterogeneity in
Sharp
Spot
Sensitivity
Sharp
Spot
20%
110%
x 1.11 =
x 0.91 =
100%
22%

## Slide 128: SHSSS! The dominant source of error for anomalous difference measurements
Spatial
Heterogeneity in
Sharp
Spot
Sensitivity
Sharp
Spot
90%
30%
x 1.11 =
x 0.91 =
27.3%
100%
Integral: 127.3

## Slide 129: SHSSS! The dominant source of error for anomalous difference measurements
Spatial
Heterogeneity in
Sharp
Spot
Sensitivity
Sharp
Spot
20%
110%
x 1.11 =
x 0.91 =
100%
22.2%
Integral: 122.2

## Slide 130: SHSSS! The dominant source of error for anomalous difference measurements
Spatial
Heterogeneity in
Sharp
Spot
Sensitivity
Sharp
Spot
90%
30%
x 1.11 =
x 0.91 =
27.3%
100%
Integral: 127.3

## Slide 131: Detector calibration: 7235 eV

## Slide 132: Detector calibration: 7247 eV

## Slide 133: Gadox calibration vs energy
photon energy (keV)
Relative absorption depth
same = good!
bad!

## Slide 134
Pick-up tool mark
Bragg
glitch
oxygen
inclusions
“tree rings”
1% high
average
1% low
~3x105 photon/pixel
Pilatus: subtract smooth baseline

## Slide 135
N
N
N
θ
λ
detection
event
incoming
photon
PAD: need more than flood-field

## Slide 136
thickness
width
N
N
N
Bragg
glitch
oxygen
inclusion
θ
λ
detection
event
incoming
photon
PAD: need more than flood-field

## Slide 137: Pilatus: 1-module shift

## Slide 138
Pilatus: 1-module shift

## Slide 139
Pilatus: 1-module shift

## Slide 140
Pilatus: 1-module shift - aligned

## Slide 141
3% high
average
3% low
Pilatus: 1-module shift - ratio

## Slide 142
3.5
3.0
2.5
2.0
1.5
1.0
0.5
SHSSS
Systematic component of Rsseparate (%)
distance between spots (mm)
0.1 1 10 100
CCD detector
anomalous mates are always on
different modules
different
modules
Pilatus
SN113
Pilatus
SN001

## Slide 143: Fractional error
mult > (—)2
~3%
<ΔF/F>

## Slide 144: Systematic error does not “average out”

## Slide 145
Holton & Frankel (2010) Acta D 66 393-408.

## Slide 146: Dose slicing
crystal’s useful life
N
photons
N
photons
N
photons
unacceptable
damage
unacceptable
read noise

## Slide 147: Data collection parameters:
16 crystals
360° each, inverse beam
7235 eV photon energy
< 1 MGy per xtal
Australian Synchrotron MX1
  35 kGy/s into 100 μm x 100 μm

## Slide 148
RESOLUTION COMPLETENESS R-FACTOR I/SIGMA R-meas CC(1/2) Anomal SigAno Nano
LIMIT OF DATA observed Corr
9.17 99.1% 3.9% 257.47 3.9% 100.0* 91* 5.024 450
6.49 100.0% 5.2% 214.33 5.2% 100.0* 86* 3.836 882
5.30 100.0% 7.2% 165.13 7.2% 100.0* 76* 3.257 1175
4.59 100.0% 7.2% 175.42 7.3% 100.0* 67* 2.589 1403
4.10 99.9% 7.7% 174.13 7.7% 100.0* 59* 2.264 1594
3.74 99.9% 9.4% 143.09 9.4% 100.0* 49* 1.953 1783
3.47 100.0% 11.2% 120.17 11.2% 100.0* 39* 1.696 1942
3.24 100.0% 14.1% 91.14 14.1% 100.0* 30* 1.333 2103
3.06 99.9% 19.5% 65.79 19.5% 100.0* 23* 1.117 2214
2.90 99.9% 29.0% 44.85 29.1% 99.9* 17* 1.008 2369
2.77 99.9% 40.5% 32.58 40.6% 99.8* 11* 0.901 2493
2.65 99.9% 52.8% 25.16 52.9% 99.8* 10* 0.866 2605
2.54 100.0% 67.4% 19.47 67.6% 99.6* 2 0.804 2705
2.45 100.0% 88.9% 14.58 89.2% 99.2* 4 0.831 2859
2.37 100.0% 109.3% 9.97 109.7% 98.1* 5 0.829 2925
2.29 100.0% 138.2% 6.87 138.9% 96.1* 1 0.760 3037
2.22 100.0% 197.1% 4.03 198.6% 83.5* -1 0.721 3159
2.16 100.0% 227.3% 2.41 230.8% 46.9* -1 0.677 3224
2.10 61.2% 154.4% 1.28 163.6% 47.0* -2 0.660 1999
2.05 47.9% 170.1% 0.68 196.5% 25.7* 3 0.629 1578
total 93.3% 15.7% 54.30 15.8% 100.0* 12* 1.217 42499
140-fold multiplicity: 16 crystals, 360° each, inverse beam, 7235 eV

## Slide 149: 140-fold multiplicity
18 σ
Phased anomalous difference Fourier
16 σ

## Slide 150: 140-fold multiplicity
15σ = PO4
Phased anomalous difference Fourier

## Slide 151: 140-fold multiplicity
~2σ = Mg?
Phased anomalous difference Fourier

## Slide 152: 140-fold multiplicity
8.2σ = Mg?
DELFAN residual anomalous difference

## Slide 153: Discerning Na+ from Mg++
f’’ (electrons)
DELFAN peak height (σ)
Mg
Ne
Na
F
O
N

## Slide 154: 140-fold multiplicity
7.4σ = Na
DELFAN residual anomalous difference

## Slide 155: Suggested anomalous protocol:
2 wavelengths are better than 1
- (peak + inf)/2, and remote
MAD, not M-SAD!
360° in < 5 MGy
move detector
4X exposure
goto 2

## Slide 156: Summary
100 ADU/pixel
10 μm for lysozyme
~3% error per spot, 1%/MGy
7235 eV for S-SAD
“Attenu-wait”: dose slicing
http://bl831.als.lbl.gov/xtalsize.html
http://bl831.als.lbl.gov/~jamesh/mlfsom/
http://bl831.als.lbl.gov/~jamesh/powerpoint/ACMC_GBU_2013.ppt

## Slide 157: Summary
15 ADU/pixel
10 μm for lysozyme
~3% error per spot, 1%/MGy
7235 eV for S-SAD
“Attenu-wait”: dose slicing
http://bl831.als.lbl.gov/xtalsize.html
http://bl831.als.lbl.gov/xtallife.html
http://bl831.als.lbl.gov/~jamesh/mlfsom/
http://bl831.als.lbl.gov/~jamesh/powerpoint/CSHL_tipsNtricks_2015.pptx

## Slide 158: Why harvest at all?
Radiation damage
  10-100x reduced at ~100 K
Background scatter
    tray >> xtal
Timing
    nothing lasts forever
Shipping
    cryo samples are glass beads

## Slide 159: Why harvest at all?
Radiation damage
  10-100x reduced at ~100 K
Background scatter
    tray >> xtal
Timing
    nothing lasts forever
Shipping
    cryo samples are glass beads

## Slide 160: X-ray transmission of materials
Depth into material (μm)
transmission
glass
protein
crystal
water
air

## Slide 161: X-ray scattering of materials
Differential cross section
(scattered photons/solid angle/incident photon/meter thickness)
 7.5 3.8 2.5 1.9 1.5
d-spacing (Ǻ)
water
25% glycerol
paratone-N oil
plastic
glass

## Slide 162: naked xtal

## Slide 163: Crowfoot Cell

## Slide 164: Crowfoot Cell

## Slide 165: Crowfoot Cell
J. D. Bernal and D. Crowfoot, Nature (London) 133, 794 (1934).

## Slide 166: Fluidigm RT/cryo mount
“Diffraction Capable” chip
http://www.fluidigm.com/products/topaz-chips.html

## Slide 167: Plate goniometers
Notes: The non-rotating “plate goniometer” (left) was developed last year. A new incarnation for taking rotation-camera data is now being developed (right). Either of these devices will accept most any crystallization tray that can be rotated 90 degrees to gravity and probe arbitrary points in each crystallization drop with a 50 um x-ray beam. Installing a plate goniometer takes ~30 minutes and the location ~300 mm downbeam from the “normal” protein crystallography sample mounting point makes this technology transferable to most any beamline.
Current status is that a 95% complete data set has been collected from a 50 um lysozyme crystal at room temperature, and trials are now underway to collect data sets from “real” crystals.

## Slide 168
Notes: non-rotating tray goniometer stage installed in ALS 8.3.1. Note that the “normal” cryo-pin goniometer on the left is still in place and a brass flight tube has been added to contain the air scatter from the main beam on its way to the tray goniometer.

## Slide 170: Why harvest at all?
Radiation damage
  10-100x reduced at ~100 K
Background scatter
    tray >> xtal
Timing
    nothing lasts forever
Shipping
    cryo samples are glass beads

## Slide 172
8000 Gy/s

## Slide 173: Example Data
lysozyme
50 μm beam
37 Gy/s (0.775 Å)
30s exposures at 20C
90° of data, 97% complete
I/ = 1.5 at 1.9 Å
Rmerge 18% (overall) 5% (low)
ΔB same as 20 min at 100K
B factor
-2
-4
-6
scale vs batch
2
1
0

## Slide 174: Room temperature damage rates
dose rate (kGy/s)
dose at half intensity (MGy)

## Slide 175: Room temperature damage rates
dose rate (kGy/s)
dose at half intensity (MGy)

## Slide 176: Every spot is an unpaired partial!
Ispot = k Ncells2
Ewald sphere
range

## Slide 177: Every spot is an unpaired partial!
Ewald sphere
range
Ispot = k Ncells1
Woolfson, M. M. (1997). An introduction to X-ray crystallography. Ch. 2 & 6

## Slide 178: Internal consistency of data
1,850,000 images
112,000 > 10 peaks
33,000 indexed
16,500 good preds
rigid-body:
Rcryst = 0.252
Rfree = 0.232

## Slide 179: what is “partiality”?
F(h,k,l)
Ewald sphere
λ1*
λ2*
F(0,0,0)
F(h,k,l)

## Slide 180: Why harvest at all?
Radiation damage
  10-100x reduced at ~100 K
Background scatter
    tray >> xtal
Timing
    nothing lasts forever
Shipping
    cryo samples are glass beads

## Slide 181: Why harvest at all?
Radiation damage
  10-100x reduced at ~100 K
Background scatter
    tray >> xtal
Timing
    nothing lasts forever
Shipping
    cryo samples are glass beads

## Slide 182: MiTeGen RT cell
http://www.mitegen.com/products/micrort/micrort.shtml

## Slide 183: Free-mounting system
http://www.proteros.de/

## Slide 184: Summary
assay, control and open mind
purity, purity, purity
oil is your friend
Warkentin method
try room temp!
http://bl831.als.lbl.gov/~jamesh/powerpoint/
CSHL_tipsNtricks_2015.pptx

## Slide 185: avoiding overlaps
c
c

## Slide 186: avoiding overlaps
c

## Slide 187: At the beamline…
Resolution
  problem: background
  solution: use as few pixels as possible
Phases
  problem: fractional errors
  solution: use as many pixels as possible

## Slide 188: Major Phasing techniques
Molecular Replacement
Multiple Isomorphous Replacement
Multiwavelength Anomalous Diffraction
Single-wavelength Anomalous Diffraction

## Slide 189
inverse Fourier Transform
no phase

## Slide 190
inverse Fourier Transform
no phase

## Slide 191: Major Phasing techniques
Molecular Replacement
Multiple Isomorphous Replacement
Multiwavelength Anomalous Diffraction
Single-wavelength Anomalous Diffraction

## Slide 192: “R” factors
R = % error

## Slide 193: “R” factors
completely random: 0.59
starting MR solution: 0.4-0.55
something still wrong?: > 0.3
correct chain trace: < 0.2
small molecule: ~ 0.05

## Slide 194: “R” factors
Rcryst (or just “R”)
observed vs calculated data (Fs)
Rfree
cross-check with “random” subset of data
should be < 0.3 and < Rcryst + 0.1

## Slide 195
structure factor (F)
spot index (h)
Fitting data

## Slide 196
structure factor (F)
spot index (h)
Fitting data

## Slide 197
structure factor (F)
spot index (h)
Fitting data

## Slide 198
structure factor (F)
spot index (h)
Fitting data

## Slide 199: “R” factors
Rcryst (or just “R”)
observed vs calculated data (Fs)
Rfree
cross-check with “random” subset of data
should be < 0.3 and < Rcryst + 0.1
Rsym = Rmerge (self-consistency of data: Is)

## Slide 200: Rmerge
completely random: 0.59
weak data (high angle): 0.7- ∞
wrong symmetry choice?: ~0.2-0.55
small or disordered crystal: ~0.1-0.2
typical: ~ 0.05
blows up
as Iobs → 0

## Slide 201: “R” factors
Rcryst (or just “R”)
observed vs calculated data (Fs)
Rfree
cross-check with “random” subset of data
should be < 0.3 and < Rcryst + 0.1
Rsym = Rmerge (self-consistency of data: Is)
Rrim Rpim Rmeas
Riso Ranom
Rdiff
Rsym is “unfair” to high multiplicity

## Slide 202: Is the structure any good?
Molecular replacement:
  R/Rfree (compare with Rmerge)
MAD/SAD/MIR
  R and FOM
All cases
  does the map look like the model?!

## Slide 203: “B” factors
B ≈ 4d2 + 12
essentially, the “resolution” of an atom
d = resolution in Å

## Slide 204: Is the structure any good?
“do the data presented justify the conclusions drawn?”

## Slide 205: Detector calibration: 7247 eV
target:
oil
distance:
900 mm
2θ: 12°

## Slide 206: Detector calibration: 7235 eV
target:
oil
distance:
900 mm
2θ: 12°

## Slide 207: Detector calibration: ALS 8.3.1

## Slide 208: Detector calibration errors: detector 2

## Slide 209: Detector calibration errors: detector 3