## Slide 1: PowerPoint File available:
http://bl831.als.lbl.gov/
~jamesh/powerpoint/
CSHL_SvN_2015.pptx

## Slide 2: Acknowledgements
UCSF LBNL SLAC
ALS 8.3.1 creator: Tom Alber
Center for Structure of Membrane Proteins (CSMP)
Membrane Protein Expression Center II (MPEC)
Center for HIV Accessory and Regulatory Complexes (HARC)
UC Multicampus Research Programs and Initiatives (MRPI)
UCSF Program for Breakthrough Biomedical Research (PBBR) Integrated Diffraction Analysis Technologies (IDAT)
Plexxikon, Inc.
M D Anderson CRC
Synchrotron Radiation Structural Biology Resource (SLAC)
The Advanced Light Source is supported by the Director, Office of Science, Office of Basic Energy Sciences, Materials Sciences Division, of the US Department of Energy under contract No. DE-AC02-05CH11231 at Lawrence Berkeley National Laboratory.
Robert Stroud James Fraser John Spence
Chris Nielsen Clemens Schulze-Briese Aina Cohen Ana Gonzalez

## Slide 3: The “success rate” of structure determination
100 s/dataset
200 days/year
~150 beamlines
~26,000,000 datasets/year
9640 PDBs in 2014

## Slide 4: signal vs noise
“If you don’t have good data, then you have no data at all.”
-Sung-Hou Kim

## Slide 5: signal vs noise
easy
hard
impossible
threshold of “solvability”

## Slide 6: signal vs noise
“If you don’t have good data, then you must learn statistics.”
-James Holton

## Slide 7: Adding noise
12 + 12 = 1.42
32 + 12 = 3.22
102 + 12 = 10.052

## Slide 8: 100% SeMet incorporation
Trivial to solve

## Slide 9: 100% S -Met incorporation
Impossible to solve

## Slide 10: Phase Problem: Se vs S
fraction Sulfur
correlation coefficient

## Slide 11
21% Se 79% S

## Slide 12
11% Se 89% S

## Slide 13: Finding sites
fraction Sulfur
Number of correct sites
25
20
15
10
5
0
peak height ()

## Slide 14: Finding sites
fraction Sulfur
Number of correct sites

## Slide 15: MR simulation
Rmsd from perfect search model (Å)
corrupted model
Correlation coefficient to correct density

## Slide 16: The transitions are sharp!
How can we predict success/failure?
Know Thy Experiment

## Slide 17: Elastic scattering

## Slide 18: Inelastic scattering

## Slide 19: Photoelectric absorption
e-
+

## Slide 20: Fluorescence
e-
+

## Slide 21: Metal identification

## Slide 22: Auger emission
++
e-

## Slide 23: Ionization track
e-
e-
e-
e-
e-
e-
e-
e-
e-
+
+
+
+
+
+
+
+
+
e-
+

## Slide 25: Where do photons go?
beamstop
elastic scattering (6%)
Transmitted (98%)
useful/absorbed energy: 7.3%
inelastic scattering (7%)
Photoelectric (87%)
Protein
1A x-rays
Re-emitted (~0%)
Absorbed (99%)
Re-emitted (99%)
Absorbed (~0%)

## Slide 26
sample
detector
x-ray beam
scattering

## Slide 28
scattering from a lattice
colored by phase
sample
detector

## Slide 29
scattering from a crystal structure
colored by phase
sample
detector

## Slide 30: Resolution
http://bl831.als.lbl.gov/~jamesh/movies/resolution.mpeg

## Slide 31: Resolution: low-angle cutoff
http://bl831.als.lbl.gov/~jamesh/movies/lo_cut.mpeg

## Slide 32: R-factor
http://bl831.als.lbl.gov/~jamesh/movies/rfactor.mpeg

## Slide 33: Figure of Merit
http://bl831.als.lbl.gov/~jamesh/movies/dephase.mpeg

## Slide 34: Overloads
http://bl831.als.lbl.gov/~jamesh/movies/overloads.mpeg

## Slide 35: Completeness: missing wedge
http://bl831.als.lbl.gov/~jamesh/movies/osc.mpeg

## Slide 36: Completeness: random deletion
http://bl831.als.lbl.gov/~jamesh/movies/completeness.mpeg

## Slide 37: At the beamline…
Resolution
  problem: background
  solution: use as few pixels as possible
Phases
  problem: fractional errors
  solution: use as many pixels as possible

## Slide 38: shoot the whole crystal

## Slide 39: shoot the whole crystal

## Slide 40: shoot the whole crystal

## Slide 41: shoot nothing but the crystal

## Slide 42: shoot nothing but the crystal

## Slide 43: X-ray scattering “rules”:
1 μm crystal ≈ 1 μm water
≈ 1 μm plastic
≈ 0.1 μm glass
≈ 1000 μm air

## Slide 44
X-ray beam size (μm)
Resolution (Å)
too big
too small
just right

## Slide 45: Background scattering

## Slide 46: Background scattering
$100,000.00
$100,000.00
$100,000.00
$100,000.00
$100,000.00
$100,000.00
$100,000.00
$100,000.00
real estate is
expensive
use it!

## Slide 47: The truth about x-ray beams
Term units significance
Flux photons/s duration of experiment
Beam Size μm match to crystal
Divergence mrad spot size vs distance
Wavelength Å resolution and absorption
Dispersion Δλ/λ spot size
Flux density ph/s/area scattering/damage rate
Fluence ph/area scattering/damage

## Slide 48: The truth about x-ray beams
Term units significance
Flux photons/s duration of experiment
Beam Size μm match to crystal
Divergence mrad spot size vs distance
Wavelength Å resolution and absorption
Dispersion Δλ/λ spot size
Flux density ph/s/area scattering/damage rate
Fluence ph/area scattering/damage

## Slide 49: The truth about x-ray beams
quantity | units | home source | APS 22-ID
flux | Photons/second | 7.6 x 103 | 7 x 1012
exposure | time | 52 days | 1 second
Dispersion | wavelength range / wavelength | 0.014% (Si111) | 0.014% (Si111)
Divergence | milliRadian | 0.1 | 0.1
Beam size | microns | 56 | 40 (h) 80 (v)
Spectral brightness | Photons/s/mm2/ mR2/0.1%BW | 1.7 x 1010 | 1.5 x 1019

## Slide 50: Fine Slicing
Pflugrath, J. W. (1999)."The finer things in X-ray diffraction data collection", Acta Cryst. D 55, 1718-1725.
background
background

## Slide 51: Classes of error in MX
Dependence on signal
Time
none sqrt proportional
none
1/sqrt
1/prop.
CCD
Read-out
Photon
counting
Beam flicker
Shutter jitter
Sample vibration
Detector calibration
attenuation
partiality
Non-isomorphism
Radiation damage

## Slide 52: Optimal exposure time (faint spots)
thr Optimal exposure time for data set (s)
tref exposure time of reference image (s)
bgref background level near weak spots on
reference image (ADU)
bg0 ADC offset of detector (ADU)
bghr optimal background level (via thr)
σ0 rms read-out noise (ADU)
gain ADU/photon
m multiplicity of data set (including partials)
adjust exposure
so this is ~15

## Slide 53: beam size vs xtal size
  Put your crystal into the beam
  Shoot the whole crystal
  Shoot nothing but the crystal
  Back off!
  The crystal must rotate

## Slide 54: Get thee to a microbeam?
Evans et al. (2011)."Macromolecular microcrystallography", Crystallography Reviews 17, 105-142.

## Slide 55
anomalous signal
Crick, F. H. C. & Magdoff, B. S. (1956) Acta Crystallogr. 9, 901-908.
Hendrickson, W. A. & Teeter, M. M. (1981) Nature 290, 107-113.
# sites
MW (Da)
ΔF
F
≈ 1.2 f”
√
World record!
ΔF/F = 0.5%
Wang, Dauter & Dauter (2006) Acta Cryst. D 62, 1475-1483.

## Slide 56: Fractional error
mult > (—)2
~3%
<ΔF/F>

## Slide 57: Can you count to 1,000,000 ?
= 0.1%
sqrt(1,000,000)
1,000,000
= 3%
sqrt(1,000)
1,000
> 1000 is a waste!
photon
spot
Theoretically:
In reality:
ISa ~ 33
Rmeas ≈ 0.1% ?
ISa = 1000
Rmeas = ≈ 3%

## Slide 58
Required signal-to-noise (I/σ)
Solve-able proteins (%)
Current
technology
Goal
Threshold of a revolution in phasing

## Slide 59
Source of error | realistic simulation | No SHSSS | Perfect detector
Photon counting | + | + | +
Shutter jitter | + | + | -
Beam flicker | + | + | -
Sample absorption | + | + | -
Radiation damage | + | + | -
Imperfect spindle | + | + | -
vignette | + | + | -
Corner correction | + | + | -
SHSSS | + | - | -
Rmeas (∞-10 Å) | 2.8% | 0.7% | 0.7%
I/σ asymptotic | 26.8 | 74.2 | 81.0
Threshold of a revolution in phasing
Holton et al (2014) "R-factor gap", FEBS Journal 281, 4046-4060.

## Slide 60
Spot centroid position (pixels)
Relative spot intensity
Spatial Heterogeneity in Sharp Spot Sensitivity

## Slide 61: SHSSS! The dominant source of error for anomalous difference measurements
Spatial
Heterogeneity in
Sharp
Spot
Sensitivity
Flat
Field

## Slide 62: Gadox calibration vs energy
photon energy (keV)
Relative absorption depth
same = good!
bad!

## Slide 63
Pick-up tool mark
Bragg
glitch
oxygen
inclusions
“tree rings”
1% high
average
1% low
~3x105 photon/pixel
Pilatus: subtract smooth baseline

## Slide 64
thickness
width
N
N
N
Bragg
glitch
oxygen
inclusion
θ
λ
detection
event
incoming
photon
PAD: need more than flood-field

## Slide 65
3.5
3.0
2.5
2.0
1.5
1.0
0.5
SHSSS
Systematic component of Rsseparate (%)
distance between spots (mm)
0.1 1 10 100
CCD detector
anomalous mates are always on
different modules
different
modules
Pilatus
SN113
Pilatus
SN001

## Slide 66
Holton & Frankel (2010) Acta D 66 393-408.

## Slide 67: Dose slicing
crystal’s useful life
N
photons
N
photons
N
photons
unacceptable
damage
unacceptable
read noise

## Slide 68
RESOLUTION COMPLETENESS R-FACTOR I/SIGMA R-meas CC(1/2) Anomal SigAno Nano
LIMIT OF DATA observed Corr
9.17 99.1% 3.9% 257.47 3.9% 100.0* 91* 5.024 450
6.49 100.0% 5.2% 214.33 5.2% 100.0* 86* 3.836 882
5.30 100.0% 7.2% 165.13 7.2% 100.0* 76* 3.257 1175
4.59 100.0% 7.2% 175.42 7.3% 100.0* 67* 2.589 1403
4.10 99.9% 7.7% 174.13 7.7% 100.0* 59* 2.264 1594
3.74 99.9% 9.4% 143.09 9.4% 100.0* 49* 1.953 1783
3.47 100.0% 11.2% 120.17 11.2% 100.0* 39* 1.696 1942
3.24 100.0% 14.1% 91.14 14.1% 100.0* 30* 1.333 2103
3.06 99.9% 19.5% 65.79 19.5% 100.0* 23* 1.117 2214
2.90 99.9% 29.0% 44.85 29.1% 99.9* 17* 1.008 2369
2.77 99.9% 40.5% 32.58 40.6% 99.8* 11* 0.901 2493
2.65 99.9% 52.8% 25.16 52.9% 99.8* 10* 0.866 2605
2.54 100.0% 67.4% 19.47 67.6% 99.6* 2 0.804 2705
2.45 100.0% 88.9% 14.58 89.2% 99.2* 4 0.831 2859
2.37 100.0% 109.3% 9.97 109.7% 98.1* 5 0.829 2925
2.29 100.0% 138.2% 6.87 138.9% 96.1* 1 0.760 3037
2.22 100.0% 197.1% 4.03 198.6% 83.5* -1 0.721 3159
2.16 100.0% 227.3% 2.41 230.8% 46.9* -1 0.677 3224
2.10 61.2% 154.4% 1.28 163.6% 47.0* -2 0.660 1999
2.05 47.9% 170.1% 0.68 196.5% 25.7* 3 0.629 1578
total 93.3% 15.7% 54.30 15.8% 100.0* 12* 1.217 42499
140-fold multiplicity: 16 crystals, 360° each, inverse beam, 7235 eV

## Slide 69: 140-fold multiplicity
18 σ
Phased anomalous difference Fourier
16 σ

## Slide 70: 140-fold multiplicity
7.4σ = Na
DELFAN residual anomalous difference

## Slide 71: Suggested anomalous protocol:
2 wavelengths are better than 1
- (peak + inf)/2, and remote
MAD, not M-SAD!
360° in < 5 MGy
move detector
4X exposure
goto 2

## Slide 72: Summary
15 ADU/pixel
10 μm for lysozyme
~3% error per spot, 1%/MGy
7235 eV for S-SAD
“Attenu-wait”: dose slicing
http://bl831.als.lbl.gov/xtalsize.html
http://bl831.als.lbl.gov/xtallife.html
http://bl831.als.lbl.gov/~jamesh/mlfsom/
http://bl831.als.lbl.gov/~jamesh/powerpoint/CSHL_tipsNtricks_2015.pptx