## Slide 1: And the dominant source of error is…
Resolution
  Disorder (including Radiation Damage)
  Background
Phases
  Non-isomorphism (including Radiation Damage)
  Vibration (beam, xtal, spindle, etc.)
  Detector calibration

## Slide 2: One xtal, no idea? Suggested anomalous protocol:
2 wavelengths are better than 1
- (peak + inf)/2, and remote
MAD, not M-SAD!
Wedge size: < 10% overhead
360° in < 5 MGy
move detector
4X exposure
goto 2
Se-Met

## Slide 3
Bijvoet And Dispersive Anomalous Scattering
300 eV
higher
Halfway between
Inflection + peak

## Slide 4: Riso vs dose
Riso (%)
change in dose (MGy)
data taken from Banumathi, et al. (2004) Acta Cryst. D 60, 1085-1093.

## Slide 5: Riso vs dose
Riso (%)
change in dose (MGy)
data taken from Banumathi, et al. (2004) Acta Cryst. D 60, 1085-1093.
Riso ≈ 0.7 %/MGy

## Slide 6
anomalous signal
Crick, F. H. C. & Magdoff, B. S. (1956) Acta Crystallogr. 9, 901-908.
Hendrickson, W. A. & Teeter, M. M. (1981) Nature 290, 107-113.
# sites
MW (Da)
ΔF
F
≈ 1.2 f”
√
World record!
ΔF/F = 0.5%
Wang, Dauter & Dauter (2006) Acta Cryst. D 62, 1475-1483.

## Slide 7: And the dominant source of error is…
Resolution
  Disorder (including Radiation Damage)
  Background
Phases
  Non-isomorphism (including Radiation Damage)
  Vibration (beam, xtal, spindle, etc.)
  Detector calibration

## Slide 8: Beam Flicker
time (seconds)
normalized flux through pinhole
pinhole
removed

## Slide 9: Shutter Jitter
open
closed
shutter jitter

## Slide 10: Optimal exposure time (anomalous on a PAD)
time
intensity
2.03 ms

## Slide 11: PAD full vs partial
iteration
Summed intensity

## Slide 12: Optimal exposure time (anomalous on a PAD)
tmad Optimal exposure time (s)
tro read-out time (s)
2.03 ms on Pilatus3 S
0.95 ms on Pilatus3 X
0.003 ms on Eiger
Attenuate for MAD
PILATUS
>0.2 s/frame
EIGER
>0.0003 s/frame
= 20,000 frames
350 GB/dataset

## Slide 13: And the dominant source of error is…
Resolution
  Disorder (including Radiation Damage)
  Background
Phases
  Non-isomorphism (including Radiation Damage)
  Vibration (beam, xtal, spindle, etc.)
  Detector calibration

## Slide 14: Detector calibration: 7235 eV

## Slide 15: Detector calibration: 7247 eV

## Slide 16: Gadox calibration vs energy
photon energy (keV)
Relative absorption depth
same = good!
bad!

## Slide 17
Pick-up tool mark
Bragg
glitch
oxygen
inclusions
“tree rings”
1% high
average
1% low
~3x105 photon/pixel Pilatus detector
Spatial Heterogeneity in Sharp Spot Sensitivity

## Slide 18: Systematic error does not “average out”

## Slide 19
Required signal-to-noise (I/σ)
Solve-able proteins (%)
CCD
old Pilatus
Goal
Threshold of a revolution in phasing
Best
Pilatus

## Slide 20: Fractional error
mult > (—)2
~3%
<ΔF/F>

## Slide 21
Holton & Frankel (2010) Acta D 66 393-408.

## Slide 22: Data collection parameters:
16 crystals
360° each, inverse beam
7235 eV photon energy
< 1 MGy per xtal
Australian Synchrotron MX1
  35 kGy/s into 100 μm x 100 μm

## Slide 23
RESOLUTION COMPLETENESS R-FACTOR I/SIGMA R-meas CC(1/2) Anomal SigAno Nano
LIMIT OF DATA observed Corr
9.17 99.1% 3.9% 257.47 3.9% 100.0* 91* 5.024 450
6.49 100.0% 5.2% 214.33 5.2% 100.0* 86* 3.836 882
5.30 100.0% 7.2% 165.13 7.2% 100.0* 76* 3.257 1175
4.59 100.0% 7.2% 175.42 7.3% 100.0* 67* 2.589 1403
4.10 99.9% 7.7% 174.13 7.7% 100.0* 59* 2.264 1594
3.74 99.9% 9.4% 143.09 9.4% 100.0* 49* 1.953 1783
3.47 100.0% 11.2% 120.17 11.2% 100.0* 39* 1.696 1942
3.24 100.0% 14.1% 91.14 14.1% 100.0* 30* 1.333 2103
3.06 99.9% 19.5% 65.79 19.5% 100.0* 23* 1.117 2214
2.90 99.9% 29.0% 44.85 29.1% 99.9* 17* 1.008 2369
2.77 99.9% 40.5% 32.58 40.6% 99.8* 11* 0.901 2493
2.65 99.9% 52.8% 25.16 52.9% 99.8* 10* 0.866 2605
2.54 100.0% 67.4% 19.47 67.6% 99.6* 2 0.804 2705
2.45 100.0% 88.9% 14.58 89.2% 99.2* 4 0.831 2859
2.37 100.0% 109.3% 9.97 109.7% 98.1* 5 0.829 2925
2.29 100.0% 138.2% 6.87 138.9% 96.1* 1 0.760 3037
2.22 100.0% 197.1% 4.03 198.6% 83.5* -1 0.721 3159
2.16 100.0% 227.3% 2.41 230.8% 46.9* -1 0.677 3224
2.10 61.2% 154.4% 1.28 163.6% 47.0* -2 0.660 1999
2.05 47.9% 170.1% 0.68 196.5% 25.7* 3 0.629 1578
total 93.3% 15.7% 54.30 15.8% 100.0* 12* 1.217 42499
140-fold multiplicity: 16 crystals, 360° each, inverse beam, 7235 eV

## Slide 24: 140-fold multiplicity
18 σ
Phased anomalous difference Fourier
16 σ

## Slide 25: 140-fold multiplicity
15σ = PO4
Phased anomalous difference Fourier

## Slide 26: 140-fold multiplicity
~2σ = Mg?
Phased anomalous difference Fourier

## Slide 27: 140-fold multiplicity
8.2σ = Mg?
DELFAN residual anomalous difference

## Slide 28: Discerning Na+ from Mg++
f’’ (electrons)
DELFAN peak height (σ)
Mg
Ne
Na
F
O
N

## Slide 29: 140-fold multiplicity
7.4σ = Na
DELFAN residual anomalous difference

## Slide 30: S-SAD merging 20 xtals
Knott et al. (2017). "Guide-bound CRISPR-Cas13a." Nature S&MB online

## Slide 31
Chart: 
| 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10
F+ | 52.3965 | 57.5371 | 48.523 | 47.5964 | 34.153 | 48.9106 | 70.3415 | 52.7871 | 59.3327 | 50.0622
F- | 46.3965 | 51.5371 | 42.523 | 41.5964 | 28.153 | 42.9106 | 64.3415 | 46.7871 | 53.3327 | 44.0622
observation
intensity
Systematic vs random error

## Slide 32: Summary
100 ADU/pixel
10 μm for lysozyme
~3% error per spot, 1%/MGy
7235 eV for S-SAD
“Attenu-wait”: dose slicing
http://bl831.als.lbl.gov/xtalsize.html
http://bl831.als.lbl.gov/~jamesh/mlfsom/
http://bl831.als.lbl.gov/~jamesh/powerpoint/ACMC_GBU_2013.ppt

## Slide 33: Summary
15 ADU/pixel on Rayonix HS 300
> 0.1 s exposures (HS 300)
10 μm for lysozyme
~3% error per spot, 1%/MGy
7235 eV for S-SAD on CCD
“Attenu-wait”: dose slicing
http://bl831.als.lbl.gov/xtalsize.html
http://bl831.als.lbl.gov/xtallife.html
http://bl831.als.lbl.gov/~jamesh/damage_rates.pdf
http://bl831.als.lbl.gov/~jamesh/powerpoint/CSHL_tipsNtrick_2017.pptx

## Slide 34: Why harvest at all?
Radiation damage
  10-100x reduced at ~100 K
Background scatter
    tray >> xtal
Timing
    nothing lasts forever
Shipping
    cryo samples are glass beads

## Slide 35: Why harvest at all?
Radiation damage
  10-100x reduced at ~100 K
Background scatter
    tray >> xtal
Timing
    nothing lasts forever
Shipping
    cryo samples are glass beads

## Slide 36: X-ray transmission of materials
Depth into material (μm)
transmission
glass
protein
crystal
water
air

## Slide 37: X-ray scattering of materials
Differential cross section
(scattered photons/solid angle/incident photon/meter thickness)
 7.5 3.8 2.5 1.9 1.5
d-spacing (Ǻ)
water
25% glycerol
paratone-N oil
plastic
glass

## Slide 38: Crowfoot Cell

## Slide 39: Crowfoot Cell

## Slide 40: Crowfoot Cell
J. D. Bernal and D. Crowfoot, Nature (London) 133, 794 (1934).

## Slide 41
pinhole
mirror
microscope
SBS
format
crystallization
tray
Notes: Inspired by the ESRF microcrystal camera, the superbend endstations are all equipped with zero-paralax optics. The user sees their sample from the point of view of the x-ray beam. Sample alignment is not only more stable but more intuitive than an off-axis camera. Placing the crystal into the microscope crosshairs gaurantees that it will be hit by the beam.

## Slide 42: Plate goniometers
Notes: The non-rotating “plate goniometer” (left) was developed last year. A new incarnation for taking rotation-camera data is now being developed (right). Either of these devices will accept most any crystallization tray that can be rotated 90 degrees to gravity and probe arbitrary points in each crystallization drop with a 50 um x-ray beam. Installing a plate goniometer takes ~30 minutes and the location ~300 mm downbeam from the “normal” protein crystallography sample mounting point makes this technology transferable to most any beamline.
Current status is that a 95% complete data set has been collected from a 50 um lysozyme crystal at room temperature, and trials are now underway to collect data sets from “real” crystals.

## Slide 43
Notes: non-rotating tray goniometer stage installed in ALS 8.3.1. Note that the “normal” cryo-pin goniometer on the left is still in place and a brass flight tube has been added to contain the air scatter from the main beam on its way to the tray goniometer.

## Slide 45
Fluidigm
Topaz® DC
http://www.fluidigm.com/home/fluidigm/docs/Datasheet_1.96DC.pdf
200 μm
Soliman, A. et al. (2011). Acta Cryst. D 67, 646-656.
250 μm
NatX-ray
CrystalQuick™ X
~50 μm tape
MiTeGen
MicroRT™
25 μm wall
25 μm wall
http://www.mitegen.com/products/micrort/micrort.shtml
PDMS
PDMS
http://www.mitegen.com/mic_catalog.php?c=insituplates
100 μm
MiTeGen
in-situ-1
~50 μm tape
in-situ diffraction systems
Márquez, J. & Cipriani, F. (2014) Structural Genomics (Chen, Y. W., ed) pp. 197-203.
Crystal DirectTM
~50 μm tape
25 μm
Notes: Next slide please

## Slide 46: What do we need?
Thin plastic = water = protein
Light low atomic number
Tight evaporation = nonisomorphism
Cheap $1-20/tray

## Slide 47: What we really need
GRAPHENE
Thin
Light
Tight
Cheap

## Slide 48: Why harvest at all?
Radiation damage
  10-100x reduced at ~100 K
Background scatter
    tray >> xtal
Timing
    nothing lasts forever
Shipping
    cryo samples are glass beads

## Slide 50
8000 Gy/s

## Slide 51: Example Data
lysozyme
50 μm beam
37 Gy/s (0.775 Å)
30s exposures at 20C
90° of data, 97% complete
I/ = 1.5 at 1.9 Å
Rmerge 18% (overall) 5% (low)
ΔB same as 20 min at 100K
B factor
-2
-4
-6
scale vs batch
2
1
0

## Slide 52: Why harvest at all?
Radiation damage
  10-100x reduced at ~100 K
Background scatter
    tray >> xtal
Timing
    nothing lasts forever
Shipping
    cryo samples are glass beads

## Slide 53: Why harvest at all?
Radiation damage
  10-100x reduced at ~100 K
Background scatter
    tray >> xtal
Timing
    nothing lasts forever
Shipping
    cryo samples are glass beads

## Slide 54: MiTeGen RT cell
http://www.mitegen.com/products/micrort/micrort.shtml

## Slide 55: Summary
assay, control and open mind
purity, purity, purity
oil is your friend
Warkentin method
try room temp!
http://bl831.als.lbl.gov/~jamesh/powerpoint/
CSHL_tipsNtricks_2015.pptx

## Slide 56: Pilatus pile-up for RT MX? same photons, different speeds
Sum of 193 shots with
193-fold attenuation

## Slide 57: Pilatus pile-up for RT MX? same photons, different speeds
1 shot with
no attenuation

## Slide 58: Pilatus pile-up for RT MX? same photons, different speeds
0.016° mosaic