## Slide 1: Acknowledgements
UCSF LBNL SLAC
ALS 8.3.1: TomAlberTron
UC Multicampus Research Programs and Initiatives (MRPI)
NIH NIGMS R01 GM124149
NIH NIGMS P30 GM124169
Integrated Diffraction Analysis Technologies (IDAT)
Plexxikon, Inc.
Synchrotron Radiation Structural Biology Resource (SLAC)
The Advanced Light Source is supported by the Director, Office of Science, Office of Basic Energy Sciences, Materials Sciences Division, of the US Department of Energy under contract No. DE-AC02-05CH11231 at Lawrence Berkeley National Laboratory.
Robert Stroud James Fraser John Spence
Chris Nielsen Clemens Schulze-Briese Aina Cohen Ana Gonzalez

## Slide 2: PowerPoint File available:
http://bl831.als.lbl.gov/
~jamesh/powerpoint/
CSHL_beamline_2018.pptx

## Slide 3: The “success rate” of data collection
100 s/dataset
200 days/year
~150 beamlines
~26,000,000 datasets/year
~10,000 PDBs in 2017

## Slide 4: signal vs noise
“If you don’t have good data, then you have no data at all.”
-Sung-Hou Kim

## Slide 5: signal vs noise
easy
hard
impossible

## Slide 6: signal vs noise
easy
hard
impossible
threshold of “solvability”

## Slide 7: signal vs noise
“If you don’t have good data, then you must learn statistics.”
-James Holton

## Slide 8: Adding noise

## Slide 9: Adding noise
12 + 12 = 1.42

## Slide 10: Adding noise
12 + 12 = 1.42
32 + 12 = 3.22
σtotal2 = σ12 + σ22

## Slide 11: Adding noise
12 + 12 = 1.42
32 + 12 = 3.22
σtotal2 = σ12 + σ22

## Slide 12: Adding noise
12 + 12 = 1.42
32 + 12 = 3.22
σtotal2 = σ12 + σ22

## Slide 13: Adding noise
12 + 12 = 1.42
32 + 12 = 3.22
102 + 12 = 10.052

## Slide 14: And the dominant source of error is…
Resolution
  Disorder (including Radiation Damage)
  Background
Phases
  Non-isomorphism (including Radiation Damage)
  Vibration (beam, xtal, spindle, etc.)
  Detector calibration

## Slide 15: 100% SeMet incorporation
Trivial to solve

## Slide 16: 100% S -Met incorporation
Impossible to solve

## Slide 17: Phase Problem: Se vs S
fraction not Selenium
correlation coefficient

## Slide 18: Phase Problem: Se vs S
fraction not Selenium
correlation coefficient

## Slide 19: Phase Problem: Se vs S
fraction not Selenium
correlation coefficient

## Slide 20: Phase Problem: Se vs S
fraction not Selenium
correlation coefficient

## Slide 21: Phase Problem: Se vs S
fraction not Selenium
correlation coefficient

## Slide 22: MR simulation
Rmsd from perfect search model (Å)
corrupted model
Correlation coefficient to correct density

## Slide 23: The “threshold of solvability” is sharp!
How can we predict success/failure?
Know Thy Experiment

## Slide 24: The “beam line”
Area Detector
Protein Crystal
Cite the “beamline paper!”
X-ray
Source
Notes: Most notable acheivements are high brightness with very little “wasted” light. The discovery made by Howard Padmore? At ALS showed that a 2:1 demagnification cancels nearly all of the spherical abberations in the parabolic/torroid system. 30-50% of the total x-ray intensity passes through the 100um pinhole. The flux through the 100um pinhole at the superbend beamlines is comparable to that of the 5.0.2 wiggler or an APS bending magnet, but only 150 Watts of power must be dissipated by the beamline optics. In contrast, wiggler beamlines must disipate kilowatts. The optics here are much simpler and not exotic in any way. The system is stable and robust. There is never a need to re-focus the beam at a new wavelength. Only a 2-minute automatic tuneup is recommended for very large energy changes (more than 1-2 keV).

## Slide 25: Where do the X-rays come from?

## Slide 26: X-rays (aka “Röntgenstrahlung”)
Röntgen, W. C. (1898)."Ueber eine neue Art von Strahlen", Annalen der Physik 300, 1-11.
The discovery soon after this of X-rays by Röntgen (22), the first investigator to use the type of tube described above, is generally considered to be a good example of a lucky discovery.
P.E.A. von Lenard
1906 Nobel Lecture

## Slide 27: Electricity + vacuum = X-rays
1654 ~10 Torr
Otto von Guericke’s
“Magdeburg hemispheres”
Benvenuti, C. & Chiggiato, P. (1993). "Obtention of pressures in the 10−14 torr range by means of a Zr-V-Fe non evaporable getter", Vacuum 44, 511-513.
1993 ~10-14 Torr
1865 < 10-2 Torr
Hg “fall tube” pump
1870s < 10-3 Torr
1894 < 10-5 Torr
Sprengel pump
Sprengel, H. (1865). "Researches on the vacuum", J. Chem. Soc. 18, 9b-21.
Redhead, Paul A., ed. Vacuum Science and Technology: Pioneers of the 20th Century. Vol. 2. Springer, 1994.

## Slide 28: The “Home source”
e-

## Slide 29: The “Home source”
e-
+
e-

## Slide 30: The “Home source”
+

## Slide 31: The “Home source”
e-
+

## Slide 32: The “Home source”

## Slide 33: Metal identification
scattering
same
λ as
incident

## Slide 34
“MAD scan” – absorption spectrum
“counts” (aka arbitrary units)
0.980888
0.980500
0.980113
0.979725
0.979338
0.978952
0.978565
0.978179
0.977794
0.977408
0.977023
0.976638
0.976254
Wavelength (Å)
fluorescence = yield * absorption
12398.42
energy

## Slide 35
maximize
absolute
value of
fpp
maximize
difference
in fp
“MAD scan” – absorption spectrum

## Slide 36
4-wavelength MAD

## Slide 37
3-wavelength MAD

## Slide 38
2-wavelength MAD

## Slide 39
2-wavelength MAD

## Slide 40
1-wavelength SAD

## Slide 41
2-wavelength metal validation
phased
anomalous
difference
Fourier

## Slide 42: Electric charge
balloon
Wool Sweater
-
-
-
-
-
-
+ + + + + +

## Slide 43: Electric field of a moving charge

## Slide 44: Electric field of a moving charge
accelerating charges
make
electromagnetic waves
(light)

## Slide 45: Bending Magnet
N
electron beam
X-rays !
S

## Slide 46: Two Bending Magnets
N
electron beam
S
S
N
2x X-rays !

## Slide 47: Wiggler
N
S
N
S
N
S
N
S
S
N
S
N
S
N
S
N
8x X-rays !

## Slide 48: Undulator
N
S
N
S
N
S
N
S
S
N
S
N
S
N
S
N
> 8x X-rays !

## Slide 49: Undulator emission spectrum

## Slide 50: XFEL Undulator
N
S
N
S
N
S
N
S
S
N
S
N
S
N
S
N
same X-rays
short pulse
S
N
S
N
S
N
S
N
N
S
N
S
N
S
N
S

## Slide 51: LCLS undulator hall

## Slide 52: LCLS at SLAC

## Slide 53: Accelerator-based light source terminology
Code name Translation
First Generation Bending Magnet
Second Generation Wiggler
Third Generation Undulator
Fourth Generation Free electron laser
..and diffraction-limited synchrotron

## Slide 54: Diffraction-limited light source
λ ≈ 2 d sinθ
λ = x-ray wavelength
d = electron/x-ray beam size
θ = beam divergence
d = 1 Å / 100 µrad / 2 = 0.5 µm

## Slide 55: The “beam line”
Area Detector
Source
Primary
mirror
Secondary
mirror
Si(111)
monochromator
Sample
focused
beam
slits
Scatter
guard
Unfocused
beam
slits

## Slide 56: The “beam line”
Area Detector
Source
Primary
mirror
Secondary
mirror
Si(111)
monochromator
Sample
focused
beam
slits
Scatter
guard
Unfocused
beam
slits
Notes: Most notable acheivements are high brightness with very little “wasted” light. The discovery made by Howard Padmore? At ALS showed that a 2:1 demagnification cancels nearly all of the spherical abberations in the parabolic/torroid system. 30-50% of the total x-ray intensity passes through the 100um pinhole. The flux through the 100um pinhole at the superbend beamlines is comparable to that of the 5.0.2 wiggler or an APS bending magnet, but only 150 Watts of power must be dissipated by the beamline optics. In contrast, wiggler beamlines must disipate kilowatts. The optics here are much simpler and not exotic in any way. The system is stable and robust. There is never a need to re-focus the beam at a new wavelength. Only a 2-minute automatic tuneup is recommended for very large energy changes (more than 1-2 keV).

## Slide 57: X-ray optics: mirrors
Notes: Next slide please

## Slide 58: Why are the mirrors so long?
3 mm
Must be < 0.2°
1000 mm
Notes: Next slide please

## Slide 59
Single crystal

## Slide 60: Monochromators
λ = 1 Å
Si
Notes: Next slide please

## Slide 61: Monochromators
λ = 1 Å
Si
d = 3.135 Å
θ = 9.177°
λ = 1 Å
nλ = 2d sin(θ)
Notes: Next slide please

## Slide 62: Monochromators
nλ = 2d sin(θ)
λ = 1.1 Å
Si
d = 3.135 Å
θ = 9.177°
Notes: Next slide please

## Slide 63: Monochromators
nλ = 2d sin(θ)
λ = 1.1 Å
Si
d = 3.135 Å
θ = 10.104°
λ = 1.1 Å
Notes: Next slide please

## Slide 64: Monochromators
nλ = 2d sin(θ)
λ = 1 Å
Si
d = 3.135 Å
θ = 9.177°
λ = 1 Å
Notes: Next slide please

## Slide 65: Monochromators
Si
λ = 1 Å
λ = 1 Å
Si
d = 3.135 Å
DCM = Double Crystal Monochromator
Notes: Next slide please

## Slide 66: Monochromators
Si
λ = 1.1 Å
λ = 1.1 Å
Si
d = 3.135 Å
DCM = Double Crystal Monochromator
Notes: Next slide please

## Slide 67: Monochromators
nλ = 2d sin(θ)
λ = 1.54 Å
θ = 1.84°
λ = 1.54 Å
Mo/C
d = 24 Å
Multilayer optics
Notes: Next slide please

## Slide 68: The truth about x-ray beams
Term units significance
Flux photons/s duration of experiment
Beam Size μm match to crystal
Divergence mrad spot size vs distance
Wavelength Å resolution and absorption
Dispersion Δλ/λ spot size
Flux density ph/s/area scattering/damage rate
Fluence ph/area scattering/damage

## Slide 69: The truth about x-ray beams
Term units significance
Flux photons/s duration of experiment
Beam Size μm match to crystal
Divergence mrad spot size vs distance
Wavelength Å resolution and absorption
Dispersion Δλ/λ spot size
Flux density ph/s/area scattering/damage rate
Fluence ph/area scattering/damage

## Slide 70
The number of photons scattered
before crystal is dead
is independent
of data collection time
1 um3 = 106 photons
Henderson, 1990; Gonzalez & Nave, 1994; Glaeser et al., 2000; Sliz et al., 2003; Leiros et al., 2006; Owen et al., 2006; Garman & McSweeney, 2006; Garman & Nave, 2009; Holton, 2009

## Slide 71: The truth about x-ray beams
quantity | units | home source | MX2
flux | Photons/second | 2.5 x 109 | 2 x 1012
exposure | seconds | 400 | 1

## Slide 72: The truth about x-ray beams
quantity | units | home source | MX2
flux | Photons/second | 2.5 x 109 | 2 x 1012
exposure | seconds | 400 | 1
Dispersion | wavelength range / wavelength | 0.2% (Kα1 - Kα2) | 0.014% (Si111)
Divergence | milliRadian | 4.8 | 1.8 (h) 0.18 (v)
Beam size | microns | 100 | 24 (h) 12 (v)
Spectral brightness | Photons/s/mm2/ mR2/0.1%BW | 5.4 x 109 | 1.5 x 1017

## Slide 73: The truth about x-ray beams
quantity | units | home source | MX2
flux | Photons/second | 72 x 106 | 2 x 1012
exposure | time | 4 hours | 1 second
Dispersion | wavelength range / wavelength | 0.2% (Kα1 - Kα2) | 0.014% (Si111)
Divergence | milliRadian | 4.8 | 1.8 (h) 0.18 (v)
Beam size | microns | 17 | 24 (h) 12 (v)
Spectral brightness | Photons/s/mm2/ mR2/0.1%BW | 5.4 x 109 | 1.5 x 1017

## Slide 74: The truth about x-ray beams
quantity | units | home source | MX2
flux | Photons/second | 1 x 106 | 2 x 1012
exposure | time | 10 days | 1 second
Dispersion | wavelength range / wavelength | 0.2% (Kα1 - Kα2) | 0.014% (Si111)
Divergence | milliRadian | 0.6 | 1.8 (h) 0.18 (v)
Beam size | microns | 17 | 24 (h) 12 (v)
Spectral brightness | Photons/s/mm2/ mR2/0.1%BW | 5.4 x 109 | 1.5 x 1017

## Slide 75: The truth about x-ray beams
quantity | units | home source | MX2
flux | Photons/second | 8 x 104 | 2 x 1012
exposure | time | 5 months | 1 second
Dispersion | wavelength range / wavelength | 0.014% (Si111) | 0.014% (Si111)
Divergence | milliRadian | 0.6 | 1.8 (h) 0.18 (v)
Beam size | microns | 17 | 24 (h) 12 (v)
Spectral brightness | Photons/s/mm2/ mR2/0.1%BW | 5.4 x 109 | 1.5 x 1017

## Slide 76: The truth about x-ray beams
Term units significance
Flux photons/s duration of experiment
Beam Size μm match to crystal
Divergence mrad spot size vs distance
Wavelength Å resolution and absorption
Dispersion Δλ/λ spot size
Flux density ph/s/area scattering/damage rate
Fluence ph/area scattering/damage

## Slide 77
Two-color experiment (ca 1914)
(Cu source)
20 eV

## Slide 78: Spot Shape
Ewald sphere
(h,k,l)

## Slide 79: Spot Shape
Ewald sphere
(h,k,l)
d*

## Slide 80: Spot Shape
Ewald sphere
spindle axis
(h,k,l)
d*

## Slide 81: Spot Shape
Ewald sphere
spindle axis
Φ circle
(h,k,l)

## Slide 82: Spot Shape
Ewald sphere
spindle axis
Φ circle
(h,k,l)

## Slide 83: Spot Shape
Ewald sphere
spindle axis
Φ circle
diffracted ray
(h,k,l)

## Slide 84: Spot Shape
Ewald sphere
spindle axis
Φ circle
diffracted ray
(h,k,l)
d*
λ*
λ*

## Slide 85: Spot Shape
Ewald sphere
spindle axis
Φ circle
diffracted ray
(h,k,l)
d*
λ*
λ*
θ
λ = 2 d sin(θ)

## Slide 86: spectral dispersion
Ewald sphere
spindle axis
Φ circle
diffracted ray
(h,k,l)
d*
λ*
λ*

## Slide 87: spectral dispersion
Ewald sphere
spindle axis
Φ circle
diffracted ray
(h,k,l)
d*
λ’*
λ’*

## Slide 88: dispersion = 0

## Slide 89: dispersion = 0.014% (Si 111)

## Slide 90: dispersion = 0.25% (CuKα)

## Slide 91: dispersion = 0.6%

## Slide 92: dispersion = 1.3%

## Slide 93: dispersion = 2.6%

## Slide 94: dispersion = 5.1%

## Slide 95: The truth about x-ray beams
Term units significance
Flux photons/s duration of experiment
Beam Size μm match to crystal
Divergence mrad spot size vs distance
Wavelength Å resolution and absorption
Dispersion Δλ/λ spot size
Flux density ph/s/area scattering/damage rate
Fluence ph/area scattering/damage

## Slide 96: beam divergence
Ewald sphere
spindle axis
Φ circle
diffracted ray
(h,k,l)
d*
λ*
λ*

## Slide 97: beam divergence
spindle axis
Φ circle
diffracted ray
(h,k,l)
d*
Ewald sphere
λ*
λ*

## Slide 98: divergence = 0 º

## Slide 99: divergence = 0.3 º

## Slide 100: Divergence Arndt & Wonacott (1977)

## Slide 101: The truth about x-ray beams
Term units significance
Flux photons/s duration of experiment
Beam Size μm match to crystal
Divergence mrad spot size vs distance
Wavelength Å resolution and absorption
Dispersion Δλ/λ spot size
Flux density ph/s/area scattering/damage rate
Fluence ph/area scattering/damage

## Slide 102: Air absorption
Distance from sample (mm)
transmittance

## Slide 103: The truth about x-ray beams
Term units significance
Flux photons/s duration of experiment
Beam Size μm match to crystal
Divergence mrad spot size vs distance
Wavelength Å resolution and absorption
Dispersion Δλ/λ spot size
Flux density ph/s/area scattering/damage rate
Fluence ph/area scattering/damage

## Slide 104: Data quality - simple rules:
  Put your crystal into the beam

## Slide 105: Your sample is a lens! (but not for x-rays)
90°

## Slide 106: Your sample is a lens! (but not for x-rays)
90°

## Slide 107: Data quality - simple rules:
  Put your crystal into the beam
  Shoot the whole crystal

## Slide 108: shoot the whole crystal

## Slide 109: shoot the whole crystal

## Slide 110: shoot the whole crystal

## Slide 111: shoot the whole crystal

## Slide 112: shoot the whole crystal

## Slide 113: How many crystals do you see?
Shoot the crystal (singular)

## Slide 114
Ewald’s “mosaic” picture

## Slide 115
Ewald’s “mosaic” picture

## Slide 116
Ewald’s “mosaic” picture
What is
this stuff?

## Slide 117
Darwin’s original picture

## Slide 118
“mosaicity” with visible light

## Slide 119: mosaic spread = 0 º

## Slide 120: mosaic spread = 0.1º

## Slide 121: mosaic spread = 0.2º

## Slide 122: mosaic spread = 0.4º

## Slide 123: mosaic spread = 0.6º

## Slide 124: mosaic spread = 0.8º

## Slide 125: mosaic spread = 1.0º

## Slide 126: mosaic spread = 1.5º

## Slide 127: mosaic spread = 2.0º

## Slide 128: mosaic spread = 2.5º

## Slide 129: mosaic spread = 3.2º

## Slide 130: mosaic spread = 6.4º

## Slide 131: mosaic spread = 12.8º

## Slide 132: mosaic spread
Ewald sphere
spindle axis
Φ circle
diffracted ray
(h,k,l)
λ*
λ*
d*
d*

## Slide 133: mosaic spread
Ewald sphere
spindle axis
Φ circle
diffracted rays
(h,k,l)
d*
d*

## Slide 134: Data quality - simple rules:
  Put your crystal into the beam
  Shoot the whole crystal
  Shoot nothing but the crystal

## Slide 135: shoot nothing but the crystal

## Slide 136: shoot nothing but the crystal

## Slide 137: shoot nothing but the crystal

## Slide 138: shoot nothing but the crystal

## Slide 139: shoot nothing but the crystal

## Slide 140: Gaussian beam
50% of flux
the other 50%

## Slide 141: Collimated beam
100% of flux
0%

## Slide 142: Tipping point for µfocus:
10 µm
~10 mm
~1 mrad =0.06°

## Slide 143: scattering/absorptoin “rules”:
1 μm crystal ≈ 1 μm water
≈ 1 μm plastic
≈ 0.1 μm glass
≈ 1000 μm air
Optimum: support thickness = crystal thickness

## Slide 144
X-ray beam size (μm)
Resolution (Å)
too big
too small
just right

## Slide 145: Data quality - simple rules:
  Put your crystal into the beam
  Shoot the whole crystal
  Shoot nothing but the crystal
  Back off!

## Slide 146: Background scattering

## Slide 147: Mosquito tips cost $0.15 each
So do detector pixels!
We bought 6,224,001
Use them!
(they are reusable)

## Slide 148: Background scattering
$100,000.00
$100,000.00
$100,000.00
$100,000.00
$100,000.00
$100,000.00
$100,000.00
$100,000.00
real estate is
expensive
use it!

## Slide 149: Fine Slicing
Pflugrath, J. W. (1999)."The finer things in X-ray diffraction data collection", Acta Cryst. D 55, 1718-1725.
background
background

## Slide 150: Optimal exposure time (faint spots)
thr Optimal exposure time for data set (s)
tref exposure time of reference image (s)
bgref background level near weak spots on
reference image (ADU)
bg0 ADC offset of detector (ADU)
bghr optimal background level (via thr)
σ0 rms read-out noise (ADU)
gain ADU/photon
m multiplicity of data set (including partials)
adjust exposure
so this is ~100

## Slide 151: Optimal exposure time (anomalous on a PAD)
time
intensity
2.03 ms

## Slide 152: PAD full vs partial
iteration
Summed intensity

## Slide 153: Optimal exposure time (anomalous on a PAD)
tmad Optimal exposure time (s)
tro read-out time (s)
2.03 ms on Pilatus3 S
0.95 ms on Pilatus3 X
0.003 ms on Eiger
Attenuate for MAD
PILATUS
>0.2 s/frame
EIGER
>0.0003 s/frame

## Slide 154: And the dominant source of error is…
Resolution
  Disorder (including Radiation Damage)
  Background
Phases
  Non-isomorphism (including Radiation Damage)
  Vibration (beam, xtal, spindle, etc.)
  Detector calibration

## Slide 155
Pick-up tool mark
Bragg
glitch
oxygen
inclusions
“tree rings”
1% high
average
1% low
~3x105 photon/pixel Pilatus detector
Spatial Heterogeneity in Sharp Spot Sensitivity

## Slide 156: Detector calibration: 7235 eV

## Slide 157: Detector calibration: 7247 eV

## Slide 158: Gadox calibration vs energy
photon energy (keV)
Relative absorption depth
same = good!
bad!

## Slide 159: Systematic error does not “average out”

## Slide 160
Required signal-to-noise (I/σ)
Solve-able proteins (%)
CCD
old Pilatus
Goal
Threshold of a revolution in phasing
Best
Pilatus

## Slide 161: Can you count to 1,000,000 ?
= 0.1%
sqrt(1,000,000)
1,000,000
= 3%
sqrt(1,000)
1,000
> 1000 is a waste!
photon
spot
Theoretically:
In reality:
ISa ~ 33
Rmeas ≈ 0.1% ?
ISa = 1000
Rmeas = ≈ 3%

## Slide 162: Fractional error
mult > (—)2
~3%
<ΔF/F>

## Slide 163
anomalous signal
Crick, F. H. C. & Magdoff, B. S. (1956) Acta Crystallogr. 9, 901-908.
Hendrickson, W. A. & Teeter, M. M. (1981) Nature 290, 107-113.
# sites
MW (Da)
ΔF
F
≈ 1.2 f”
√
World record!
ΔF/F = 0.5%
Wang, Dauter & Dauter (2006) Acta Cryst. D 62, 1475-1483.

## Slide 164
Holton & Frankel (2010) Acta D 66 393-408.

## Slide 165: S-SAD merging 20 xtals
Knott et al. (2017). "Guide-bound CRISPR-Cas13a." Nature S&MB online

## Slide 166
Chart: 
| 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10
F+ | 52.3965 | 57.5371 | 48.523 | 47.5964 | 34.153 | 48.9106 | 70.3415 | 52.7871 | 59.3327 | 50.0622
F- | 46.3965 | 51.5371 | 42.523 | 41.5964 | 28.153 | 42.9106 | 64.3415 | 46.7871 | 53.3327 | 44.0622
observation
intensity
Systematic vs random error

## Slide 167: Data collection parameters:
16 crystals
360° each, inverse beam
7235 eV photon energy
< 1 MGy per xtal
Australian Synchrotron MX1
  35 kGy/s into 100 μm x 100 μm

## Slide 168
RESOLUTION COMPLETENESS R-FACTOR I/SIGMA R-meas CC(1/2) Anomal SigAno Nano
LIMIT OF DATA observed Corr
9.17 99.1% 3.9% 257.47 3.9% 100.0* 91* 5.024 450
6.49 100.0% 5.2% 214.33 5.2% 100.0* 86* 3.836 882
5.30 100.0% 7.2% 165.13 7.2% 100.0* 76* 3.257 1175
4.59 100.0% 7.2% 175.42 7.3% 100.0* 67* 2.589 1403
4.10 99.9% 7.7% 174.13 7.7% 100.0* 59* 2.264 1594
3.74 99.9% 9.4% 143.09 9.4% 100.0* 49* 1.953 1783
3.47 100.0% 11.2% 120.17 11.2% 100.0* 39* 1.696 1942
3.24 100.0% 14.1% 91.14 14.1% 100.0* 30* 1.333 2103
3.06 99.9% 19.5% 65.79 19.5% 100.0* 23* 1.117 2214
2.90 99.9% 29.0% 44.85 29.1% 99.9* 17* 1.008 2369
2.77 99.9% 40.5% 32.58 40.6% 99.8* 11* 0.901 2493
2.65 99.9% 52.8% 25.16 52.9% 99.8* 10* 0.866 2605
2.54 100.0% 67.4% 19.47 67.6% 99.6* 2 0.804 2705
2.45 100.0% 88.9% 14.58 89.2% 99.2* 4 0.831 2859
2.37 100.0% 109.3% 9.97 109.7% 98.1* 5 0.829 2925
2.29 100.0% 138.2% 6.87 138.9% 96.1* 1 0.760 3037
2.22 100.0% 197.1% 4.03 198.6% 83.5* -1 0.721 3159
2.16 100.0% 227.3% 2.41 230.8% 46.9* -1 0.677 3224
2.10 61.2% 154.4% 1.28 163.6% 47.0* -2 0.660 1999
2.05 47.9% 170.1% 0.68 196.5% 25.7* 3 0.629 1578
total 93.3% 15.7% 54.30 15.8% 100.0* 12* 1.217 42499
140-fold multiplicity: 16 crystals, 360° each, inverse beam, 7235 eV

## Slide 169: 140-fold multiplicity
18 σ
Phased anomalous difference Fourier
16 σ

## Slide 170: 140-fold multiplicity
15σ = PO4
Phased anomalous difference Fourier

## Slide 171: 140-fold multiplicity
~2σ = Mg?
Phased anomalous difference Fourier

## Slide 172: 140-fold multiplicity
8.2σ = Mg?
DELFAN residual anomalous difference

## Slide 173: Discerning Na+ from Mg++
f’’ (electrons)
DELFAN peak height (σ)
Mg
Ne
Na
F
O
N

## Slide 174: 140-fold multiplicity
7.4σ = Na
DELFAN residual anomalous difference

## Slide 175
FOM from “dm”
Correlation to final map
XFEL Phasing success vs Figure of Merit
Hunter et al. (2016)."Selenium anomalous phasing XFEL", Nature Comm. 7, 13388.
http://cxidb.org/id54-SfjMJjaBny/id-54.html http://cxidb.org/id55-BWuBDSUrBy/id-55.html
Adjusting:
Solvent content
Resolution
SDfac Sdadd
Cycles of dm
“mode”
“scheme”
NCS radius
NCS cycles
Phase blur

## Slide 176: Data quality - simple rules:
  Put your crystal into the beam
  Shoot the whole crystal
  Shoot nothing but the crystal
  Back off!
  The crystal must rotate

## Slide 177: The crystal rotates!

## Slide 178: The crystal rotates!
Membrane Protein Expression Center © 2012

## Slide 179: avoiding overlaps
c
c

## Slide 180: avoiding overlaps
c

## Slide 181: avoiding overlaps
c
http://www.mitegen.com/mic_catalog.php?c=DTMicroMounts

## Slide 182: avoiding overlaps
c
http://www.mitegen.com/mic_catalog.php?c=orientedmounts

## Slide 183: One xtal, no idea? Suggested native protocol:
Do strategy
Start at specified phi
Complete fast as you can
Keep going to 360°
Move detector
Advance phi 10-30 deg
Increase photons/image 4x
goto 4

## Slide 184: Adding images
% cat << EOF > MERGE2CBF.INP
NAME_TEMPLATE_OF_DATA_FRAMES= /data/you/weak_?????.cbf
DATA_RANGE= 1 10
NAME_TEMPLATE_OF_OUTPUT_FRAMES=/data/you/sum_???.cbf
NUMBER_OF_DATA_FRAMES_COVERED_BY_EACH_OUTPUT_FRAME= 10
EOF
% merge2cbf
Which is better?:
% xds_runme.com /data/you/weak_?????.cbf
% xds_runme.com /data/you/sum_?????.cbf
“lossy”
compression

## Slide 185: “pretty” crystals

## Slide 186: “pretty” crystals

## Slide 187: “pretty” crystals

## Slide 188: Summary
1 + 1 = 1.4
Match beam to “crystal”
Match goop thickness to crystal
Better beam = less time
Look at your spots!
Avoid overlaps
Crystal must rotate

## Slide 189: Correct beam center
155.063 155.647
159.353 155.063
159.301 155.011
1512.73 1554.57
155.063 159.356 -250
ADXV
HKL2000
MOSFLM
XDS
DIALS

## Slide 190: Where do photons go?
beamstop
Transmitted (98%)
Protein
1A x-rays

## Slide 191: Elastic scattering

## Slide 192: Elastic scattering

## Slide 193: Where do photons go?
beamstop
elastic scattering (6%)
Transmitted (98%)
Protein
1A x-rays

## Slide 194: Elastic scattering

## Slide 195: Inelastic scattering

## Slide 196: Where do photons go?
beamstop
elastic scattering (6%)
Transmitted (98%)
Protein
1A x-rays

## Slide 197: Where do photons go?
beamstop
elastic scattering (6%)
Transmitted (98%)
inelastic scattering (7%)
Protein
1A x-rays

## Slide 198: Where do photons go?
beamstop
elastic scattering (6%)
Transmitted (98%)
inelastic scattering (7%)
Protein
1A x-rays
Re-emitted (99%)
Absorbed (~0%)

## Slide 199: Photoelectric absorption

## Slide 200: Photoelectric absorption
e-
+

## Slide 201: Fluorescence
+

## Slide 202: Fluorescence
e-
+

## Slide 203: Auger emission
+

## Slide 204: Auger emission
++
e-

## Slide 205: Sample atoms

## Slide 206: Sample atoms

## Slide 207: Primary ionization
e-
+

## Slide 208: Secondary ionizations
e-
e-
e-
e-
e-
e-
+
+
+
+
+
+

## Slide 209: Ionization track
e-
e-
e-
e-
e-
e-
e-
e-
e-
+
+
+
+
+
+
+
+
+
e-
+

## Slide 212: Where do photons go?
beamstop
elastic scattering (6%)
Transmitted (98%)
inelastic scattering (7%)
Protein
1A x-rays
Re-emitted (99%)
Absorbed (~0%)

## Slide 213: Where do photons go?
beamstop
elastic scattering (6%)
Transmitted (98%)
inelastic scattering (7%)
Photoelectric (87%)
Protein
1A x-rays
Re-emitted (99%)
Absorbed (~0%)

## Slide 214: Where do photons go?
beamstop
elastic scattering (6%)
Transmitted (98%)
inelastic scattering (7%)
Photoelectric (87%)
Protein
1A x-rays
Re-emitted (~0%)
Absorbed (99%)
Re-emitted (99%)
Absorbed (~0%)

## Slide 215: Where do photons go?
beamstop
elastic scattering (6%)
Transmitted (98%)
useful/absorbed energy: 7.3%
inelastic scattering (7%)
Photoelectric (87%)
Protein
1A x-rays
Re-emitted (~0%)
Absorbed (99%)
Re-emitted (99%)
Absorbed (~0%)

## Slide 216: What do we want?
Resolution
Mosaicity
Nice, round spots
No overlaps
No twinning
Lots of crystals
…

## Slide 218: What do we want?
Resolution
Mosaicity
Nice, round spots
No overlaps
No twinning
Lots of crystals
…
2. Phases

## Slide 219: Phases
http://bl831.als.lbl.gov/~jamesh/movies/dephase.mpeg

## Slide 220: And the dominant source of error is…
Resolution
  Disorder (including Radiation Damage)
  Background
Phases
  Non-isomorphism (including Radiation Damage)
  Vibration (beam, xtal, spindle, etc.)
  Detector calibration

## Slide 221: What is “disorder”?

## Slide 222: What is “disorder”?

## Slide 223: What is “disorder”?

## Slide 224: What is “disorder”?
order
order
disorder
B-factor

## Slide 225: “B” factors
ATOM 122 N LEU A 13 -3.244 25.808 19.998 1.00 16.96 N
ATOM 123 CA LEU A 13 -2.877 25.448 21.355 1.00 15.29 C
ATOM 124 C LEU A 13 -2.792 23.966 21.561 1.00 17.54 C
ATOM 125 O LEU A 13 -1.814 23.493 22.143 1.00 16.35 O
ATOM 126 CB LEU A 13 -3.907 26.164 22.268 1.00 18.72 C
ATOM 127 CG LEU A 13 -3.577 25.982 23.738 1.00 21.19 C
ATOM 128 CD1 LEU A 13 -2.283 26.820 24.019 1.00 19.43 C
ATOM 129 CD2 LEU A 13 -4.702 26.474 24.639 1.00 24.65 C
ATOM 130 N SER A 14 -3.677 23.149 20.979 1.00 15.96 N
ATOM 131 CA SER A 14 -3.646 21.711 21.061 1.00 18.26 C
ATOM 132 C SER A 14 -2.373 21.203 20.360 1.00 18.71 C
ATOM 133 O SER A 14 -1.747 20.315 20.930 1.00 17.47 O
ATOM 134 CB SER A 14 -4.875 21.077 20.419 1.00 17.62 C
ATOM 135 OG ASER A 14 -4.825 19.665 20.388 0.50 20.89 O
ATOM 136 OG BSER A 14 -6.027 21.408 21.164 0.50 18.67 O
ATOM 137 N LYS A 15 -2.045 21.772 19.215 1.00 18.03 N
ATOM 138 CA LYS A 15 -0.799 21.361 18.555 1.00 18.12 C
ATOM 139 C LYS A 15 0.446 21.707 19.351 1.00 18.81 C
ATOM 140 O LYS A 15 1.400 20.948 19.411 1.00 17.77 O
ATOM 141 CB LYS A 15 -0.700 22.034 17.177 1.00 14.49 C
ATOM 142 CG LYS A 15 -1.727 21.368 16.256 1.00 16.12 C
ATOM 143 CD LYS A 15 -1.663 22.147 14.936 1.00 19.40 C
ATOM 144 CE ALYS A 15 -2.725 21.614 13.986 0.50 17.42 C
ATOM 145 CE BLYS A 15 -1.750 21.211 13.750 0.50 17.01 C
ATOM 146 NZ ALYS A 15 -2.346 21.674 12.559 0.50 18.61 N
ATOM 147 NZ BLYS A 15 -3.052 20.513 13.741 0.50 18.76 N

## Slide 226: The B factor
Bragg, W. H. (1914)."XCVII. The intensity of reflexion of X rays by crystals", Philosophical Magazine Series 6 27, 881-899. Published May 1 1914

## Slide 227: B factor from image analysis
B = 500

## Slide 228: B factor from image analysis
B = 20

## Slide 229
Crystal diameter (micron)
Resolution (Å)
B = 0
B = 20
B = 60
B = 170
B = 330
B = 560
B = 850
Bigger is better, but not by much

## Slide 230
disorder
+ averaging
= poor resolution
Diffraction Take-Home Lesson:

## Slide 231: What Causes Disorder?
Function
Radiation damage
Mosaicity
Strain
Xtal abuse
Dehydration

## Slide 232: What Causes Disorder?
Function
Radiation damage
Mosaicity
Strain
Xtal abuse
Dehydration

## Slide 233: Muybridge’s galloping horse (1878)

## Slide 234: realistic “crystal” of horses

## Slide 235: average structure: galloping horse

## Slide 236: Ways to improve order
New crystal form
Purify
  - add a column
  - heat shock
Bind something
  - Known ligand
  - Silver bullets
Change the clone
  - chop off floppy bits
  - homolog
  - Ala -> Lys mutants
Make Protein Sit Still!!!
Newman J. (2006) Acta Cryst. D62 27-31.
8000 cases in PDB of resolution improvement
80% of these are different crystal forms
> 1 Å improvement → 94% different forms

## Slide 237: What Causes Disorder?
Function
Radiation damage
Mosaicity
Strain
Xtal abuse
Dehydration

## Slide 238: Types of radiation damage
Global
Specific

## Slide 239: Types of radiation damage
Global
  scale factor
  B factor
  non-isomorphism
Specific
  depends on chemistry
  destroys phasing signal

## Slide 240
Global damage
Howells, et al. (2009) J. Electron Spectrosc. Relat. Phenom. 170, 4-12.

## Slide 241: resolution dependence of global damage
resolution (Å)
maximum tolerable dose (MGy)
1 2 3 5 7 10 20 40 70 100
1
10
100
103
Howells et al. (2009) J. Electron. Spectrosc. Relat. Phenom. 170 4-12

## Slide 242: resolution dependence of global damage
resolution (Å)
maximum tolerable dose (MGy)
1 2 3 5 7 10 20 40 70 100
1
10
100
103
Howells et al. (2009) J. Electron. Spectrosc. Relat. Phenom. 170 4-12
10 MGy/Å

## Slide 243: what the is a MGy?
http://bl831.als.lbl.gov/
damage_rates.pdf
Holton J. M. (2009) J. Synchrotron Rad. 16 133-42
1 MGy = 106 J/kg

## Slide 244
How long will my crystal last?
synch line type flux beamsize flux density dose max xtal min site
ph/s μm ph/μm2/s rate lifetime lifetime
APS 5ID-B POWDER 5e12 1000x300 1.7e+07 8.25 kGy/s 61 m 4 m
APS 8-BM MAD 1e11 200 2.5e+06 858 Gy/s 9.7 h 39 m
APS 14-BM-C MONO 6e11 130x200 2.3e+07 11 kGy/s 46 m 3 m
APS 14-ID-B LAUE 7e13 15x20 2.3e+11 132 MGy/s 0.23 s 0.015 s
APS 17-BM POWDER 8e11 200 2.0e+07 8.72 kGy/s 57 m 3.8 m
APS 17-ID MAD 7e11 200 1.8e+07 9.87 kGy/s 51 m 3.4 m
APS 18-ID SAXS 2e13 50x150 2.7e+09 2.69 MGy/s 11 s 0.74 s
APS 19-BM MAD 6e+08 60x50 2.0e+05 77.3 Gy/s 4.5 d 7.2 h
APS 19-ID MAD 3e+10 100x20 1.5e+07 7.81 kGy/s 64 m 4.3 m
APS 21-ID-D MAD 5e12 50 2.0e+09 1.13 MGy/s 27 s 1.8 s
APS 21-ID-F MONO 1e12 50 4.0e+08 190 kGy/s 2.6 m 11 s
APS 21-ID-G MONO 1e12 50 4.0e+08 190 kGy/s 2.6 m 11 s
APS 22-BM MAD 7e12 40x80 2.2e+09 750 kGy/s 40 s 2.7 s
APS 22-ID MAD 7e12 40x80 2.2e+09 750 kGy/s 40 s 2.7 s
APS 23-BM-B MAD 2e11 200? 5.0e+06 2.47 kGy/s 3.4 h 13 m
APS 23-ID-B MAD 6e12 80x20 3.8e+09 2.11 MGy/s 14 s 0.95 s
APS 23-ID-D MAD 1e13 75x25 5.3e+09 3.01 MGy/s 10 s 0.66 s
APS 24-ID-C MAD 2.5e12 40x60 1.0e+09 357 kGy/s 84 s 5.6 s
APS 24-ID-E MONO 2.85e12 20x120 1.2e+09 563 kGy/s 53 s 3.5 s
APS 31-ID MAD 2.6e12 120x80 2.7e+08 105 kGy/s 4.8 m 19 s
NSLS-II 17-ID-1 MAD 5e12 7x5 1.4e+11 67 MGy/s 0.45 s 0.03 s
NSLS-II 17-ID-2 MAD 2e12 1x1 1.3e+12 626 MGy/s 0.048 s 0.0032 s
NSLS X3A MAD 3.8e10 200 9.5e+05 451 Gy/s 18 h 74 m
NSLS X3B EXAFS 9e11 200x700 6.4e+06 6.48 kGy/s 77 m 5.1 m
NSLS X4A MAD 2e10 300 2.2e+05 136 Gy/s 61 h 4.1 h
NSLS X6A MAD 1.5e12 100x500 3.0e+07 10.3 kGy/s 49 m 3.2 m
NSLS X8C MAD 1.1e10 200? 2.8e+05 204 Gy/s 41 h 2.7 h

## Slide 246
Holton & Frankel (2010) Acta D 66 393-408.

## Slide 247: Multi-crystal strategies
Kendrew et al. (1960) "Structure of Myoglobin” Nature 185, 422-427.

## Slide 248: Specific damage
final, refined phases

## Slide 249: Specific damage
experimentally-obtained phases

## Slide 250
Holton, J. M. (2007) J. Synch. Rad. 14, 51-72.

## Slide 251
1.0
0.0
fraction
unconverted
Holton, J. M. (2007) J. Synch. Rad. 14, 51-72.

## Slide 252
Holton (2009) J. Synchrotron Rad. 16 133-42
Radiation Damage World Records
MGy reaction reference
~45 global damage Owen et al. (2006)
10/Å global damage Howells et al. (2009)
5 Se-Met Holton (2007)
4 Hg-S Ramagopal et al. (2004)
4 R-C-COOH Garman et al. (2015)
3 S-S Murray et al. (2002)
1 Br-RNA Olieric et al. (2007)
~1? Cl-C ???
0.5 Mn in PS II Yano et al. (2005)
0.06 putidaredoxin Corbett et al. (2007)
0.02 Fe in myoglobin Denisov et al. (2007)

## Slide 253: Life-halving concentration
Na | 44
Mg | 27
P | 9
S | 7
Cl | 5
K | 3.3
Ca | 2.7
Fe | 1.0
Cu | 0.67
Zn | 0.58
molar, at the Se edge based on RADDOSE
As | 0.44
Se | 0.42
Br | 2.9
I | 0.47
Gd | 0.21
Ta | 0.13
Pt | 0.17
Au | 0.16
Hg | 0.16
U | 0.23

## Slide 254: Life-halving concentration
Na | 44
Mg | 27
P | 9
S | 7
Cl | 5
K | 3.3
Ca | 2.7
Fe | 1.0
Cu | 0.67
Zn | 0.58
molar, at the Se edge based on RADDOSE
As | 0.44
Se | 0.42
Br | 2.9
I | 0.47
Gd | 0.21
Ta | 0.13
Pt | 0.17
Au | 0.16
Hg | 0.16
U | 0.23

## Slide 255: No Damage in the Dark!
Experiment time (hours)
Peak absorbance (electrons)
4.6 5.0 5.4 5.8 6.2 6.6
0 5 10 15 20 25

## Slide 256
Where:
IDL - average damage-limited intensity (photons/hkl) at a given resolution
105 - converting R from μm to m, re from m to Å, ρ from g/cm3 to kg/m3 and MGy to Gy
re - classical electron radius (2.818 x 10-15 m/electron)
h - Planck’s constant (6.626 x 10-34 J∙s)
c - speed of light (299792458 m/s)
fdecayed - fractional progress toward completely faded spots at end of data set
ρ - density of crystal (~1.2 g/cm3)
R - radius of the spherical crystal (μm)
λ - X-ray wavelength (Å)
fNH - the Nave & Hill (2005) dose capture fraction (1 for large crystals)
nASU - number of proteins in the asymmetric unit
Mr - molecular weight of the protein (Daltons or g/mol)
VM - Matthews’s coefficient (~2.4 Å3/Dalton)
H - Howells’s criterion (10 MGy/Å)
θ - Bragg angle
a2 - number-averaged squared structure factor per protein atom (electron2)
Ma - number-averaged atomic weight of a protein atom (~7.1 Daltons)
B - average (Wilson) temperature factor (Å2)
μ - attenuation coefficient of sphere material (m-1)
μen - mass energy-absorption coefficient of sphere material (m-1)
Self-calibrated damage limit
Holton & Frankel (2010) Acta D 66 393-408.
No flux
No symmetry
No exposure time

## Slide 257
The number of photons scattered
before crystal is dead
is independent
of flux & time
1 um3 = 105 photons (roomT)
1 um3 = 106 photons (synch)
1 um3 = 108 photons (XFEL)
Henderson, 1990; Gonzalez & Nave, 1994; Glaeser et al., 2000; Sliz et al., 2003; Leiros et al., 2006; Owen et al., 2006; Garman & McSweeney, 2006; Garman & Nave, 2009; Holton, 2009
(synch)

## Slide 258: Dose slicing
crystal’s useful life
N
photons
N
photons
N
photons
unacceptable
completeness
unacceptable
read noise

## Slide 259: X-ray data are 3D!

## Slide 260: Data collection “strategy”
Dauter, Z. (1999). "Data-collection strategies." Acta Cryst. D 55(10): 1703-1717.
The “American Method”:
360° no matter what
(shoot first, questions later)

## Slide 261: Dose slicing: is that a spot?

## Slide 262: Dose slicing: is that a spot?

## Slide 263: Dose slicing: is that a spot?

## Slide 264: Dose slicing: is that a spot?

## Slide 265: Dose slicing: is that a spot?

## Slide 266: Dose slicing: is that a spot?

## Slide 267: What if?
You only have a few small crystals…
Should you:
  Collect 360° from each?
  Collect 10° from each at 36x exposure?
  Glue 36 xtals together, then collect 360° ?
  Glue, and do 12960° faint exposures?

## Slide 268
Chart: 
Resolution (Å)
CC vs right answer
“true” resolution limit

## Slide 269
Dose slice collection scenario |  |  | 
fine | coarse | glue | Glue & fine
3.17700 | 3.14057 | 3.17604 | 3.12086
“true” resolution vs strategy
301,640,334 photons/xtal

## Slide 270
processing procedure | resolution limit |  |  | 
 | fine | coarse | glue | glue&fine
XDS/XSCALE | 3.17700 | 3.14057 | 3.17604 | 3.12086
XDS/aimless | 3.16251 | 3.15922 | 3.16675 | 3.12162
XDS/noscale | 3.27289 | 3.15617 | 3.38929 | 3.12162
MOSFLM-defaults | x | x | x | x
MOSFLM-cheat | 3.19755 | 3.18166 | 3.20354 | 3.19755
HKL2000-defaults | x | x | x | x
HKL2000-expert | 3.21803 | 3.12646 | 3.11551 | 3.21803
HKL2000-cheat | 3.16119 | 3.10338 | 3.10338 | 3.16300
DIALS-defaults | x | x | 3.42125 | x
DIALS-expert | 3.10296 | 3.09547 | 3.07653 | 3.08465
d*Trek | ? | ? | 4.34076 | ?
EVAL15 | 3.14352 | 3.10383 | 3.09936 | 3.14105
EVAL15-sadabs | 3.13920 | 3.10797 | 3.10171 | 3.13700

## Slide 271: Minimum exposure time:
~1 photon/pixel
  PAD w XDS, DIALS
~10 photon/pixel
PAD w HKL2k, MOSFLM
~30 photon/pixel
CCD detector

## Slide 272: Optimal exposure time (faint spots on CCD or IP)
thr Optimal exposure time for data set (s)
tref exposure time of reference image (s)
bgref background level near weak spots on
reference image (ADU)
bg0 ADC offset of detector (ADU)
bghr optimal background level (via thr)
σ0 rms read-out noise (ADU)
gain ADU/photon
m multiplicity of data set (including partials)
adjust exposure
so this is ~15

## Slide 273: What Causes Disorder?
Function
Radiation damage
Mosaicity
Strain
Xtal abuse
Dehydration

## Slide 274: mosaic spread = 0 º

## Slide 275: mosaic spread = 0.1º

## Slide 276: mosaic spread = 0.2º

## Slide 277: mosaic spread = 0.4º

## Slide 278: mosaic spread = 0.6º

## Slide 279: mosaic spread = 0.8º

## Slide 280: mosaic spread = 1.0º

## Slide 281: mosaic spread = 1.5º

## Slide 282: mosaic spread = 2.0º

## Slide 283: mosaic spread = 2.5º

## Slide 284: mosaic spread = 3.2º

## Slide 285: mosaic spread = 6.4º

## Slide 286: mosaic spread = 12.8º

## Slide 287: What Causes Disorder?
Function
Radiation damage
Mosaicity
Strain
Xtal abuse
Dehydration

## Slide 288
stress and strain
radiation damage = defects
What about
undamaged crystals?

## Slide 289: Purity is crucial!
McPherson, A., Malkin, A. J., Kuznetsov, Y. G. & Plomp, M. (2001)."Atomic force microscopy applications in macromolecular crystallography", Acta Cryst. D 57, 1053-1060.
not important for
initial hits
important for
resolution

## Slide 290: What can I improve?
Purity!
is 95% good enough? 99%?
Purity!
conformational (homogeneous)
Purity!
kinetic (stable over time)

## Slide 291: What Causes Disorder?
Function
Radiation damage
Mosaicity (η) > ~3 μm
Strain (B) < ~3 μm
Xtal abuse
Dehydration

## Slide 292: What Causes Disorder?
Function
Radiation damage
Mosaicity > ~3 μm
Strain < ~3 μm
Xtal abuse
Dehydration

## Slide 293: Xtal Abuse
physical contact
don’t touch the part you intend to shoot
osmotic shock
equilibrate, or calculate matching solution
changes in dielectric constant
Petsko (1975) J. Mol. Biol. 96, 381-388.
cooled density mismatch
Juers & Matthews (2004) Acta Cryst. D 60, 412-421.
basically: no sudden moves!

## Slide 294: What Causes Disorder?
Function
Radiation damage
Mosaicity > ~3 μm
Strain < ~3 μm
Xtal abuse
Dehydration

## Slide 295: opened drop: you have ~30 seconds!
equilibrated drop

## Slide 296: OK! OK! Now I am afraid!
What should I do?

## Slide 297: Basic Principles
“Hell, there are NO RULES here - we're trying to accomplish something.”
Thomas A. Edison – inventor
“You’ve got to have an ASSAY.”
Arthur Kornberg – Nobel Laureate
“Control, control, you must learn CONTROL!”
Yoda – Jedi Master

## Slide 298: The Best Assay

## Slide 299: And the dominant source of error is…
Resolution
  Disorder (including Radiation Damage)
  Background
Phases
  Non-isomorphism (including Radiation Damage)
  Vibration (beam, xtal, spindle, etc.)
  Detector calibration

## Slide 300: Native Summary
100 ADU/pixel
10 μm for lysozyme
Do “strategy”
Follow by “American Method”
“Attenu-wait”: dose slicing
Assay, control and open mind
http://bl831.als.lbl.gov/xtalsize.html
http://bl831.als.lbl.gov/xtallife.html
http://bl831.als.lbl.gov/~jamesh/powerpoint/Uruguay_dtacoll_2017.pptx

## Slide 301: And the dominant source of error is…
Resolution
  Disorder (including Radiation Damage)
  Background
Phases
  Non-isomorphism (including Radiation Damage)
  Vibration (beam, xtal, spindle, etc.)
  Detector calibration

## Slide 302
anomalous signal
Crick, F. H. C. & Magdoff, B. S. (1956) Acta Crystallogr. 9, 901-908.
Hendrickson, W. A. & Teeter, M. M. (1981) Nature 290, 107-113.
# sites
MW (Da)
ΔF
F
≈ 1.2 f”
√
World record!
ΔF/F = 0.5%
Wang, Dauter & Dauter (2006) Acta Cryst. D 62, 1475-1483.

## Slide 303: Riso vs dose
Riso (%)
change in dose (MGy)
data taken from Banumathi, et al. (2004) Acta Cryst. D 60, 1085-1093.

## Slide 304: Riso vs dose
Riso (%)
change in dose (MGy)
data taken from Banumathi, et al. (2004) Acta Cryst. D 60, 1085-1093.
Riso ≈ 0.7 %/MGy

## Slide 305: First diffraction from protein xtal
Bernal, J. & Crowfoot, D. (1934). "X-ray photographs of crystalline pepsin", Nature 133, 794-795.
1934

## Slide 306: Dehydration: 1934
100 μm
2 mm
= 1.5 μL
= 1.0 nL
10 μm
= 1.0 pL
and 2017?
1 μm
= 1.0 fL

## Slide 307
Dear James
The story of the two forms of lysozyme crystals goes back to about 1964 when it was found that the diffraction patterns from different crystals could be placed in one of two classes depending on their intensities. This discovery was a big set back at the time and I can remember a lecture title being changed from the 'The structure of lysozyme' to 'The structure of lysozyme two steps forward and one step back'. Thereafter the crystals were screened based on intensities of the (11,11,l) rows to distinguish them (e.g. 11,11,4 > 11,11,5 in one form and vice versa in another). Data were collected only for those that fulfilled the Type II criteria. (These reflections were easy to measure on the linear diffractometer because crystals were mounted to rotate about the diagonal axis). As I recall both Type I and Type II could be found in the same crystallisation batch . Although sometimes the external morphology allowed recognition this was not infallible.
The structure was based on Type II crystals. Later a graduate student Helen Handoll examined Type I. The work, which was in the early days and before refinement programmes, seemed to suggest that the differences lay in the arrangement of water or chloride molecules (Lysozyme was crystallised from NaCl). But the work was never written up. Keith Wilson at one stage was following this up as lysozyme was being used to test data collection strategies but I do not know the outcome.
An account of this is given in International Table Volume F (Rossmann and Arnold edited 2001) p760.
Tony North was much involved in sorting this out and if you wanted more info he would be the person to contact. I hope this is helpful. Do let me know if you need more.
Best wishes
Louise
Non-isomorphism in lysozyme

## Slide 308
Johnson’s ratio
Structure factor (e-)
Non-isomorphism in lysozyme

## Slide 309
Johnson’s ratio
Structure factor (e-)
Non-isomorphism in lysozyme

## Slide 311: Non-isomorphism in lysozyme
RH: 84.2%
vs 71.9%
Riso = 44.5%
3aw6
3aw7
Δcell = 0.7 %
RMSD = 1.70 Å
RMSD = 0.18 Å

## Slide 312
0.5% cell change → 15%
Crick & Magdoff (1956) “theory method isomorphous replacement”, Acta Cryst. 9, 901-8.
Same structure, different cell

## Slide 313
Same structure, different cell
0.5% cell change → 15%
Crick & Magdoff (1956) “theory method isomorphous replacement”, Acta Cryst. 9, 901-8.

## Slide 314
Same cell, rotated protein
0.5° rotation → 16%
Crick & Magdoff (1956) “theory method isomorphous replacement”, Acta Cryst. 9, 901-8.

## Slide 315
Same cell, rotated protein
0.5° rotation → 16%
Crick & Magdoff (1956) “theory method isomorphous replacement”, Acta Cryst. 9, 901-8.

## Slide 316
Uniform stretch
5% change → 0%

## Slide 317
5% change → 0%
Uniform stretch

## Slide 318
Magdoff & Crick (1955) “Ribonuclease II. Accuracy of measurement and shrinkage”, Acta Cryst. 8, 461-8.
Crick & Magdoff (1956) “The theory of the method of isomorphous replacement for protein crystals”, Acta Cryst. 9, 901-8.

## Slide 319: Elastic deformation = non-isomorphism

## Slide 320: Elastic deformation = non-isomorphism

## Slide 321: Plastic deformation = poor diffraction

## Slide 322: Plastic deformation = poor diffraction

## Slide 323: … but does it work for membrane proteins?
Efremov, R. G. & Sazanov, L. A. (2011)."Structure of the membrane domain of respiratory complex I", Nature 476, 414-420.
“To improve diffraction properties, crystals were slowly cooled to 4 C, placed into micro-dialysis buttons, and mother liquor was exchanged by dialysis for 0.1 M sodium acetate (pH 4.8), 0.15 M sodium tartrate (pH 5.0), 30% PEG4000, 0.1% sodium cholate and 12.5% glycerol over a period of 5 days.”
CCP4BB:
“We used it for a large membrane protein complex and diffraction improved from ~7 to 2.7 A. The crystal is fished out and put into mother liquor solution in the button, sealed with dialysis membrane and the button is then placed into about 5 mls of mother liquor with slightly higher PEG concentration. Then you just exchange outside buffer every day or so for solutions containing higher concentrations of PEG. We went from ~9 to 30 % PEG4000 in about a week. You can easily observe crystal under microscope and if it cracks - you went too far/too quickly with PEG and need to use a bit less next time. Also, this method allows you to control all other components of the dehydrating solution - we needed to decrease salt concentration at the same time as increasing PEG. You can also introduce/increase cryo-protectant concentration at the same time. With these crystals, otherwise excellent dehydration machines already mentioned did not work, possibly because the process had to be really slow.”

## Slide 324: And the dominant source of error is…
Resolution
  Disorder (including Radiation Damage)
  Background
Phases
  Non-isomorphism (including Radiation Damage)
  Vibration (beam, xtal, spindle, etc.)
  Detector calibration

## Slide 325: Phasing Summary
~3% error per spot, 1%/MGy
7235 eV for S-SAD
Multiplicity multiplicity multiplicity
“Attenu-wait”: dose slicing
>1000 photons/spot is a waste
http://bl831.als.lbl.gov/xtalsize.html
http://bl831.als.lbl.gov/xtallife.html
http://bl831.als.lbl.gov/~jamesh/powerpoint/Uruguay_datacoll_2017.pptx

## Slide 326: But my reviewer says…
Rmerge in high-resolution bin is >100% !
Answer: this is expected

## Slide 327: Expected Rmerge as Iobs → 0

## Slide 328
Average value
Averaging Gaussian error
Number of samples averaged

## Slide 329
Average value
Averaging Gauss/Gauss error
Number of samples averaged

## Slide 330
Average outer-shell Rmerge
Rmerge at the resolution limit in PDB
Year

## Slide 331: Take-home lesson:
R factors are
undefined
as Iobs → 0
Report as “ – “ in outer bin

## Slide 332: Re-indexing cheat sheet
Dauter, Z. (1999). "Data-collection strategies." Acta Cryst. D 55(10): 1703-1717.

## Slide 333: CCP4: aimless log
$TABLE: Analysis against resolution, XDSdataset:
$GRAPHS:I/sigma, Mean Mn(I)/sd(Mn(I)):0|0.216023x0|137.14:2,13,14:
:Rmerge, Rfull, Rmeas, Rpim v Resolution:0|0.216023x0|1.70834:2,4,5,6,7:
:Average I, RMSdeviation and Sd:0|0.216023x0|1650.8:2,10,11,12:
:Fractional bias:0|0.216023x0|0:2,15:
$$
N 1/d^2 Dmid Rmrg Rfull Rcum Rmeas Rpim Nmeas AvI RMSdev sd I/RMS Mn(I/sd) FrcBias $$ $$
1 0.0064 12.55 0.020 0.020 0.020 0.021 0.006 13115 1651 57 42 29.2 137.1 -
2 0.0191 7.24 0.027 0.027 0.024 0.028 0.008 24753 1171 47 42 25.0 105.2 -
3 0.0318 5.61 0.038 0.038 0.029 0.040 0.012 32197 857 46 43 18.4 79.6 -
4 0.0445 4.74 0.034 0.034 0.031 0.035 0.010 37743 1212 57 53 21.4 91.2 -
5 0.0572 4.18 0.036 0.036 0.032 0.038 0.011 42642 1181 59 57 19.9 83.8 -
6 0.0699 3.78 0.049 0.049 0.036 0.052 0.015 47224 883 59 57 15.1 65.1 -
7 0.0826 3.48 0.065 0.065 0.040 0.068 0.020 51052 685 59 58 11.7 50.9 -
8 0.0953 3.24 0.096 0.096 0.045 0.100 0.029 54636 448 56 56 8.0 35.0 -
9 0.1080 3.04 0.151 0.151 0.050 0.158 0.046 58072 268 53 53 5.1 22.7 -
10 0.1207 2.88 0.229 0.229 0.056 0.240 0.070 60731 171 51 51 3.3 15.4 -
11 0.1334 2.74 0.314 0.314 0.063 0.329 0.097 63807 125 51 51 2.4 11.3 -
12 0.1461 2.62 0.406 0.406 0.070 0.425 0.125 66241 98 51 52 1.9 8.7 -
13 0.1588 2.51 0.537 0.537 0.078 0.562 0.166 68272 76 53 53 1.4 6.5 -
14 0.1715 2.41 0.685 0.685 0.084 0.727 0.237 54170 61 53 54 1.1 4.4 -
15 0.1843 2.33 0.886 0.886 0.088 0.950 0.333 39781 48 55 56 0.9 3.1 -
16 0.1970 2.25 1.249 1.249 0.092 1.355 0.508 29754 35 57 59 0.6 2.1 -

## Slide 334: CCP4: aimless log
$TABLE: Correlations CC(1/2) within dataset, XDSdataset:
$GRAPHS: Anom & Imean CCs v resolution:0|0.216023x0|1:2,4,7:
: RMS correlation ratio :0|0.216023x0|2.20344:2,6:
$$
N 1/d^2 Dmid CCanom Nanom RCRanom CC1/2 NImean $$ $$
1 0.0064 12.55 0.659 499 2.203 1.000 669
2 0.0191 7.24 0.550 975 1.853 1.000 1155
3 0.0318 5.61 0.527 1295 1.798 1.000 1479
16 0.1970 2.25 0.037 2123 1.038 0.711 2275
17 0.2097 2.18 0.043 1682 1.044 0.460 1877

## Slide 335: XDS: CORRECT.LP or XSCALE.LP
SUBSET OF INTENSITY DATA WITH SIGNAL/NOISE >= -3.0 AS FUNCTION OF RESOLUTION
RESOLUTION NUMBER OF REFLECTIONS COMPLETENESS R-FACTOR R-FACTOR COMPARED I/SIGMA R-meas CC(1/2) Anomal SigAno Nano
LIMIT OBSERVED UNIQUE POSSIBLE OF DATA observed expected Corr
6.39 4103 2517 3135 80.3% 1.7% 1.8% 3018 33.77 2.3% 99.9* 21* 1.012 533
4.54 7339 4970 5544 89.6% 2.6% 2.6% 4585 22.56 3.6% 99.8* 9 0.914 796
3.72 9442 6724 7186 93.6% 2.8% 2.9% 5327 19.99 4.0% 99.7* 9 0.859 877
3.22 11077 7997 8486 94.2% 4.8% 4.9% 6094 12.27 6.8% 99.3* -1 0.784 913
2.88 12408 8851 9595 92.2% 10.0% 10.1% 7068 6.01 14.2% 97.8* -2 0.799 1012
2.63 13673 9561 10662 89.7% 20.8% 20.8% 8185 3.10 29.4% 88.8* -4 0.776 1274
2.44 14360 9852 11510 85.6% 34.5% 34.7% 8981 1.90 48.8% 75.9* 2 0.765 1525
2.28 9501 6495 12416 52.3% 61.8% 59.6% 5991 1.14 87.3% 53.5* -2 0.722 1031
2.15 4567 3307 13210 25.0% 120.5% 120.3% 2520 0.59 170.4% 21.9* 4 0.693 408
total 86470 60274 81744 73.7% 6.7% 6.7% 51769 8.97 9.5% 99.5* 2 0.804 8369

## Slide 336: XDS: CORRECT.LP or XSCALE.LP
******************************************************************************
CORRECTION PARAMETERS FOR THE STANDARD ERROR OF REFLECTION INTENSITIES
******************************************************************************
The variance v0(I) of the intensity I obtained from counting statistics is
replaced by v(I)=a*(v0(I)+b*I^2). The model parameters a, b are chosen to
minimize the discrepancies between v(I) and the variance estimated from
sample statistics of symmetry related reflections. This model implicates
an asymptotic limit ISa=1/SQRT(a*b) for the highest I/Sigma(I) that the
experimental setup can produce (Diederichs (2010) Acta Cryst D66, 733-740).
a b ISa
3.806E+00 1.080E-04 49.32

## Slide 337: Hygrostatic gradients
loop
high
water activity
low

## Slide 338
1% high
average
1% low
Relative Pixel
Calibration
ALS 8.3.1
Pilatus3 6M
0.9793 Å x-rays
RMS = 0.8%