## Slide 1: Acknowledgements
http://bl831.als.lbl.gov/~jamesh/powerpoint/CSHL_beamline_2023.pptx
UCSF LBNL SLAC
ALS 8.3.1: TomAlberTron
NIH NIGMS R01 GM124149 to Holton
NIH NIAID P50 AI150476 to Krogan
NIH NIAID U19 AI171110 to Krogan
NIH NIGMS P30 GM124169 to Adams
DOE-BER IDAT to Hura
NIH NIGMS P41 GM103393 to Hodgson
The Advanced Light Source is supported by the Director, Office of Science, Office of Basic Energy Sciences, Materials Sciences Division, of the US Department of Energy under contract No. DE-AC02-05CH11231 at Lawrence Berkeley National Laboratory.
Robert Stroud James Fraser
Paul Adams Greg Hura
Aina Cohen Keith Hodgson

## Slide 3: PowerPoint File available:
http://bl831.als.lbl.gov/
~jamesh/powerpoint/
CSHL_beamline_2023.pptx

## Slide 4: The “success rate” of data collection
100 s/dataset
200 days/year
~150 beamlines
~26,000,000 datasets/year
~14,000 PDBs in 2022

## Slide 5: signal vs noise
“If you don’t have good data, then you have no data at all.”
-Sung-Hou Kim

## Slide 6: signal vs noise
easy
hard
impossible

## Slide 7: signal vs noise
easy
hard
impossible
threshold of “solvability”

## Slide 8: signal vs noise
“If you don’t have good data, then you must learn statistics.”
-James Holton

## Slide 9: Adding noise

## Slide 10: Adding noise
12 + 12 = 1.42

## Slide 11: Adding noise
12 + 12 = 1.42
32 + 12 = 3.22
σtotal2 = σ12 + σ22

## Slide 12: Adding noise
12 + 12 = 1.42
32 + 12 = 3.22
σtotal2 = σ12 + σ22

## Slide 13: Adding noise
12 + 12 = 1.42
32 + 12 = 3.22
σtotal2 = σ12 + σ22

## Slide 14: Adding noise
12 + 12 = 1.42
32 + 12 = 3.22
102 + 12 = 10.052

## Slide 15: And the dominant source of error is…
Resolution
  Disorder (including Radiation Damage)
  Background
Phases
  Non-isomorphism (including Radiation Damage)
  Vibration (beam, xtal, spindle, etc.)
  Detector calibration

## Slide 16: Data quality - simple rules:
  Put your crystal into the beam

## Slide 17: Your sample is a lens! (but not for x-rays)
90°

## Slide 18: Your sample is a lens! (but not for x-rays)
90°

## Slide 19: Data quality - simple rules:
  Put your crystal into the beam
  Shoot the whole crystal

## Slide 20: shoot the whole crystal

## Slide 21: shoot the whole crystal

## Slide 22: shoot the whole crystal

## Slide 23: shoot the whole crystal

## Slide 24: shoot the whole crystal

## Slide 25: How many “crystals” do you see?
Shoot the crystal (singular)

## Slide 26: Data quality - simple rules:
  Put your crystal into the beam
  Shoot the whole crystal
  Shoot nothing but the crystal

## Slide 27: shoot nothing but the crystal

## Slide 28: shoot nothing but the crystal

## Slide 29: scattering/absorptoin “rules”:
1 μm crystal ≈ 1 μm water
≈ 1 μm plastic
≈ 0.1 μm glass
≈ 1000 μm air
Optimum: support thickness = crystal thickness
1 + 1 = 1.4

## Slide 30
X-ray beam size (μm)
Resolution (Å)
too big
too small
just right
limited by
radiation
damage
limited by
excess
background

## Slide 31: platy crystals

## Slide 32: platy crystals

## Slide 33: Data quality - simple rules:
  Put your crystal into the beam
  Shoot the whole crystal
  Shoot nothing but the crystal
  The crystal must rotate

## Slide 34: “pretty” crystals

## Slide 35: “pretty” crystals

## Slide 36: “pretty” crystals

## Slide 37: The crystal rotates!

## Slide 38: The crystal rotates!

## Slide 39: avoiding overlaps
c
c

## Slide 40: avoiding overlaps
c

## Slide 41: avoiding overlaps
c
http://www.mitegen.com/mic_catalog.php?c=DTMicroMounts

## Slide 42: avoiding overlaps
c
http://www.mitegen.com/mic_catalog.php?c=orientedmounts

## Slide 43: Data quality - simple rules:
  Put your crystal into the beam
  Shoot the whole crystal
  Shoot nothing but the crystal
  The crystal must rotate
  Back off!

## Slide 44: Background scattering

## Slide 45: Mosquito tips cost $0.15 each
So do detector pixels!
We bought 6,224,001
Use them!
(they are reusable)

## Slide 46: Background scattering
$100,000.00
$100,000.00
$100,000.00
$100,000.00
$100,000.00
$100,000.00
$100,000.00
$100,000.00
real estate is
expensive
use it!

## Slide 47: Fine Slicing
Pflugrath, J. W. (1999)."The finer things in X-ray diffraction data collection", Acta Cryst. D 55, 1718-1725.
background
background

## Slide 49: Optimal exposure time (anomalous on a PAD)
tmad Optimal exposure time (s)
tro read-out time (s)
2.03 ms on Pilatus3 S
0.95 ms on Pilatus3 X
0.003 ms on Eiger
Attenuate for MAD
PILATUS
>0.2 s/frame
EIGER
>0.0003 s/frame

## Slide 50: Optimal exposure time (anomalous on a PAD)
time
intensity
2.03 ms

## Slide 51: Optimal exposure time (faint spots on CCD detector)
thr Optimal exposure time for data set (s)
tref exposure time of reference image (s)
bgref background level near weak spots on
reference image (ADU)
bg0 ADC offset of detector (ADU)
bghr optimal background level (via thr)
σ0 rms read-out noise (ADU)
gain ADU/photon
m multiplicity of data set (including partials)
adjust exposure
so this is ~100

## Slide 52: And the dominant source of error is…
Resolution
  Disorder (including Radiation Damage)
  Background
Phases
  Non-isomorphism (including Radiation Damage)
  Vibration (beam, xtal, spindle, etc.)
  Detector calibration

## Slide 53
Pick-up tool mark
Bragg
glitch
oxygen
inclusions
“tree rings”
1% high
average
1% low
~3x105 photon/pixel Pilatus detector
Spatial Heterogeneity in Sharp Spot Sensitivity

## Slide 54: Detector calibration: 7235 eV

## Slide 55: Detector calibration: 7247 eV

## Slide 56: Gadox calibration vs energy
photon energy (keV)
Relative absorption depth
same = good!
bad!

## Slide 57: Systematic error does not “average out”
Unless you can randomize their source

## Slide 58: Fractional error
mult > (—)2
~3%
<ΔF/F>

## Slide 59
anomalous signal
Crick, F. H. C. & Magdoff, B. S. (1956) Acta Crystallogr. 9, 901-908.
Hendrickson, W. A. & Teeter, M. M. (1981) Nature 290, 107-113.
# sites
MW (Da)
ΔF
F
≈ 1.2 f”
√
World record!
ΔF/F = 0.5%
Wang, Dauter & Dauter (2006) Acta Cryst. D 62, 1475-1483.

## Slide 60
Holton & Frankel (2010) Acta D 66 393-408.

## Slide 61: 100% SeMet incorporation
Trivial to solve

## Slide 62: 100% S -Met incorporation
Impossible to solve

## Slide 63: Phase Problem: Se vs S
fraction not Selenium
correlation coefficient

## Slide 64: Phase Problem: Se vs S
fraction not Selenium
correlation coefficient

## Slide 65: Phase Problem: Se vs S
fraction not Selenium
correlation coefficient

## Slide 66: Phase Problem: Se vs S
fraction not Selenium
correlation coefficient

## Slide 67: Phase Problem: Se vs S
fraction not Selenium
correlation coefficient

## Slide 68: One xtal, no idea? Suggested native protocol:
Do strategy
Start at specified phi
100% fast as you can
Keep going to 360°
Move detector
Advance phi 10-30 deg
Increase photons/image 4x
goto 4

## Slide 69: Suggested anomalous protocol:
2 wavelengths are better than 1
- (peak + inf)/2, and remote
MAD, not M-SAD!
360° in < ~5 MGy (for Se)
move detector
4X exposure
goto 2

## Slide 70: The “threshold of solvability” is sharp!
How can we predict success/failure?
Know Thy Experiment

## Slide 71: The “beam line”
Area Detector
Protein Crystal
Cite the “beamline paper!”
X-ray
Source
Notes: Most notable acheivements are high brightness with very little “wasted” light. The discovery made by Howard Padmore? At ALS showed that a 2:1 demagnification cancels nearly all of the spherical abberations in the parabolic/torroid system. 30-50% of the total x-ray intensity passes through the 100um pinhole. The flux through the 100um pinhole at the superbend beamlines is comparable to that of the 5.0.2 wiggler or an APS bending magnet, but only 150 Watts of power must be dissipated by the beamline optics. In contrast, wiggler beamlines must disipate kilowatts. The optics here are much simpler and not exotic in any way. The system is stable and robust. There is never a need to re-focus the beam at a new wavelength. Only a 2-minute automatic tuneup is recommended for very large energy changes (more than 1-2 keV).

## Slide 72: The “beam line”
Area Detector
Source
Primary
mirror
Secondary
mirror
Si(111)
monochromator
Sample
focused
beam
slits
Scatter
guard
Unfocused
beam
slits

## Slide 73: The “beam line”
Area Detector
Source
Primary
mirror
Secondary
mirror
Si(111)
monochromator
Sample
focused
beam
slits
Scatter
guard
Unfocused
beam
slits
Notes: Most notable acheivements are high brightness with very little “wasted” light. The discovery made by Howard Padmore? At ALS showed that a 2:1 demagnification cancels nearly all of the spherical abberations in the parabolic/torroid system. 30-50% of the total x-ray intensity passes through the 100um pinhole. The flux through the 100um pinhole at the superbend beamlines is comparable to that of the 5.0.2 wiggler or an APS bending magnet, but only 150 Watts of power must be dissipated by the beamline optics. In contrast, wiggler beamlines must disipate kilowatts. The optics here are much simpler and not exotic in any way. The system is stable and robust. There is never a need to re-focus the beam at a new wavelength. Only a 2-minute automatic tuneup is recommended for very large energy changes (more than 1-2 keV).

## Slide 74: Where do the X-rays come from?

## Slide 75: The “Home source”
e-

## Slide 76: The “Home source”
e-
+
e-

## Slide 77: The “Home source”
+

## Slide 78: The “Home source”
e-
+

## Slide 79: The “Home source”

## Slide 80: Metal identification
scattering
same
λ as
incident

## Slide 81
“MAD scan” – absorption spectrum
“counts” (aka arbitrary units)
0.980888
0.980500
0.980113
0.979725
0.979338
0.978952
0.978565
0.978179
0.977794
0.977408
0.977023
0.976638
0.976254
Wavelength (Å)
fluorescence = yield * absorption
12398.42
energy

## Slide 82
maximize
absolute
value of
f’’
maximize
difference
in f’
“MAD scan” – absorption spectrum

## Slide 83
4-wavelength MAD

## Slide 84
3-wavelength MAD

## Slide 85
2-wavelength MAD

## Slide 86
2-wavelength MAD

## Slide 87
1-wavelength SAD

## Slide 88
2-wavelength metal validation
phased
anomalous
difference
Fourier

## Slide 89: The “beam line”
Area Detector
Source
Primary
mirror
Secondary
mirror
Si(111)
monochromator
Sample
focused
beam
slits
Scatter
guard
Unfocused
beam
slits
Notes: Most notable acheivements are high brightness with very little “wasted” light. The discovery made by Howard Padmore? At ALS showed that a 2:1 demagnification cancels nearly all of the spherical abberations in the parabolic/torroid system. 30-50% of the total x-ray intensity passes through the 100um pinhole. The flux through the 100um pinhole at the superbend beamlines is comparable to that of the 5.0.2 wiggler or an APS bending magnet, but only 150 Watts of power must be dissipated by the beamline optics. In contrast, wiggler beamlines must disipate kilowatts. The optics here are much simpler and not exotic in any way. The system is stable and robust. There is never a need to re-focus the beam at a new wavelength. Only a 2-minute automatic tuneup is recommended for very large energy changes (more than 1-2 keV).

## Slide 90: Electric field of a moving charge

## Slide 91: Electric field of a moving charge
accelerating charges
make
electromagnetic waves
(light)

## Slide 92: Bending Magnet
N
electron beam
X-rays !
S

## Slide 93: Two Bending Magnets
N
electron beam
S
S
N
2x X-rays !

## Slide 94: Wiggler
N
S
N
S
N
S
N
S
S
N
S
N
S
N
S
N
8x X-rays !

## Slide 95: Undulator
N
S
N
S
N
S
N
S
S
N
S
N
S
N
S
N
> 8x X-rays !

## Slide 96: Undulator emission spectrum

## Slide 97: XFEL Undulator
N
S
N
S
N
S
N
S
S
N
S
N
S
N
S
N
same X-rays
short pulse
S
N
S
N
S
N
S
N
N
S
N
S
N
S
N
S

## Slide 98: LCLS undulator hall

## Slide 99: LCLS at SLAC

## Slide 100: Accelerator-based light source terminology
Code name Translation
First Generation Bending Magnet
Second Generation Wiggler
Third Generation Undulator
Fourth Generation Free electron laser
..and diffraction-limited synchrotron

## Slide 101: Diffraction-limited light source
λ ≈ 2 d sinθ
λ = x-ray wavelength
d = electron/x-ray beam size
θ = beam divergence
d = 1 Å / 100 µrad / 2 = 0.5 µm

## Slide 102: The “beam line”
Area Detector
Source
Primary
mirror
Secondary
mirror
Si(111)
monochromator
Sample
focused
beam
slits
Scatter
guard
Unfocused
beam
slits
Notes: Most notable acheivements are high brightness with very little “wasted” light. The discovery made by Howard Padmore? At ALS showed that a 2:1 demagnification cancels nearly all of the spherical abberations in the parabolic/torroid system. 30-50% of the total x-ray intensity passes through the 100um pinhole. The flux through the 100um pinhole at the superbend beamlines is comparable to that of the 5.0.2 wiggler or an APS bending magnet, but only 150 Watts of power must be dissipated by the beamline optics. In contrast, wiggler beamlines must disipate kilowatts. The optics here are much simpler and not exotic in any way. The system is stable and robust. There is never a need to re-focus the beam at a new wavelength. Only a 2-minute automatic tuneup is recommended for very large energy changes (more than 1-2 keV).

## Slide 103: X-ray optics: mirrors
Notes: Next slide please

## Slide 104: Why are the mirrors so long?
3 mm
Must be < 0.2°
1000 mm
Notes: Next slide please

## Slide 105: There is a lens for X-rays!
1 mm
0.1 mm
Be
Notes: Next slide please

## Slide 106: The “beam line”
Area Detector
Source
Primary
mirror
Secondary
mirror
Si(111)
monochromator
Sample
focused
beam
slits
Scatter
guard
Unfocused
beam
slits
Notes: Most notable acheivements are high brightness with very little “wasted” light. The discovery made by Howard Padmore At ALS showed that a 2:1 demagnification cancels nearly all of the spherical abberations in the parabolic/torroid system. 30-50% of the total x-ray intensity passes through the 100um pinhole. The flux through the 100um pinhole at the superbend beamlines is comparable to that of the 5.0.2 wiggler or an APS bending magnet, but only 150 Watts of power must be dissipated by the beamline optics. In contrast, wiggler beamlines must disipate kilowatts. The optics here are much simpler and not exotic in any way. The system is stable and robust. There is never a need to re-focus the beam at a new wavelength. Only a 2-minute automatic tuneup is recommended for very large energy changes (more than 1-2 keV).

## Slide 107
Single crystal

## Slide 108: Monochromators
λ = 1 Å
Si
Notes: Next slide please

## Slide 109: Monochromators
λ = 1 Å
Si
d = 3.135 Å
θ = 9.177°
λ = 1 Å
nλ = 2d sin(θ)
Notes: Next slide please

## Slide 110: Monochromators
nλ = 2d sin(θ)
λ = 1.1 Å
Si
d = 3.135 Å
θ = 9.177°
Notes: Next slide please

## Slide 111: Monochromators
nλ = 2d sin(θ)
λ = 1.1 Å
Si
d = 3.135 Å
θ = 10.104°
λ = 1.1 Å
Notes: Next slide please

## Slide 112: Monochromators
nλ = 2d sin(θ)
λ = 1 Å
Si
d = 3.135 Å
θ = 9.177°
λ = 1 Å
Notes: Next slide please

## Slide 113: Monochromators
Si
λ = 1 Å
λ = 1 Å
Si
d = 3.135 Å
DCM = Double Crystal Monochromator
Notes: Next slide please

## Slide 114: Monochromators
Si
λ = 1.1 Å
λ = 1.1 Å
Si
d = 3.135 Å
DCM = Double Crystal Monochromator
Notes: Next slide please

## Slide 115: Monochromators
nλ = 2d sin(θ)
λ = 1.54 Å
θ = 1.84°
λ = 1.54 Å
Mo/C
d = 24 Å
Multilayer optics
Notes: Next slide please

## Slide 116: The truth about x-ray beams
Term units significance
Flux photons/s duration of experiment
Beam Size μm match to crystal
Divergence mrad spot size vs distance
Wavelength Å resolution and absorption
Dispersion Δλ/λ spot size
Flux density ph/s/area scattering/damage rate
Fluence ph/area scattering/damage

## Slide 117: The truth about x-ray beams
Term units significance
Flux photons/s duration of experiment
Beam Size μm match to crystal
Divergence mrad spot size vs distance
Wavelength Å resolution and absorption
Dispersion Δλ/λ spot size
Flux density ph/s/area scattering/damage rate
Fluence ph/area scattering/damage

## Slide 118
The number of photons scattered
before crystal is dead
is independent
of data collection time
1 um3 = 106 photons
Henderson, 1990; Gonzalez & Nave, 1994; Glaeser et al., 2000; Sliz et al., 2003; Leiros et al., 2006; Owen et al., 2006; Garman & McSweeney, 2006; Garman & Nave, 2009; Holton, 2009

## Slide 119: The truth about x-ray beams
quantity | units | home source | synchrotron
flux | Photons/second | 2.5 x 109 | 2 x 1012
exposure | seconds | 400 | 1

## Slide 120: The truth about x-ray beams
quantity | units | home source | synchrotron
flux | Photons/second | 2.5 x 109 | 2 x 1012
exposure | seconds | 400 | 1
Dispersion | wavelength range / wavelength | 0.2% (Kα1 - Kα2) | 0.014% (Si111)
Divergence | milliRadian | 4.8 | 1.8 (h) 0.18 (v)
Beam size | microns | 100 | 24 (h) 12 (v)
Spectral brightness | Photons/s/mm2/ mR2/0.1%BW | 5.4 x 109 | 1.5 x 1017

## Slide 121: The truth about x-ray beams
quantity | units | home source | synchrotron
flux | Photons/second | 72 x 106 | 2 x 1012
exposure | time | 4 hours | 1 second
Dispersion | wavelength range / wavelength | 0.2% (Kα1 - Kα2) | 0.014% (Si111)
Divergence | milliRadian | 4.8 | 1.8 (h) 0.18 (v)
Beam size | microns | 17 | 24 (h) 12 (v)
Spectral brightness | Photons/s/mm2/ mR2/0.1%BW | 5.4 x 109 | 1.5 x 1017

## Slide 122: The truth about x-ray beams
quantity | units | home source | synchrotron
flux | Photons/second | 1 x 106 | 2 x 1012
exposure | time | 10 days | 1 second
Dispersion | wavelength range / wavelength | 0.2% (Kα1 - Kα2) | 0.014% (Si111)
Divergence | milliRadian | 0.6 | 1.8 (h) 0.18 (v)
Beam size | microns | 17 | 24 (h) 12 (v)
Spectral brightness | Photons/s/mm2/ mR2/0.1%BW | 5.4 x 109 | 1.5 x 1017

## Slide 123: The truth about x-ray beams
quantity | units | home source | synchrotron
flux | Photons/second | 8 x 104 | 2 x 1012
exposure | time | 5 months | 1 second
Dispersion | wavelength range / wavelength | 0.014% (Si111) | 0.014% (Si111)
Divergence | milliRadian | 0.6 | 1.8 (h) 0.18 (v)
Beam size | microns | 17 | 24 (h) 12 (v)
Spectral brightness | Photons/s/mm2/ mR2/0.1%BW | 5.4 x 109 | 1.5 x 1017

## Slide 124: The truth about x-ray beams
Term units significance
Flux photons/s duration of experiment
Beam Size μm match to crystal
Divergence mrad spot size vs distance
Wavelength Å resolution and absorption
Dispersion Δλ/λ spot size
Flux density ph/s/area scattering/damage rate
Fluence ph/area scattering/damage

## Slide 125: Out of time

## Slide 126
Two-color experiment (ca 1914)
(Cu source)
20 eV

## Slide 127: spectral dispersion
Ewald sphere
spindle axis
Φ circle
diffracted ray
(h,k,l)
d*
λ*
λ*

## Slide 128: spectral dispersion
Ewald sphere
spindle axis
Φ circle
diffracted ray
(h,k,l)
d*
λ’*
λ’*

## Slide 129: dispersion = 0

## Slide 130: dispersion = 0.014% (Si 111)

## Slide 131: dispersion = 0.25% (CuKα)

## Slide 132: dispersion = 0.6%

## Slide 133: dispersion = 1.3%

## Slide 134: dispersion = 2.6%

## Slide 135: dispersion = 5.1%

## Slide 136: The truth about x-ray beams
Term units significance
Flux photons/s duration of experiment
Beam Size μm match to crystal
Divergence mrad spot size vs distance
Wavelength Å resolution and absorption
Dispersion Δλ/λ spot size
Flux density ph/s/area scattering/damage rate
Fluence ph/area scattering/damage

## Slide 137: beam divergence
Ewald sphere
spindle axis
Φ circle
diffracted ray
(h,k,l)
d*
λ*
λ*

## Slide 138: beam divergence
spindle axis
Φ circle
diffracted ray
(h,k,l)
d*
Ewald sphere
λ*
λ*

## Slide 139: divergence = 0 º

## Slide 140: divergence = 0.3 º

## Slide 141: mosaic spread
Ewald sphere
spindle axis
Φ circle
diffracted ray
(h,k,l)
λ*
λ*
d*
d*

## Slide 142: mosaic spread
Ewald sphere
spindle axis
Φ circle
diffracted rays
(h,k,l)
d*
d*

## Slide 143: The truth about x-ray beams
Term units significance
Flux photons/s duration of experiment
Beam Size μm match to crystal
Divergence mrad spot size vs distance
Wavelength Å resolution and absorption
Dispersion Δλ/λ spot size
Flux density ph/s/area scattering/damage rate
Fluence ph/area scattering/damage

## Slide 144: Air absorption
Distance from sample (mm)
transmittance

## Slide 145: Summary
1 + 1 = 1.4
Match beam to “crystal”
Match goop thickness to crystal
Better beam = less time
Look at your spots!
“true multiplicity” for anomalous
Crystal must rotate

## Slide 146: Correct beam center
155.063 155.647
159.353 155.063
159.301 155.011
1512.73 1554.57
155.063 159.356 -250
ADXV
HKL2000
MOSFLM
XDS
DIALS

## Slide 147: There are 16 different “beam centers”
Origin
Beam X
Beam Y

## Slide 148: There are 16 different “beam centers”
Origin
Beam X
Beam Y
Cameraman
View
Beam
View

## Slide 149: Re-indexing cheat sheet
Dauter, Z. (1999). "Data-collection strategies." Acta Cryst. D 55(10): 1703-1717.

## Slide 150: CCP4: aimless log
$TABLE: Analysis against resolution, XDSdataset:
$GRAPHS:I/sigma, Mean Mn(I)/sd(Mn(I)):0|0.216023x0|137.14:2,13,14:
:Rmerge, Rfull, Rmeas, Rpim v Resolution:0|0.216023x0|1.70834:2,4,5,6,7:
:Average I, RMSdeviation and Sd:0|0.216023x0|1650.8:2,10,11,12:
:Fractional bias:0|0.216023x0|0:2,15:
$$
N 1/d^2 Dmid Rmrg Rfull Rcum Rmeas Rpim Nmeas AvI RMSdev sd I/RMS Mn(I/sd) FrcBias $$ $$
1 0.0064 12.55 0.020 0.020 0.020 0.021 0.006 13115 1651 57 42 29.2 137.1 -
2 0.0191 7.24 0.027 0.027 0.024 0.028 0.008 24753 1171 47 42 25.0 105.2 -
3 0.0318 5.61 0.038 0.038 0.029 0.040 0.012 32197 857 46 43 18.4 79.6 -
4 0.0445 4.74 0.034 0.034 0.031 0.035 0.010 37743 1212 57 53 21.4 91.2 -
5 0.0572 4.18 0.036 0.036 0.032 0.038 0.011 42642 1181 59 57 19.9 83.8 -
6 0.0699 3.78 0.049 0.049 0.036 0.052 0.015 47224 883 59 57 15.1 65.1 -
7 0.0826 3.48 0.065 0.065 0.040 0.068 0.020 51052 685 59 58 11.7 50.9 -
8 0.0953 3.24 0.096 0.096 0.045 0.100 0.029 54636 448 56 56 8.0 35.0 -
9 0.1080 3.04 0.151 0.151 0.050 0.158 0.046 58072 268 53 53 5.1 22.7 -
10 0.1207 2.88 0.229 0.229 0.056 0.240 0.070 60731 171 51 51 3.3 15.4 -
11 0.1334 2.74 0.314 0.314 0.063 0.329 0.097 63807 125 51 51 2.4 11.3 -
12 0.1461 2.62 0.406 0.406 0.070 0.425 0.125 66241 98 51 52 1.9 8.7 -
13 0.1588 2.51 0.537 0.537 0.078 0.562 0.166 68272 76 53 53 1.4 6.5 -
14 0.1715 2.41 0.685 0.685 0.084 0.727 0.237 54170 61 53 54 1.1 4.4 -
15 0.1843 2.33 0.886 0.886 0.088 0.950 0.333 39781 48 55 56 0.9 3.1 -
16 0.1970 2.25 1.249 1.249 0.092 1.355 0.508 29754 35 57 59 0.6 2.1 -

## Slide 151: CCP4: aimless log
$TABLE: Correlations CC(1/2) within dataset, XDSdataset:
$GRAPHS: Anom & Imean CCs v resolution:0|0.216023x0|1:2,4,7:
: RMS correlation ratio :0|0.216023x0|2.20344:2,6:
$$
N 1/d^2 Dmid CCanom Nanom RCRanom CC1/2 NImean $$ $$
1 0.0064 12.55 0.659 499 2.203 1.000 669
2 0.0191 7.24 0.550 975 1.853 1.000 1155
3 0.0318 5.61 0.527 1295 1.798 1.000 1479
16 0.1970 2.25 0.037 2123 1.038 0.711 2275
17 0.2097 2.18 0.043 1682 1.044 0.460 1877

## Slide 152: XDS: CORRECT.LP or XSCALE.LP
SUBSET OF INTENSITY DATA WITH SIGNAL/NOISE >= -3.0 AS FUNCTION OF RESOLUTION
RESOLUTION NUMBER OF REFLECTIONS COMPLETENESS R-FACTOR R-FACTOR COMPARED I/SIGMA R-meas CC(1/2) Anomal SigAno Nano
LIMIT OBSERVED UNIQUE POSSIBLE OF DATA observed expected Corr
6.39 4103 2517 3135 80.3% 1.7% 1.8% 3018 33.77 2.3% 99.9* 21* 1.012 533
4.54 7339 4970 5544 89.6% 2.6% 2.6% 4585 22.56 3.6% 99.8* 9 0.914 796
3.72 9442 6724 7186 93.6% 2.8% 2.9% 5327 19.99 4.0% 99.7* 9 0.859 877
3.22 11077 7997 8486 94.2% 4.8% 4.9% 6094 12.27 6.8% 99.3* -1 0.784 913
2.88 12408 8851 9595 92.2% 10.0% 10.1% 7068 6.01 14.2% 97.8* -2 0.799 1012
2.63 13673 9561 10662 89.7% 20.8% 20.8% 8185 3.10 29.4% 88.8* -4 0.776 1274
2.44 14360 9852 11510 85.6% 34.5% 34.7% 8981 1.90 48.8% 75.9* 2 0.765 1525
2.28 9501 6495 12416 52.3% 61.8% 59.6% 5991 1.14 87.3% 53.5* -2 0.722 1031
2.15 4567 3307 13210 25.0% 120.5% 120.3% 2520 0.59 170.4% 21.9* 4 0.693 408
total 86470 60274 81744 73.7% 6.7% 6.7% 51769 8.97 9.5% 99.5* 2 0.804 8369

## Slide 153: XDS: CORRECT.LP or XSCALE.LP
******************************************************************************
CORRECTION PARAMETERS FOR THE STANDARD ERROR OF REFLECTION INTENSITIES
******************************************************************************
The variance v0(I) of the intensity I obtained from counting statistics is
replaced by v(I)=a*(v0(I)+b*I^2). The model parameters a, b are chosen to
minimize the discrepancies between v(I) and the variance estimated from
sample statistics of symmetry related reflections. This model implicates
an asymptotic limit ISa=1/SQRT(a*b) for the highest I/Sigma(I) that the
experimental setup can produce (Diederichs (2010) Acta Cryst D66, 733-740).
a b ISa
3.806E+00 1.080E-04 49.32

## Slide 154: Lorentz factor: note added in proof
Debye, P. J. W. (1914)."Interferenz von Röntgenstrahlen und Wärmebewegung", Annalen der Physik 348, 49-92.
Lorentz

## Slide 155: Lorentz Factor
Ewald sphere
Φ circle
diffracted ray
(h,k,l)
spindle axis

## Slide 156: Lorentz Factor
Ewald sphere
diffracted ray
spindle axis

## Slide 157: Lorentz Factor
Ewald sphere
diffracted ray
spindle axis

## Slide 158: Lorentz Factor
Ewald sphere
spindle axis
diffracted ray

## Slide 159: The “Lorentz zone”

## Slide 160: Lorentz factor
+
50x
1.2x
550x
3x
10x

## Slide 161: Polarization factor
1.0
0.45
+

## Slide 162: attenuation factor
Bouguer, P. (1729). Essai d'optique sur la gradation de la lumière.
Lambert, J. H. (1760). Photometria: sive De mensura et gradibus luminis, colorum et umbrae. E. Klett.
Beer, A. (1852)."Bestimmung der Absorption des rothen Lichts in farbigen Flüssigkeiten", Ann. Phys. Chem 86, 78-90.
A = = exp[-μxtal(txi+ txo)
-μsolvent(tsi + tso)]
IT
Ibeam
μxtal
txi
txo
tsi
tso
txi
txo
tsi
tso
txi
txo
tsi
tso
μsolvent

## Slide 163: Where do photons go?
beamstop
98%
Transmitted
100 µm Protein
1 Å X-rays
error < total < 2%

## Slide 164: Air absorption
Distance from sample (mm)
transmittance

## Slide 165
Pick-up tool mark
Bragg
glitch
oxygen
inclusions
“tree rings”
1% high
average
1% low
~3x105 photon/pixel Pilatus detector
Spatial Heterogeneity in Sharp Spot Sensitivity

## Slide 166: A few other corrections…
Meisburger, et. al. Nat Commun 11, 1271 (2020). https://doi.org/10.1038

## Slide 167
CORRECT – after all the “factors”

## Slide 168: What you know when:
Autoindexing: Lattice (14)
Scaling: Point Group (43)
Phasing: Space Group (84)

## Slide 169: Asymmetric Unit (ASU)
Smallest part of the crystal
that you can use to build up the rest of the crystal
with symmetry operations
Unit Cell
Smallest part of the crystal
that you can use to build up the rest of the crystal
with translation symmetry only
“NCS ASU”
Builds ASU with Non-crystallographic symmetry operations

## Slide 170: Frequency of Space Groups in PDB
13007 | P212121 | 566 | R32 | 191 | P6322 | 74 | P4322 | 36 | P3212
9629 | P21 | 491 | P6522 | 187 | P213 | 71 | P4122 | 30 | P622
5300 | C2 | 474 | P61 | 161 | P6 | 70 | P64 | 24 | F4132
3021 | P21212 | 421 | P65 | 144 | P4212 | 70 | P6222 | 22 | P4332
2543 | C2221 | 353 | P41 | 142 | I23 | 67 | P6422 | 22 | P422
2109 | P1 | 351 | P63 | 132 | F432 | 66 | F222 | 22 | P23
2030 | P43212 | 286 | P31 | 124 | P321 | 63 | P62 | 19 | P3112
1565 | P3221 | 257 | I422 | 115 | I41 | 59 | P4132 | 18 | P4222
1359 | P41212 | 253 | P42212 | 103 | P3 | 46 | P2221 | 18 | I4132
1198 | P3121 | 249 | I4 | 99 | I213 | 43 | P4 | 16 | P432
1147 | I222 | 246 | P43 | 94 | C222 | 43 | I432 | 8 | P4232
825 | R3 | 239 | I4122 | 93 | P2 | 40 | P42 | 3 | P222
620 | P6122 | 193 | P32 | 82 | I212121 | 37 | F23 | 2 | P312

## Slide 171: Pitfalls: sub-space groups
P622
What else?
P6122 P6522 P6222 P6422 P6322
P6 P61 P65 P62 P64 P63
P312 P321 P3112 P3121 P3212 P3221
P3 P31 P32
P2 P21 C2 C222 C2221
P1
zanuda

## Slide 172: Trick Question: What is the space group?
a = 63 b = 63 c = 63
α = 90 β = 90 γ = 90

## Slide 173: Diffraction from a 3D lattice
real space
reciprocal space
+

## Slide 174: Diffraction from a 3D lattice
real space
reciprocal space
+

## Slide 175: Diffraction from a 3D lattice
real space
reciprocal space
+
+

## Slide 176: Diffraction from a 3D lattice
real space
reciprocal space
+
+

## Slide 177
Base-centered lattice

## Slide 178
Base-centered lattice

## Slide 179
Face-centered lattice

## Slide 180
Face-centered lattice

## Slide 181
Face-centered lattice

## Slide 182
Face-centered lattice
Body-centered lattice

## Slide 183
Face-centered lattice
All half-way points

## Slide 184
Face-centered lattice
All half-way points

## Slide 185
cubic lattice