## Slide 1: Acknowledgements
http://bl831.als.lbl.gov/~jamesh/powerpoint/CSHL_beamline_2025.pptx
UCSF LBNL SLAC
ALS 8.3.1: TomAlberTron
NIH NIGMS R35 GM158447 to Holton
NIH NIAID P50 AI150476 to Krogan
NIH NIAID U19 AI171110 to Krogan
NIH NIGMS P30 GM124169 to Adams
DOE-BER IDAT to Hura
NIH NIGMS P41 GM103393 to Hodgson
The Advanced Light Source is supported by the Director, Office of Science, Office of Basic Energy Sciences, Materials Sciences Division, of the US Department of Energy under contract No. DE-AC02-05CH11231 at Lawrence Berkeley National Laboratory.
Robert Stroud James Fraser
Paul Adams Greg Hura
Aina Cohen Keith Hodgson

## Slide 3: PowerPoint File available:
http://bl831.als.lbl.gov/
~jamesh/powerpoint/
CSHL_beamline_2025.pptx

## Slide 4: The “success rate” of data collection
100 s/dataset
200 days/year
~150 beamlines
~26,000,000 datasets/year
~14,000 PDBs in 2024

## Slide 5: signal vs noise
“If you don’t have good data, then you have no data at all.”
-Sung-Hou Kim

## Slide 6: signal vs noise
easy
hard
impossible

## Slide 7: signal vs noise
easy
hard
impossible
threshold of “solvability”

## Slide 8: signal vs noise
“If you don’t have good data, then you must learn statistics.”
-James Holton

## Slide 9: Adding noise

## Slide 10: Adding noise
12 + 12 = 1.42

## Slide 11: Adding noise
12 + 12 = 1.42
32 + 12 = 3.22
σtotal2 = σ12 + σ22

## Slide 12: Adding noise
12 + 12 = 1.42
32 + 12 = 3.22
σtotal2 = σ12 + σ22

## Slide 13: Adding noise
12 + 12 = 1.42
32 + 12 = 3.22
σtotal2 = σ12 + σ22

## Slide 14: Adding noise
12 + 12 = 1.42
32 + 12 = 3.22
102 + 12 = 10.052

## Slide 15: And the dominant source of error is…
Resolution
  Disorder (including Radiation Damage)
  Background
Phases
  Non-isomorphism (including Radiation Damage)
  Vibration (beam, xtal, spindle, etc.)
  Detector calibration

## Slide 16: Correctable sources of error: “scaling”
Lorentz Factor
Polarization Factor
Smooth stuff
  Illuminated volume vs rotation
  Source decay
  Attenuation Factor
  Low-frequency calibration errors

## Slide 17: Lorentz factor: note added in proof
Debye, P. J. W. (1914)."Interferenz von Röntgenstrahlen und Wärmebewegung", Annalen der Physik 348, 49-92.
Lorentz

## Slide 18: Lorentz Factor
Ewald sphere
Φ circle
diffracted ray
(h,k,l)
spindle axis

## Slide 19: Lorentz Factor
Ewald sphere
diffracted ray
spindle axis

## Slide 20: Lorentz Factor
Ewald sphere
diffracted ray
spindle axis

## Slide 21: Lorentz Factor
Ewald sphere
spindle axis
diffracted ray

## Slide 22: The “Lorentz zone”

## Slide 23: Lorentz factor
+
50x
1.2x
550x
3x
10x

## Slide 24: Polarization factor
1.0
0.45
+

## Slide 25: Air absorption
Chart: 
| 0 | 20 | 40 | 60 | 80 | 100 | 120 | 140 | 160 | 180 | 200 | 220 | 240 | 260 | 280 | 300 | 320 | 340 | 360 | 380 | 400 | 420 | 440 | 460 | 480 | 500 | 520 | 540 | 560 | 580 | 600 | 620 | 640 | 660 | 680 | 700 | 720 | 740 | 760 | 780 | 800 | 820 | 840 | 860 | 880 | 900 | 920 | 940 | 960 | 980 | 1000
8047 | 1 | 0.97664 | 0.95383 | 0.93155 | 0.90979 | 0.88854 | 0.86778 | 0.84751 | 0.82772 | 0.80838 | 0.7895 | 0.77106 | 0.75305 | 0.73546 | 0.71828 | 0.7015 | 0.68511 | 0.66911 | 0.65348 | 0.63822 | 0.62331 | 0.60875 | 0.59453 | 0.58064 | 0.56708 | 0.55383 | 0.5409 | 0.52826 | 0.51592 | 0.50387 | 0.4921 | 0.48061 | 0.46938 | 0.45842 | 0.44771 | 0.43725 | 0.42704 | 0.41706 | 0.40732 | 0.39781 | 0.38851 | 0.37944 | 0.37058 | 0.36192 | 0.35347 | 0.34521 | 0.33715 | 0.32927 | 0.32158 | 0.31407 | 0.30673
12398 | 1 | 0.99337 | 0.98677 | 0.98023 | 0.97372 | 0.96726 | 0.96084 | 0.95447 | 0.94814 | 0.94185 | 0.9356 | 0.92939 | 0.92322 | 0.9171 | 0.91101 | 0.90497 | 0.89896 | 0.893 | 0.88707 | 0.88119 | 0.87534 | 0.86953 | 0.86376 | 0.85803 | 0.85234 | 0.84668 | 0.84106 | 0.83548 | 0.82994 | 0.82443 | 0.81896 | 0.81353 | 0.80813 | 0.80277 | 0.79744 | 0.79215 | 0.7869 | 0.78168 | 0.77649 | 0.77134 | 0.76622 | 0.76114 | 0.75609 | 0.75107 | 0.74608 | 0.74113 | 0.73622 | 0.73133 | 0.72648 | 0.72166 | 0.71687
17479 | 1 | 0.99738 | 0.99478 | 0.99217 | 0.98958 | 0.98699 | 0.98441 | 0.98184 | 0.97927 | 0.97671 | 0.97415 | 0.9716 | 0.96906 | 0.96653 | 0.964 | 0.96148 | 0.95896 | 0.95646 | 0.95395 | 0.95146 | 0.94897 | 0.94649 | 0.94401 | 0.94154 | 0.93908 | 0.93663 | 0.93418 | 0.93173 | 0.9293 | 0.92687 | 0.92444 | 0.92202 | 0.91961 | 0.91721 | 0.91481 | 0.91241 | 0.91003 | 0.90765 | 0.90528 | 0.90291 | 0.90055 | 0.89819 | 0.89584 | 0.8935 | 0.89116 | 0.88883 | 0.88651 | 0.88419 | 0.88187 | 0.87957 | 0.87727
Distance from sample (mm)
transmittance

## Slide 26: attenuation factor
Bouguer, P. (1729). Essai d'optique sur la gradation de la lumière.
Lambert, J. H. (1760). Photometria: sive De mensura et gradibus luminis, colorum et umbrae. E. Klett.
Beer, A. (1852)."Bestimmung der Absorption des rothen Lichts in farbigen Flüssigkeiten", Ann. Phys. Chem 86, 78-90.
A = = exp[-μxtal(txi+ txo)
-μsolvent(tsi + tso)]
IT
Ibeam
μxtal
txi
txo
tsi
tso
txi
txo
tsi
tso
txi
txo
tsi
tso
μsolvent

## Slide 27: Where do photons go?
beamstop
98%
Transmitted
100 µm Protein
1 Å X-rays
error < total < 2%

## Slide 28
Pick-up tool mark
Bragg
glitch
oxygen
inclusions
“tree rings”
1% high
average
1% low
~3x105 photon/pixel Pilatus detector
Spatial Heterogeneity in Sharp Spot Sensitivity

## Slide 29: A few other corrections…
Meisburger, et. al. Nat Commun 11, 1271 (2020). https://doi.org/10.1038

## Slide 30
CORRECT – after all the “factors”

## Slide 31: Data quality - simple rules:
  Put your crystal into the beam

## Slide 32: Your sample is a lens! (but not for x-rays)
90°

## Slide 33: Your sample is a lens! (but not for x-rays)
90°

## Slide 34: Data quality - simple rules:
  Put your crystal into the beam
  Shoot the whole crystal

## Slide 35: shoot the whole crystal

## Slide 36: shoot the whole crystal

## Slide 37: shoot the whole crystal

## Slide 38: shoot the whole crystal

## Slide 39: shoot the whole crystal

## Slide 40: How many “crystals” do you see?
Shoot the crystal (singular)

## Slide 41: Data quality - simple rules:
  Put your crystal into the beam
  Shoot the whole crystal
  Shoot nothing but the crystal

## Slide 42: shoot nothing but the crystal

## Slide 43: shoot nothing but the crystal

## Slide 44: scattering/absorptoin “rules”:
1 μm crystal ≈ 1 μm water
≈ 1 μm plastic
≈ 0.1 μm glass
≈ 1000 μm air
Optimum: support thickness = crystal thickness
1 + 1 = 1.4

## Slide 45
X-ray beam size (μm)
Resolution (Å)
too big
too small
just right
limited by
radiation
damage
limited by
excess
background

## Slide 46: platy crystals

## Slide 47: platy crystals

## Slide 48: Data quality - simple rules:
  Put your crystal into the beam
  Shoot the whole crystal
  Shoot nothing but the crystal
  The crystal must rotate

## Slide 49: “pretty” crystals

## Slide 50: “pretty” crystals

## Slide 51: “pretty” crystals

## Slide 52: The crystal rotates!

## Slide 53: The crystal rotates!

## Slide 54: avoiding overlaps
c
c

## Slide 55: avoiding overlaps
c

## Slide 56: avoiding overlaps
c
http://www.mitegen.com/mic_catalog.php?c=DTMicroMounts

## Slide 57: avoiding overlaps
c
http://www.mitegen.com/mic_catalog.php?c=orientedmounts

## Slide 58: Data quality - simple rules:
  Put your crystal into the beam
  Shoot the whole crystal
  Shoot nothing but the crystal
  The crystal must rotate
  Back off!

## Slide 59: Background scattering

## Slide 60: Mosquito tips cost $0.15 each
So do detector pixels!
We bought 6,224,001
Use them!
(they are reusable)

## Slide 61: Background scattering
$100,000.00
$100,000.00
$100,000.00
$100,000.00
$100,000.00
$100,000.00
$100,000.00
$100,000.00
real estate is
expensive
use it!

## Slide 62: Fine Slicing
Pflugrath, J. W. (1999)."The finer things in X-ray diffraction data collection", Acta Cryst. D 55, 1718-1725.
background
background

## Slide 64: Optimal exposure time (anomalous on a PAD)
tmad Optimal exposure time (s)
tro read-out time (s)
2.03 ms on Pilatus3 S
0.95 ms on Pilatus3 X
0.003 ms on Eiger
Attenuate for MAD
PILATUS
>0.2 s/frame
EIGER
>0.0003 s/frame

## Slide 65: Optimal exposure time (anomalous on a PAD)
time
intensity
2.03 ms

## Slide 66: Optimal exposure time (faint spots on CCD detector)
thr Optimal exposure time for data set (s)
tref exposure time of reference image (s)
bgref background level near weak spots on
reference image (ADU)
bg0 ADC offset of detector (ADU)
bghr optimal background level (via thr)
σ0 rms read-out noise (ADU)
gain ADU/photon
m multiplicity of data set (including partials)
adjust exposure
so this is ~100

## Slide 67: And the dominant source of error is…
Resolution
  Disorder (including Radiation Damage)
  Background
Phases
  Non-isomorphism (including Radiation Damage)
  Vibration (beam, xtal, spindle, etc.)
  Detector calibration

## Slide 68: Detector calibration: 7235 eV

## Slide 69: Detector calibration: 7247 eV

## Slide 70: Gadox calibration vs energy
photon energy (keV)
Relative absorption depth
same = good!
bad!

## Slide 71: Systematic error does not “average out”
Unless you can randomize their source

## Slide 72: Fractional error
mult > (—)2
~3%
<ΔF/F>

## Slide 73
anomalous signal
Crick, F. H. C. & Magdoff, B. S. (1956) Acta Crystallogr. 9, 901-908.
Hendrickson, W. A. & Teeter, M. M. (1981) Nature 290, 107-113.
# sites
MW (Da)
ΔF
F
≈ 1.2 f”
√
World record!
ΔF/F = 0.5%
Wang, Dauter & Dauter (2006) Acta Cryst. D 62, 1475-1483.

## Slide 74
Holton & Frankel (2010) Acta D 66 393-408.

## Slide 75: 100% SeMet incorporation
Trivial to solve

## Slide 76: 100% S -Met incorporation
Impossible to solve

## Slide 77: Phase Problem: Se vs S
fraction not Selenium
correlation coefficient

## Slide 78: Phase Problem: Se vs S
fraction not Selenium
correlation coefficient

## Slide 79: Phase Problem: Se vs S
fraction not Selenium
correlation coefficient

## Slide 80: Phase Problem: Se vs S
fraction not Selenium
correlation coefficient

## Slide 81: Phase Problem: Se vs S
fraction not Selenium
correlation coefficient

## Slide 82: One xtal, no idea? Suggested native protocol:
Do strategy
Start at specified phi
100% fast as you can
Keep going to 360°
Move detector
Advance phi 10-30 deg
Increase photons/image 4x
goto 4

## Slide 83: Suggested anomalous protocol:
2 wavelengths are better than 1
- (peak + inf)/2, and remote
MAD, not M-SAD!
360° in < ~5 MGy (for Se)
move detector
4X exposure
goto 2

## Slide 84: Summary
1 + 1 = 1.4
Match beam to “crystal”
Match goop thickness to crystal
Better beam = less time
Look at your spots!
“true multiplicity” for anomalous
Crystal must rotate

## Slide 85: The “threshold of solvability” is sharp!
How can we predict success/failure?
Know Thy Experiment

## Slide 86: The “beam line”
Area Detector
Protein Crystal
Cite the “beamline paper!”
X-ray
Source
Notes: Most notable acheivements are high brightness with very little “wasted” light. The discovery made by Howard Padmore? At ALS showed that a 2:1 demagnification cancels nearly all of the spherical abberations in the parabolic/torroid system. 30-50% of the total x-ray intensity passes through the 100um pinhole. The flux through the 100um pinhole at the superbend beamlines is comparable to that of the 5.0.2 wiggler or an APS bending magnet, but only 150 Watts of power must be dissipated by the beamline optics. In contrast, wiggler beamlines must disipate kilowatts. The optics here are much simpler and not exotic in any way. The system is stable and robust. There is never a need to re-focus the beam at a new wavelength. Only a 2-minute automatic tuneup is recommended for very large energy changes (more than 1-2 keV).

## Slide 87: The “beam line”
Area Detector
Source
Primary
mirror
Secondary
mirror
Si(111)
monochromator
Sample
focused
beam
slits
Scatter
guard
Unfocused
beam
slits

## Slide 88: The “beam line”
Area Detector
Source
Primary
mirror
Secondary
mirror
Si(111)
monochromator
Sample
focused
beam
slits
Scatter
guard
Unfocused
beam
slits
Notes: Most notable acheivements are high brightness with very little “wasted” light. The discovery made by Howard Padmore? At ALS showed that a 2:1 demagnification cancels nearly all of the spherical abberations in the parabolic/torroid system. 30-50% of the total x-ray intensity passes through the 100um pinhole. The flux through the 100um pinhole at the superbend beamlines is comparable to that of the 5.0.2 wiggler or an APS bending magnet, but only 150 Watts of power must be dissipated by the beamline optics. In contrast, wiggler beamlines must disipate kilowatts. The optics here are much simpler and not exotic in any way. The system is stable and robust. There is never a need to re-focus the beam at a new wavelength. Only a 2-minute automatic tuneup is recommended for very large energy changes (more than 1-2 keV).

## Slide 89: Where do the X-rays come from?

## Slide 90: The “Home source”
e-

## Slide 91: The “Home source”
e-
+
e-

## Slide 92: The “Home source”
+

## Slide 93: The “Home source”
e-
+

## Slide 94: The “Home source”

## Slide 95: Metal identification
scattering
same
λ as
incident

## Slide 96
“MAD scan” – absorption spectrum
“counts” (aka arbitrary units)
0.980888
0.980500
0.980113
0.979725
0.979338
0.978952
0.978565
0.978179
0.977794
0.977408
0.977023
0.976638
0.976254
Wavelength (Å)
fluorescence = yield * absorption
12398.42
energy

## Slide 97
maximize
absolute
value of
f’’
maximize
difference
in f’
“MAD scan” – absorption spectrum

## Slide 98
4-wavelength MAD

## Slide 99
3-wavelength MAD

## Slide 100
2-wavelength MAD

## Slide 101
2-wavelength MAD

## Slide 102
1-wavelength SAD

## Slide 103
2-wavelength metal validation
phased
anomalous
difference
Fourier

## Slide 104: The “beam line”
Area Detector
Source
Primary
mirror
Secondary
mirror
Si(111)
monochromator
Sample
focused
beam
slits
Scatter
guard
Unfocused
beam
slits
Notes: Most notable acheivements are high brightness with very little “wasted” light. The discovery made by Howard Padmore? At ALS showed that a 2:1 demagnification cancels nearly all of the spherical abberations in the parabolic/torroid system. 30-50% of the total x-ray intensity passes through the 100um pinhole. The flux through the 100um pinhole at the superbend beamlines is comparable to that of the 5.0.2 wiggler or an APS bending magnet, but only 150 Watts of power must be dissipated by the beamline optics. In contrast, wiggler beamlines must disipate kilowatts. The optics here are much simpler and not exotic in any way. The system is stable and robust. There is never a need to re-focus the beam at a new wavelength. Only a 2-minute automatic tuneup is recommended for very large energy changes (more than 1-2 keV).

## Slide 105: Electric field of a moving charge

## Slide 106: Electric field of a moving charge
accelerating charges
make
electromagnetic waves
(light)

## Slide 107: Bending Magnet
N
electron beam
X-rays !
S

## Slide 108: Two Bending Magnets
N
electron beam
S
S
N
2x X-rays !

## Slide 109: Wiggler
N
S
N
S
N
S
N
S
S
N
S
N
S
N
S
N
8x X-rays !

## Slide 110: Undulator
N
S
N
S
N
S
N
S
S
N
S
N
S
N
S
N
> 8x X-rays !

## Slide 111: Undulator emission spectrum

## Slide 112: XFEL Undulator
N
S
N
S
N
S
N
S
S
N
S
N
S
N
S
N
same X-rays
short pulse
S
N
S
N
S
N
S
N
N
S
N
S
N
S
N
S

## Slide 113: LCLS undulator hall

## Slide 114: LCLS at SLAC

## Slide 115: Accelerator-based light source terminology
Code name Translation
First Generation Bending Magnet
Second Generation Wiggler
Third Generation Undulator
Fourth Generation Free electron laser
..and diffraction-limited synchrotron

## Slide 116: Diffraction-limited light source
λ ≈ 2 d sinθ
λ = x-ray wavelength
d = electron/x-ray beam size
θ = beam divergence
d = 1 Å / 100 µrad / 2 = 0.5 µm

## Slide 117: The “beam line”
Area Detector
Source
Primary
mirror
Secondary
mirror
Si(111)
monochromator
Sample
focused
beam
slits
Scatter
guard
Unfocused
beam
slits
Notes: Most notable acheivements are high brightness with very little “wasted” light. The discovery made by Howard Padmore? At ALS showed that a 2:1 demagnification cancels nearly all of the spherical abberations in the parabolic/torroid system. 30-50% of the total x-ray intensity passes through the 100um pinhole. The flux through the 100um pinhole at the superbend beamlines is comparable to that of the 5.0.2 wiggler or an APS bending magnet, but only 150 Watts of power must be dissipated by the beamline optics. In contrast, wiggler beamlines must disipate kilowatts. The optics here are much simpler and not exotic in any way. The system is stable and robust. There is never a need to re-focus the beam at a new wavelength. Only a 2-minute automatic tuneup is recommended for very large energy changes (more than 1-2 keV).

## Slide 118: X-ray optics: mirrors
Notes: Next slide please

## Slide 119: Why are the mirrors so long?
3 mm
Must be < 0.2°
1000 mm
Notes: Next slide please

## Slide 120: There is a lens for X-rays!
1 mm
0.1 mm
Be
Notes: Next slide please

## Slide 121: The “beam line”
Area Detector
Source
Primary
mirror
Secondary
mirror
Si(111)
monochromator
Sample
focused
beam
slits
Scatter
guard
Unfocused
beam
slits
Notes: Most notable acheivements are high brightness with very little “wasted” light. The discovery made by Howard Padmore At ALS showed that a 2:1 demagnification cancels nearly all of the spherical abberations in the parabolic/torroid system. 30-50% of the total x-ray intensity passes through the 100um pinhole. The flux through the 100um pinhole at the superbend beamlines is comparable to that of the 5.0.2 wiggler or an APS bending magnet, but only 150 Watts of power must be dissipated by the beamline optics. In contrast, wiggler beamlines must disipate kilowatts. The optics here are much simpler and not exotic in any way. The system is stable and robust. There is never a need to re-focus the beam at a new wavelength. Only a 2-minute automatic tuneup is recommended for very large energy changes (more than 1-2 keV).

## Slide 122
Single crystal

## Slide 123: Monochromators
λ = 1 Å
Si
Notes: Next slide please

## Slide 124: Monochromators
λ = 1 Å
Si
d = 3.135 Å
θ = 9.177°
λ = 1 Å
nλ = 2d sin(θ)
Notes: Next slide please

## Slide 125: Monochromators
nλ = 2d sin(θ)
λ = 1.1 Å
Si
d = 3.135 Å
θ = 9.177°
Notes: Next slide please

## Slide 126: Monochromators
nλ = 2d sin(θ)
λ = 1.1 Å
Si
d = 3.135 Å
θ = 10.104°
λ = 1.1 Å
Notes: Next slide please

## Slide 127: Monochromators
nλ = 2d sin(θ)
λ = 1 Å
Si
d = 3.135 Å
θ = 9.177°
λ = 1 Å
Notes: Next slide please

## Slide 128: Monochromators
Si
λ = 1 Å
λ = 1 Å
Si
d = 3.135 Å
DCM = Double Crystal Monochromator
Notes: Next slide please

## Slide 129: Monochromators
Si
λ = 1.1 Å
λ = 1.1 Å
Si
d = 3.135 Å
DCM = Double Crystal Monochromator
Notes: Next slide please

## Slide 130: Monochromators
nλ = 2d sin(θ)
λ = 1.54 Å
θ = 1.84°
λ = 1.54 Å
Mo/C
d = 24 Å
Multilayer optics
Notes: Next slide please

## Slide 131: The truth about x-ray beams
Term units significance
Flux photons/s duration of experiment
Beam Size μm match to crystal
Divergence mrad spot size vs distance
Wavelength Å resolution and absorption
Dispersion Δλ/λ spot size
Flux density ph/s/area scattering/damage rate
Fluence ph/area scattering/damage

## Slide 132: The truth about x-ray beams
Term units significance
Flux photons/s duration of experiment
Beam Size μm match to crystal
Divergence mrad spot size vs distance
Wavelength Å resolution and absorption
Dispersion Δλ/λ spot size
Flux density ph/s/area scattering/damage rate
Fluence ph/area scattering/damage

## Slide 133
The number of photons scattered
before crystal is dead
is independent
of data collection time
1 um3 = 106 photons
Henderson, 1990; Gonzalez & Nave, 1994; Glaeser et al., 2000; Sliz et al., 2003; Leiros et al., 2006; Owen et al., 2006; Garman & McSweeney, 2006; Garman & Nave, 2009; Holton, 2009

## Slide 134: The truth about x-ray beams
quantity | units | home source | synchrotron
flux | Photons/second | 2.5 x 109 | 2 x 1012
exposure | seconds | 400 | 1

## Slide 135: The truth about x-ray beams
quantity | units | home source | synchrotron
flux | Photons/second | 2.5 x 109 | 2 x 1012
exposure | seconds | 400 | 1
Dispersion | wavelength range / wavelength | 0.2% (Kα1 - Kα2) | 0.014% (Si111)
Divergence | milliRadian | 4.8 | 1.8 (h) 0.18 (v)
Beam size | microns | 100 | 24 (h) 12 (v)
Spectral brightness | Photons/s/mm2/ mR2/0.1%BW | 5.4 x 109 | 1.5 x 1017

## Slide 136: The truth about x-ray beams
quantity | units | home source | synchrotron
flux | Photons/second | 72 x 106 | 2 x 1012
exposure | time | 4 hours | 1 second
Dispersion | wavelength range / wavelength | 0.2% (Kα1 - Kα2) | 0.014% (Si111)
Divergence | milliRadian | 4.8 | 1.8 (h) 0.18 (v)
Beam size | microns | 17 | 24 (h) 12 (v)
Spectral brightness | Photons/s/mm2/ mR2/0.1%BW | 5.4 x 109 | 1.5 x 1017

## Slide 137: The truth about x-ray beams
quantity | units | home source | synchrotron
flux | Photons/second | 1 x 106 | 2 x 1012
exposure | time | 10 days | 1 second
Dispersion | wavelength range / wavelength | 0.2% (Kα1 - Kα2) | 0.014% (Si111)
Divergence | milliRadian | 0.6 | 1.8 (h) 0.18 (v)
Beam size | microns | 17 | 24 (h) 12 (v)
Spectral brightness | Photons/s/mm2/ mR2/0.1%BW | 5.4 x 109 | 1.5 x 1017

## Slide 138: The truth about x-ray beams
quantity | units | home source | synchrotron
flux | Photons/second | 8 x 104 | 2 x 1012
exposure | time | 5 months | 1 second
Dispersion | wavelength range / wavelength | 0.014% (Si111) | 0.014% (Si111)
Divergence | milliRadian | 0.6 | 1.8 (h) 0.18 (v)
Beam size | microns | 17 | 24 (h) 12 (v)
Spectral brightness | Photons/s/mm2/ mR2/0.1%BW | 5.4 x 109 | 1.5 x 1017

## Slide 139: The truth about x-ray beams
Term units significance
Flux photons/s duration of experiment
Beam Size μm match to crystal
Divergence mrad spot size vs distance
Wavelength Å resolution and absorption
Dispersion Δλ/λ spot size
Flux density ph/s/area scattering/damage rate
Fluence ph/area scattering/damage

## Slide 140
Two-color experiment (ca 1914)
(Cu source)
20 eV

## Slide 141: spectral dispersion
Ewald sphere
spindle axis
Φ circle
diffracted ray
(h,k,l)
d*
λ*
λ*

## Slide 142: spectral dispersion
Ewald sphere
spindle axis
Φ circle
diffracted ray
(h,k,l)
d*
λ’*
λ’*

## Slide 143: dispersion = 0

## Slide 144: dispersion = 0.014% (Si 111)

## Slide 145: dispersion = 0.25% (CuKα)

## Slide 146: dispersion = 0.6%

## Slide 147: dispersion = 1.3%

## Slide 148: dispersion = 2.6%

## Slide 149: dispersion = 5.1%

## Slide 150: The truth about x-ray beams
Term units significance
Flux photons/s duration of experiment
Beam Size μm match to crystal
Divergence mrad spot size vs distance
Wavelength Å resolution and absorption
Dispersion Δλ/λ spot size
Flux density ph/s/area scattering/damage rate
Fluence ph/area scattering/damage

## Slide 151: beam divergence
Ewald sphere
spindle axis
Φ circle
diffracted ray
(h,k,l)
d*
λ*
λ*

## Slide 152: beam divergence
spindle axis
Φ circle
diffracted ray
(h,k,l)
d*
Ewald sphere
λ*
λ*

## Slide 153: divergence = 0 º

## Slide 154: divergence = 0.3 º

## Slide 155: mosaic spread
Ewald sphere
spindle axis
Φ circle
diffracted ray
(h,k,l)
λ*
λ*
d*
d*

## Slide 156: mosaic spread
Ewald sphere
spindle axis
Φ circle
diffracted rays
(h,k,l)
d*
d*

## Slide 157: The truth about x-ray beams
Term units significance
Flux photons/s duration of experiment
Beam Size μm match to crystal
Divergence mrad spot size vs distance
Wavelength Å resolution and absorption
Dispersion Δλ/λ spot size
Flux density ph/s/area scattering/damage rate
Fluence ph/area scattering/damage

## Slide 158: Air absorption
Chart: 
| 0 | 20 | 40 | 60 | 80 | 100 | 120 | 140 | 160 | 180 | 200 | 220 | 240 | 260 | 280 | 300 | 320 | 340 | 360 | 380 | 400 | 420 | 440 | 460 | 480 | 500 | 520 | 540 | 560 | 580 | 600 | 620 | 640 | 660 | 680 | 700 | 720 | 740 | 760 | 780 | 800 | 820 | 840 | 860 | 880 | 900 | 920 | 940 | 960 | 980 | 1000
8047 | 1 | 0.97664 | 0.95383 | 0.93155 | 0.90979 | 0.88854 | 0.86778 | 0.84751 | 0.82772 | 0.80838 | 0.7895 | 0.77106 | 0.75305 | 0.73546 | 0.71828 | 0.7015 | 0.68511 | 0.66911 | 0.65348 | 0.63822 | 0.62331 | 0.60875 | 0.59453 | 0.58064 | 0.56708 | 0.55383 | 0.5409 | 0.52826 | 0.51592 | 0.50387 | 0.4921 | 0.48061 | 0.46938 | 0.45842 | 0.44771 | 0.43725 | 0.42704 | 0.41706 | 0.40732 | 0.39781 | 0.38851 | 0.37944 | 0.37058 | 0.36192 | 0.35347 | 0.34521 | 0.33715 | 0.32927 | 0.32158 | 0.31407 | 0.30673
12398 | 1 | 0.99337 | 0.98677 | 0.98023 | 0.97372 | 0.96726 | 0.96084 | 0.95447 | 0.94814 | 0.94185 | 0.9356 | 0.92939 | 0.92322 | 0.9171 | 0.91101 | 0.90497 | 0.89896 | 0.893 | 0.88707 | 0.88119 | 0.87534 | 0.86953 | 0.86376 | 0.85803 | 0.85234 | 0.84668 | 0.84106 | 0.83548 | 0.82994 | 0.82443 | 0.81896 | 0.81353 | 0.80813 | 0.80277 | 0.79744 | 0.79215 | 0.7869 | 0.78168 | 0.77649 | 0.77134 | 0.76622 | 0.76114 | 0.75609 | 0.75107 | 0.74608 | 0.74113 | 0.73622 | 0.73133 | 0.72648 | 0.72166 | 0.71687
17479 | 1 | 0.99738 | 0.99478 | 0.99217 | 0.98958 | 0.98699 | 0.98441 | 0.98184 | 0.97927 | 0.97671 | 0.97415 | 0.9716 | 0.96906 | 0.96653 | 0.964 | 0.96148 | 0.95896 | 0.95646 | 0.95395 | 0.95146 | 0.94897 | 0.94649 | 0.94401 | 0.94154 | 0.93908 | 0.93663 | 0.93418 | 0.93173 | 0.9293 | 0.92687 | 0.92444 | 0.92202 | 0.91961 | 0.91721 | 0.91481 | 0.91241 | 0.91003 | 0.90765 | 0.90528 | 0.90291 | 0.90055 | 0.89819 | 0.89584 | 0.8935 | 0.89116 | 0.88883 | 0.88651 | 0.88419 | 0.88187 | 0.87957 | 0.87727
Distance from sample (mm)
transmittance

## Slide 159: Summary
1 + 1 = 1.4
Match beam to “crystal”
Match goop thickness to crystal
Better beam = less time
Look at your spots!
“true multiplicity” for anomalous
Crystal must rotate

## Slide 160: Correct beam center
155.063 155.647
159.353 155.063
159.301 155.011
1512.73 1554.57
155.063 159.356 -250
ADXV
HKL2000
MOSFLM
XDS
DIALS

## Slide 161: There are 16 different “beam centers”
Origin
Beam X
Beam Y

## Slide 162: There are 16 different “beam centers”
Origin
Beam X
Beam Y
Cameraman
View
Beam
View

## Slide 163: Re-indexing cheat sheet
Dauter, Z. (1999). "Data-collection strategies." Acta Cryst. D 55(10): 1703-1717.

## Slide 164: CCP4: aimless log
$TABLE: Analysis against resolution, XDSdataset:
$GRAPHS:I/sigma, Mean Mn(I)/sd(Mn(I)):0|0.216023x0|137.14:2,13,14:
:Rmerge, Rfull, Rmeas, Rpim v Resolution:0|0.216023x0|1.70834:2,4,5,6,7:
:Average I, RMSdeviation and Sd:0|0.216023x0|1650.8:2,10,11,12:
:Fractional bias:0|0.216023x0|0:2,15:
$$
N 1/d^2 Dmid Rmrg Rfull Rcum Rmeas Rpim Nmeas AvI RMSdev sd I/RMS Mn(I/sd) FrcBias $$ $$
1 0.0064 12.55 0.020 0.020 0.020 0.021 0.006 13115 1651 57 42 29.2 137.1 -
2 0.0191 7.24 0.027 0.027 0.024 0.028 0.008 24753 1171 47 42 25.0 105.2 -
3 0.0318 5.61 0.038 0.038 0.029 0.040 0.012 32197 857 46 43 18.4 79.6 -
4 0.0445 4.74 0.034 0.034 0.031 0.035 0.010 37743 1212 57 53 21.4 91.2 -
5 0.0572 4.18 0.036 0.036 0.032 0.038 0.011 42642 1181 59 57 19.9 83.8 -
6 0.0699 3.78 0.049 0.049 0.036 0.052 0.015 47224 883 59 57 15.1 65.1 -
7 0.0826 3.48 0.065 0.065 0.040 0.068 0.020 51052 685 59 58 11.7 50.9 -
8 0.0953 3.24 0.096 0.096 0.045 0.100 0.029 54636 448 56 56 8.0 35.0 -
9 0.1080 3.04 0.151 0.151 0.050 0.158 0.046 58072 268 53 53 5.1 22.7 -
10 0.1207 2.88 0.229 0.229 0.056 0.240 0.070 60731 171 51 51 3.3 15.4 -
11 0.1334 2.74 0.314 0.314 0.063 0.329 0.097 63807 125 51 51 2.4 11.3 -
12 0.1461 2.62 0.406 0.406 0.070 0.425 0.125 66241 98 51 52 1.9 8.7 -
13 0.1588 2.51 0.537 0.537 0.078 0.562 0.166 68272 76 53 53 1.4 6.5 -
14 0.1715 2.41 0.685 0.685 0.084 0.727 0.237 54170 61 53 54 1.1 4.4 -
15 0.1843 2.33 0.886 0.886 0.088 0.950 0.333 39781 48 55 56 0.9 3.1 -
16 0.1970 2.25 1.249 1.249 0.092 1.355 0.508 29754 35 57 59 0.6 2.1 -

## Slide 165: CCP4: aimless log
$TABLE: Correlations CC(1/2) within dataset, XDSdataset:
$GRAPHS: Anom & Imean CCs v resolution:0|0.216023x0|1:2,4,7:
: RMS correlation ratio :0|0.216023x0|2.20344:2,6:
$$
N 1/d^2 Dmid CCanom Nanom RCRanom CC1/2 NImean $$ $$
1 0.0064 12.55 0.659 499 2.203 1.000 669
2 0.0191 7.24 0.550 975 1.853 1.000 1155
3 0.0318 5.61 0.527 1295 1.798 1.000 1479
16 0.1970 2.25 0.037 2123 1.038 0.711 2275
17 0.2097 2.18 0.043 1682 1.044 0.460 1877

## Slide 166: XDS: CORRECT.LP or XSCALE.LP
SUBSET OF INTENSITY DATA WITH SIGNAL/NOISE >= -3.0 AS FUNCTION OF RESOLUTION
RESOLUTION NUMBER OF REFLECTIONS COMPLETENESS R-FACTOR R-FACTOR COMPARED I/SIGMA R-meas CC(1/2) Anomal SigAno Nano
LIMIT OBSERVED UNIQUE POSSIBLE OF DATA observed expected Corr
6.39 4103 2517 3135 80.3% 1.7% 1.8% 3018 33.77 2.3% 99.9* 21* 1.012 533
4.54 7339 4970 5544 89.6% 2.6% 2.6% 4585 22.56 3.6% 99.8* 9 0.914 796
3.72 9442 6724 7186 93.6% 2.8% 2.9% 5327 19.99 4.0% 99.7* 9 0.859 877
3.22 11077 7997 8486 94.2% 4.8% 4.9% 6094 12.27 6.8% 99.3* -1 0.784 913
2.88 12408 8851 9595 92.2% 10.0% 10.1% 7068 6.01 14.2% 97.8* -2 0.799 1012
2.63 13673 9561 10662 89.7% 20.8% 20.8% 8185 3.10 29.4% 88.8* -4 0.776 1274
2.44 14360 9852 11510 85.6% 34.5% 34.7% 8981 1.90 48.8% 75.9* 2 0.765 1525
2.28 9501 6495 12416 52.3% 61.8% 59.6% 5991 1.14 87.3% 53.5* -2 0.722 1031
2.15 4567 3307 13210 25.0% 120.5% 120.3% 2520 0.59 170.4% 21.9* 4 0.693 408
total 86470 60274 81744 73.7% 6.7% 6.7% 51769 8.97 9.5% 99.5* 2 0.804 8369

## Slide 167: XDS: CORRECT.LP or XSCALE.LP
******************************************************************************
CORRECTION PARAMETERS FOR THE STANDARD ERROR OF REFLECTION INTENSITIES
******************************************************************************
The variance v0(I) of the intensity I obtained from counting statistics is
replaced by v(I)=a*(v0(I)+b*I^2). The model parameters a, b are chosen to
minimize the discrepancies between v(I) and the variance estimated from
sample statistics of symmetry related reflections. This model implicates
an asymptotic limit ISa=1/SQRT(a*b) for the highest I/Sigma(I) that the
experimental setup can produce (Diederichs (2010) Acta Cryst D66, 733-740).
a b ISa
3.806E+00 1.080E-04 49.32

## Slide 168: What you know when:
Autoindexing: Lattice (14)
Scaling: Point Group (43)
Phasing: Space Group (84)

## Slide 169: Asymmetric Unit (ASU)
Smallest part of the crystal
that you can use to build up the rest of the crystal
with symmetry operations
Unit Cell
Smallest part of the crystal
that you can use to build up the rest of the crystal
with translation symmetry only
“NCS ASU”
Builds ASU with Non-crystallographic symmetry operations