## Slide 1: PowerPoint File available:
http://bl831.als.lbl.gov/
~jamesh/powerpoint/
CSHL_tipsNtrick_2017.pptx

## Slide 2: Acknowledgements
UCSF LBNL SLAC
ALS 8.3.1: TomAlberTron
UC Multicampus Research Programs and Initiatives (MRPI)
NIH NIGMS R01 GM124149
NIH NIGMS P30 GM124169
Integrated Diffraction Analysis Technologies (IDAT)
Plexxikon, Inc.
Synchrotron Radiation Structural Biology Resource (SLAC)
The Advanced Light Source is supported by the Director, Office of Science, Office of Basic Energy Sciences, Materials Sciences Division, of the US Department of Energy under contract No. DE-AC02-05CH11231 at Lawrence Berkeley National Laboratory.
Robert Stroud James Fraser John Spence
Chris Nielsen Clemens Schulze-Briese Aina Cohen Ana Gonzalez

## Slide 3: The “success rate” of structure determination
100 s/dataset
200 days/year
~150 beamlines
~26,000,000 datasets/year
~10,000 PDBs in 2016

## Slide 4: signal vs noise
“If you don’t have good data, then you have no data at all.”
-Sung-Hou Kim

## Slide 5: signal vs noise
easy
hard
impossible

## Slide 6: signal vs noise
easy
hard
impossible
threshold of “solvability”

## Slide 7: signal vs noise
“If you don’t have good data, then you must learn statistics.”
-James Holton

## Slide 8: Adding noise

## Slide 9: Adding noise
12 + 12 = 1.42

## Slide 10: Adding noise
12 + 12 = 1.42
32 + 12 = 3.22
σtotal2 = σ12 + σ22

## Slide 11: Adding noise
12 + 12 = 1.42
32 + 12 = 3.22
σtotal2 = σ12 + σ22

## Slide 12: Adding noise
12 + 12 = 1.42
32 + 12 = 3.22
σtotal2 = σ12 + σ22

## Slide 13: Adding noise
12 + 12 = 1.42
32 + 12 = 3.22
102 + 12 = 10.052

## Slide 14: The “threshold of solvability” is sharp!
How can we predict success/failure?
Know Thy Experiment

## Slide 15: What do we want?
Resolution
Mosaicity
Nice, round spots
No overlaps
No twinning
Lots of crystals
…

## Slide 17: What do we want?
Resolution
Mosaicity
Nice, round spots
No overlaps
No twinning
Lots of crystals
…
2. Phases

## Slide 18: Phases
http://bl831.als.lbl.gov/~jamesh/movies/dephase.mpeg

## Slide 19: And the dominant source of error is…
Resolution
  Disorder (including Radiation Damage)
  Background
Phases
  Non-isomorphism (including Radiation Damage)
  Vibration (beam, xtal, spindle, etc.)
  Detector calibration

## Slide 20: What is “disorder”?

## Slide 21: What is “disorder”?

## Slide 22: What is “disorder”?

## Slide 23: What is “disorder”?
order
order
disorder
B-factor

## Slide 24: “B” factors
ATOM 122 N LEU A 13 -3.244 25.808 19.998 1.00 16.96 N
ATOM 123 CA LEU A 13 -2.877 25.448 21.355 1.00 15.29 C
ATOM 124 C LEU A 13 -2.792 23.966 21.561 1.00 17.54 C
ATOM 125 O LEU A 13 -1.814 23.493 22.143 1.00 16.35 O
ATOM 126 CB LEU A 13 -3.907 26.164 22.268 1.00 18.72 C
ATOM 127 CG LEU A 13 -3.577 25.982 23.738 1.00 21.19 C
ATOM 128 CD1 LEU A 13 -2.283 26.820 24.019 1.00 19.43 C
ATOM 129 CD2 LEU A 13 -4.702 26.474 24.639 1.00 24.65 C
ATOM 130 N SER A 14 -3.677 23.149 20.979 1.00 15.96 N
ATOM 131 CA SER A 14 -3.646 21.711 21.061 1.00 18.26 C
ATOM 132 C SER A 14 -2.373 21.203 20.360 1.00 18.71 C
ATOM 133 O SER A 14 -1.747 20.315 20.930 1.00 17.47 O
ATOM 134 CB SER A 14 -4.875 21.077 20.419 1.00 17.62 C
ATOM 135 OG ASER A 14 -4.825 19.665 20.388 0.50 20.89 O
ATOM 136 OG BSER A 14 -6.027 21.408 21.164 0.50 18.67 O
ATOM 137 N LYS A 15 -2.045 21.772 19.215 1.00 18.03 N
ATOM 138 CA LYS A 15 -0.799 21.361 18.555 1.00 18.12 C
ATOM 139 C LYS A 15 0.446 21.707 19.351 1.00 18.81 C
ATOM 140 O LYS A 15 1.400 20.948 19.411 1.00 17.77 O
ATOM 141 CB LYS A 15 -0.700 22.034 17.177 1.00 14.49 C
ATOM 142 CG LYS A 15 -1.727 21.368 16.256 1.00 16.12 C
ATOM 143 CD LYS A 15 -1.663 22.147 14.936 1.00 19.40 C
ATOM 144 CE ALYS A 15 -2.725 21.614 13.986 0.50 17.42 C
ATOM 145 CE BLYS A 15 -1.750 21.211 13.750 0.50 17.01 C
ATOM 146 NZ ALYS A 15 -2.346 21.674 12.559 0.50 18.61 N
ATOM 147 NZ BLYS A 15 -3.052 20.513 13.741 0.50 18.76 N

## Slide 25: The B factor
Bragg, W. H. (1914)."XCVII. The intensity of reflexion of X rays by crystals", Philosophical Magazine Series 6 27, 881-899. Published May 1 1914

## Slide 26: B factor from image analysis
B = 500

## Slide 27: B factor from image analysis
B = 20

## Slide 28
Crystal diameter (micron)
Resolution (Å)
B = 0
B = 20
B = 60
B = 170
B = 330
B = 560
B = 850
Bigger is better, but not by much

## Slide 29
disorder
+ averaging
= poor resolution
Diffraction Take-Home Lesson:

## Slide 30: What Causes Disorder?
Function
Radiation damage
Mosaicity
Strain
Xtal abuse
Dehydration

## Slide 31: What Causes Disorder?
Function
Radiation damage
Mosaicity
Strain
Xtal abuse
Dehydration

## Slide 32: Muybridge’s galloping horse (1878)

## Slide 33: realistic “crystal” of horses

## Slide 34: average structure: galloping horse

## Slide 35: Ways to improve order
New crystal form
Purify
  - add a column
  - heat shock
Bind something
  - Known ligand
  - Silver bullets
Change the clone
  - chop off floppy bits
  - homolog
  - Ala -> Lys mutants
Make Protein Sit Still!!!
Newman J. (2006) Acta Cryst. D62 27-31.
8000 cases in PDB of resolution improvement
80% of these are different crystal forms
> 1 Å improvement → 94% different forms

## Slide 36: What Causes Disorder?
Function
Radiation damage
Mosaicity
Strain
Xtal abuse
Dehydration

## Slide 37: Types of radiation damage
Global
Specific

## Slide 38: Types of radiation damage
Global
  scale factor
  B factor
  non-isomorphism
Specific
  depends on chemistry
  destroys phasing signal

## Slide 39
Global damage
Howells, et al. (2009) J. Electron Spectrosc. Relat. Phenom. 170, 4-12.

## Slide 40: resolution dependence of global damage
resolution (Å)
maximum tolerable dose (MGy)
1 2 3 5 7 10 20 40 70 100
1
10
100
103
Howells et al. (2009) J. Electron. Spectrosc. Relat. Phenom. 170 4-12

## Slide 41: resolution dependence of global damage
resolution (Å)
maximum tolerable dose (MGy)
1 2 3 5 7 10 20 40 70 100
1
10
100
103
Howells et al. (2009) J. Electron. Spectrosc. Relat. Phenom. 170 4-12
10 MGy/Å

## Slide 42: what the is a MGy?
http://bl831.als.lbl.gov/
damage_rates.pdf
Holton J. M. (2009) J. Synchrotron Rad. 16 133-42
1 MGy = 106 J/kg

## Slide 43
How long will my crystal last?
synch line type flux beamsize flux density dose max xtal min site
ph/s μm ph/μm2/s rate lifetime lifetime
ALS 4.2.2 MAD 2.2e11 75x80 3.7e+07 27.3 kGy/s 18 m 73 s
ALS 5.0.1 mono 1.6e11 100 2.0e+07 10.4 kGy/s 48 m 3.2 m
ALS 5.0.2 MAD 8e11 100 1.0e+08 51.8 kGy/s 9.6 m 39 s
ALS 5.0.3 mono 1.7e11 100 2.2e+07 11 kGy/s 45 m 3 m
ALS 8.2.1 MAD 1.8e11 100 2.3e+07 11.7 kGy/s 43 m 2.9 m
ALS 8.2.2 MAD 2.3e11 100 2.9e+07 14.9 kGy/s 34 m 2.2 m
ALS 8.3.1 MAD 9e11 70 2.3e+08 119 kGy/s 4.2 m 17 s
ALS 8.3.1 typical 6e11 70 1.6e+08 115 kGy/s 4.3 m 17 s
ALS 12.3.1 MAD 1.8e11 100 2.3e+07 11.7 kGy/s 43 m 2.9 m
ALS 12.3.1 ML 4.0e12 100 5.1e+08 513 kGy/s 58 s 3.9 s
APS 8-BM MAD 1e11 200 2.5e+06 1.27 kGy/s 6.6 h 26 m
APS 14-BM-C mono 5.8e10 200 1.4e+06 738 Gy/s 11 h 45 m
APS 14-BM-D MAD 3.3e9 200 8.2e+04 42 Gy/s 8.3 d 13 h
APS 14-ID-B MAD 6.0e10 200 1.5e+06 763 Gy/s 11 h 44 m
APS 17-BM MAD 1.1e11 200 2.8e+06 1.4 kGy/s 6 h 24 m
APS 17-ID MAD 2.3e11 200 5.8e+06 2.93 kGy/s 2.8 h 11 m
APS 19-BM MAD 2.0e11 70x60 4.8e+07 24.2 kGy/s 21 m 83 s
APS 19-ID MAD 1.3e13 80x40 4.1e+09 2.07 MGy/s 15 s 0.97 s
APS 19-ID typical 5.5e11 100x100 5.5e+07 28 kGy/s 18 m 71 s
APS 22-BM MAD 7e12 80x40 2.2e+09 1.23 MGy/s 24 s 1.6 s
APS 22-ID MAD 7e12 80x40 2.2e+09 1.23 MGy/s 24 s 1.6 s
APS 22-ID typical 1.5e12 80 2.3e+08 119 kGy/s 4.2 m 17 s
APS 23-ID-B MAD 1e13 75x25 5.3e+09 3.01 MGy/s 10 s 0.66 s
APS 23-ID typical 1.5e12 80 2.3e+08 119 kGy/s 4.2 m 17 s
APS 24-ID-C MAD 1.3e13 20x60 1.1e+10 5.23 MGy/s 5.7 s 0.38 s
APS 24-ID-E MAD 0.5e13 20x100 2.5e+09 1.19 MGy/s 25 s 1.7 s
APS 31-ID MAD 2e12 70 4.1e+08 194 kGy/s 2.6 m 10 s

## Slide 45
Holton & Frankel (2010) Acta D 66 393-408.

## Slide 46: Specific damage
final, refined phases

## Slide 47: Specific damage
experimentally-obtained phases

## Slide 48
Holton, J. M. (2007) J. Synch. Rad. 14, 51-72.

## Slide 49
1.0
0.0
fraction
unconverted
Holton, J. M. (2007) J. Synch. Rad. 14, 51-72.

## Slide 50
Holton (2009) J. Synchrotron Rad. 16 133-42
Radiation Damage World Records
MGy reaction reference
~45 global damage Owen et al. (2006)
10/Å global damage Howells et al. (2009)
5 Se-Met Holton (2007)
4 Hg-S Ramagopal et al. (2004)
4 R-C-COOH Garman et al. (2015)
3 S-S Murray et al. (2002)
1 Br-RNA Olieric et al. (2007)
~1? Cl-C ???
0.5 Mn in PS II Yano et al. (2005)
0.06 putidaredoxin Corbett et al. (2007)
0.02 Fe in myoglobin Denisov et al. (2007)

## Slide 51: Life-halving concentration
Na | 44
Mg | 27
P | 9
S | 7
Cl | 5
K | 3.3
Ca | 2.7
Fe | 1.0
Cu | 0.67
Zn | 0.58
molar, at the Se edge based on RADDOSE
As | 0.44
Se | 0.42
Br | 2.9
I | 0.47
Gd | 0.21
Ta | 0.13
Pt | 0.17
Au | 0.16
Hg | 0.16
U | 0.23

## Slide 52: Life-halving concentration
Na | 44
Mg | 27
P | 9
S | 7
Cl | 5
K | 3.3
Ca | 2.7
Fe | 1.0
Cu | 0.67
Zn | 0.58
molar, at the Se edge based on RADDOSE
As | 0.44
Se | 0.42
Br | 2.9
I | 0.47
Gd | 0.21
Ta | 0.13
Pt | 0.17
Au | 0.16
Hg | 0.16
U | 0.23

## Slide 53: No Damage in the Dark!
Experiment time (hours)
Peak absorbance (electrons)
4.6 5.0 5.4 5.8 6.2 6.6
0 5 10 15 20 25

## Slide 54
The number of photons scattered
before crystal is dead
is independent
of flux & time
1 um3 = 105 photons (roomT)
1 um3 = 106 photons (synch)
1 um3 = 108 photons (XFEL)
Henderson, 1990; Gonzalez & Nave, 1994; Glaeser et al., 2000; Sliz et al., 2003; Leiros et al., 2006; Owen et al., 2006; Garman & McSweeney, 2006; Garman & Nave, 2009; Holton, 2009
(synch)

## Slide 55: Dose slicing
crystal’s useful life
N
photons
N
photons
N
photons
unacceptable
completeness
unacceptable
read noise

## Slide 56: Dose slicing: is that a spot?

## Slide 57: Dose slicing: is that a spot?

## Slide 58: Dose slicing: is that a spot?

## Slide 59: Dose slicing: is that a spot?

## Slide 60: Dose slicing: is that a spot?

## Slide 61: Dose slicing: is that a spot?

## Slide 62: What if?
You only have a few small crystals…
Should you:
  Collect 360° from each?
  Collect 10° from each at 36x exposure?
  Glue 36 xtals together, then collect 360° ?
  Glue, and do 12960° faint exposures?

## Slide 63
Chart: 
Resolution (Å)
CC vs right answer
“true” resolution limit

## Slide 64
Dose slice collection scenario |  |  | 
fine | coarse | glue | Glue & fine
3.17700 | 3.14057 | 3.17604 | 3.12086
“true” resolution vs strategy
301,640,334 photons/xtal

## Slide 65
processing procedure | resolution limit |  |  | 
 | fine | coarse | glue | glue&fine
XDS/XSCALE | 3.17700 | 3.14057 | 3.17604 | 3.12086
XDS/aimless | 3.16251 | 3.15922 | 3.16675 | 3.12162
XDS/noscale | 3.27289 | 3.15617 | 3.38929 | 3.12162
MOSFLM-defaults | x | x | x | x
MOSFLM-cheat | 3.19755 | 3.18166 | 3.20354 | 3.19755
HKL2000-defaults | x | x | x | x
HKL2000-expert | 3.21803 | 3.12646 | 3.11551 | 3.21803
HKL2000-cheat | 3.16119 | 3.10338 | 3.10338 | 3.16300
DIALS-defaults | x | x | 3.42125 | x
DIALS-expert | 3.10296 | 3.09547 | 3.07653 | 3.08465
d*Trek | ? | ? | 4.34076 | ?
EVAL15 | 3.14352 | 3.10383 | 3.09936 | 3.14105
EVAL15-sadabs | 3.13920 | 3.10797 | 3.10171 | 3.13700

## Slide 66: Minimum exposure time:
~1 photon/pixel
  PAD w XDS, DIALS
~10 photon/pixel
PAD w HKL2k, MOSFLM
~30 photon/pixel
CCD detector
adjust exposure
so this is ~15
(Rayonix)

## Slide 67: What Causes Disorder?
Function
Radiation damage
Mosaicity
Strain
Xtal abuse
Dehydration

## Slide 68
Ewald’s “mosaic” picture

## Slide 69
Ewald’s “mosaic” picture

## Slide 70
Ewald’s “mosaic” picture
What is
this stuff?

## Slide 71
Darwin’s original picture

## Slide 72
“mosaicity” with visible light

## Slide 73: mosaic spread = 0 º

## Slide 74: mosaic spread = 0.1º

## Slide 75: mosaic spread = 0.2º

## Slide 76: mosaic spread = 0.4º

## Slide 77: mosaic spread = 0.6º

## Slide 78: mosaic spread = 0.8º

## Slide 79: mosaic spread = 1.0º

## Slide 80: mosaic spread = 1.5º

## Slide 81: mosaic spread = 2.0º

## Slide 82: mosaic spread = 2.5º

## Slide 83: mosaic spread = 3.2º

## Slide 84: mosaic spread = 6.4º

## Slide 85: mosaic spread = 12.8º

## Slide 86: What Causes Disorder?
Function
Radiation damage
Mosaicity
Strain
Xtal abuse
Dehydration

## Slide 87
stress and strain
radiation damage = defects
What about
undamaged crystals?

## Slide 88: Purity is crucial!
McPherson, A., Malkin, A. J., Kuznetsov, Y. G. & Plomp, M. (2001)."Atomic force microscopy applications in macromolecular crystallography", Acta Cryst. D 57, 1053-1060.
not important for
initial hits
important for
resolution

## Slide 89: What can I improve?
Purity!
is 95% good enough? 99%?
Purity!
conformational (homogeneous)
Purity!
kinetic (stable over time)

## Slide 90: What Causes Disorder?
Function
Radiation damage
Mosaicity (η) > ~3 μm
Strain (B) < ~3 μm
Xtal abuse
Dehydration

## Slide 91: What Causes Disorder?
Function
Radiation damage
Mosaicity > ~3 μm
Strain < ~3 μm
Xtal abuse
Dehydration

## Slide 92: Xtal Abuse
physical contact
don’t touch the part you intend to shoot
osmotic shock
equilibrate, or calculate matching solution
changes in dielectric constant
Petsko (1975) J. Mol. Biol. 96, 381-388.
cooled density mismatch
Juers & Matthews (2004) Acta Cryst. D 60, 412-421.
basically: no sudden moves!

## Slide 93: What Causes Disorder?
Function
Radiation damage
Mosaicity > ~3 μm
Strain < ~3 μm
Xtal abuse
Dehydration

## Slide 94: opened drop: you have ~30 seconds!
equilibrated drop

## Slide 95: First diffraction from protein xtal
Bernal, J. & Crowfoot, D. (1934). "X-ray photographs of crystalline pepsin", Nature 133, 794-795.
1934

## Slide 96: Dehydration: 1934
100 μm
2 mm
= 1.5 μL
= 1.0 nL
10 μm
= 1.0 pL
and 2017?
1 μm
= 1.0 fL

## Slide 97
Dear James
The story of the two forms of lysozyme crystals goes back to about 1964 when it was found that the diffraction patterns from different crystals could be placed in one of two classes depending on their intensities. This discovery was a big set back at the time and I can remember a lecture title being changed from the 'The structure of lysozyme' to 'The structure of lysozyme two steps forward and one step back'. Thereafter the crystals were screened based on intensities of the (11,11,l) rows to distinguish them (e.g. 11,11,4 > 11,11,5 in one form and vice versa in another). Data were collected only for those that fulfilled the Type II criteria. (These reflections were easy to measure on the linear diffractometer because crystals were mounted to rotate about the diagonal axis). As I recall both Type I and Type II could be found in the same crystallisation batch . Although sometimes the external morphology allowed recognition this was not infallible.
The structure was based on Type II crystals. Later a graduate student Helen Handoll examined Type I. The work, which was in the early days and before refinement programmes, seemed to suggest that the differences lay in the arrangement of water or chloride molecules (Lysozyme was crystallised from NaCl). But the work was never written up. Keith Wilson at one stage was following this up as lysozyme was being used to test data collection strategies but I do not know the outcome.
An account of this is given in International Table Volume F (Rossmann and Arnold edited 2001) p760.
Tony North was much involved in sorting this out and if you wanted more info he would be the person to contact. I hope this is helpful. Do let me know if you need more.
Best wishes
Louise
Non-isomorphism in lysozyme

## Slide 98
Johnson’s ratio
Structure factor (e-)
Non-isomorphism in lysozyme

## Slide 99
Johnson’s ratio
Structure factor (e-)
Non-isomorphism in lysozyme

## Slide 101: Non-isomorphism in lysozyme
RH: 84.2%
vs 71.9%
Riso = 44.5%
3aw6
3aw7
Δcell = 0.7 %
RMSD = 1.70 Å
RMSD = 0.18 Å

## Slide 102
0.5% cell change → 15%
Crick & Magdoff (1956) “theory method isomorphous replacement”, Acta Cryst. 9, 901-8.
Same structure, different cell

## Slide 103
Same structure, different cell
0.5% cell change → 15%
Crick & Magdoff (1956) “theory method isomorphous replacement”, Acta Cryst. 9, 901-8.

## Slide 104
Same cell, rotated protein
0.5° rotation → 16%
Crick & Magdoff (1956) “theory method isomorphous replacement”, Acta Cryst. 9, 901-8.

## Slide 105
Same cell, rotated protein
0.5° rotation → 16%
Crick & Magdoff (1956) “theory method isomorphous replacement”, Acta Cryst. 9, 901-8.

## Slide 106
Uniform stretch
5% change → 0%

## Slide 107
5% change → 0%
Uniform stretch

## Slide 108
Magdoff & Crick (1955) “Ribonuclease II. Accuracy of measurement and shrinkage”, Acta Cryst. 8, 461-8.
Crick & Magdoff (1956) “The theory of the method of isomorphous replacement for protein crystals”, Acta Cryst. 9, 901-8.

## Slide 109: Elastic deformation = non-isomorphism

## Slide 110: Elastic deformation = non-isomorphism

## Slide 111: Plastic deformation = poor diffraction

## Slide 112: Plastic deformation = poor diffraction

## Slide 113: OK! OK! Now I am afraid!
What should I do?

## Slide 114: Basic Principles
“Hell, there are NO RULES here - we're trying to accomplish something.”
Thomas A. Edison – inventor
“You’ve got to have an ASSAY.”
Arthur Kornberg – Nobel Laureate
“Control, control, you must learn CONTROL!”
Yoda – Jedi Master

## Slide 115: The Best Assay

## Slide 116: And the dominant source of error is…
Resolution
  Disorder (including Radiation Damage)
  Background
Phases
  Non-isomorphism (including Radiation Damage)
  Vibration (beam, xtal, spindle, etc.)
  Detector calibration

## Slide 117: beam size vs xtal size
  Put your crystal into the beam

## Slide 118: Get thee to a microbeam?
Evans et al. (2011)."Macromolecular microcrystallography", Crystallography Reviews 17, 105-142.

## Slide 119: beam size vs xtal size
  Put your crystal into the beam
  Shoot the whole crystal

## Slide 120: shoot the whole crystal

## Slide 121: shoot the whole crystal

## Slide 122: shoot the whole crystal

## Slide 123: beam size vs xtal size
  Put your crystal into the beam
  Shoot the whole crystal
  Shoot nothing but the crystal

## Slide 124: shoot nothing but the crystal

## Slide 125: shoot nothing but the crystal

## Slide 126: X-ray scattering “rules”:
1 μm crystal ≈ 1 μm water
≈ 1 μm plastic
≈ 0.1 μm glass
≈ 1000 μm air

## Slide 127
X-ray beam size (μm)
Resolution (Å)
too big
too small
just right

## Slide 128: Gaussian beam
50% of flux
the other 50%

## Slide 129: Collimated beam
100% of flux
0%

## Slide 130: Tipping point for µfocus:
10 µm
~10 mm
~1 mrad =0.06°

## Slide 131: beam size vs xtal size
  Put your crystal into the beam
  Shoot the whole crystal
  Shoot nothing but the crystal
  Back off!

## Slide 132: Background scattering

## Slide 133: Background scattering
$100,000.00
$100,000.00
$100,000.00
$100,000.00
$100,000.00
$100,000.00
$100,000.00
$100,000.00
real estate is
expensive
use it!

## Slide 134: Fine Slicing
Pflugrath, J. W. (1999)."The finer things in X-ray diffraction data collection", Acta Cryst. D 55, 1718-1725.
background
background

## Slide 135: beam size vs xtal size
  Put your crystal into the beam
  Shoot the whole crystal
  Shoot nothing but the crystal
  Back off!
  The crystal must rotate

## Slide 136: Get thee to a microbeam?
Evans et al. (2011)."Macromolecular microcrystallography", Crystallography Reviews 17, 105-142.

## Slide 137: naked xtal

## Slide 138: And the dominant source of error is…
Resolution
  Disorder (including Radiation Damage)
  Background
Phases
  Non-isomorphism (including Radiation Damage)
  Vibration (beam, xtal, spindle, etc.)
  Detector calibration

## Slide 139
anomalous signal
Crick, F. H. C. & Magdoff, B. S. (1956) Acta Crystallogr. 9, 901-908.
Hendrickson, W. A. & Teeter, M. M. (1981) Nature 290, 107-113.
# sites
MW (Da)
ΔF
F
≈ 1.2 f”
√
World record!
ΔF/F = 0.5%
Wang, Dauter & Dauter (2006) Acta Cryst. D 62, 1475-1483.

## Slide 140: Riso vs dose
Riso (%)
change in dose (MGy)
data taken from Banumathi, et al. (2004) Acta Cryst. D 60, 1085-1093.

## Slide 141: Riso vs dose
Riso (%)
change in dose (MGy)
data taken from Banumathi, et al. (2004) Acta Cryst. D 60, 1085-1093.
Riso ≈ 0.7 %/MGy

## Slide 142
Required signal-to-noise (I/σ)
Solve-able proteins (%)
CCD
old Pilatus
Goal
Threshold of a revolution in phasing
Best
Pilatus

## Slide 143: S-SAD merging 20 xtals
Knott et al. (2017). "Guide-bound CRISPR-Cas13a." Nature S&MB online

## Slide 144
Chart: 
| 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10
F+ | 52.3965 | 57.5371 | 48.523 | 47.5964 | 34.153 | 48.9106 | 70.3415 | 52.7871 | 59.3327 | 50.0622
F- | 46.3965 | 51.5371 | 42.523 | 41.5964 | 28.153 | 42.9106 | 64.3415 | 46.7871 | 53.3327 | 44.0622
observation
intensity
Systematic vs random error

## Slide 145: And the dominant source of error is…
Resolution
  Disorder (including Radiation Damage)
  Background
Phases
  Non-isomorphism (including Radiation Damage)
  Vibration (beam, xtal, spindle, etc.)
  Detector calibration

## Slide 146: Optimal exposure time (anomalous on a PAD)
time
intensity
2.03 ms

## Slide 147: PAD full vs partial
iteration
Summed intensity

## Slide 148: Optimal exposure time (anomalous on a PAD)
tmad Optimal exposure time (s)
tro read-out time (s)
2.03 ms on Pilatus3 S
0.95 ms on Pilatus3 X
0.003 ms on Eiger
Attenuate for MAD
PILATUS
>0.2 s/frame
EIGER
>0.0003 s/frame
= 20,000 frames
350 GB/dataset

## Slide 149: And the dominant source of error is…
Resolution
  Disorder (including Radiation Damage)
  Background
Phases
  Non-isomorphism (including Radiation Damage)
  Vibration (beam, xtal, spindle, etc.)
  Detector calibration

## Slide 150: Detector calibration: 7235 eV

## Slide 151: Detector calibration: 7247 eV

## Slide 152: Gadox calibration vs energy
photon energy (keV)
Relative absorption depth
same = good!
bad!

## Slide 153
Pick-up tool mark
Bragg
glitch
oxygen
inclusions
“tree rings”
1% high
average
1% low
~3x105 photon/pixel Pilatus detector
Spatial Heterogeneity in Sharp Spot Sensitivity

## Slide 154: Systematic error does not “average out”

## Slide 155: Fractional error
mult > (—)2
~3%
<ΔF/F>

## Slide 156
Holton & Frankel (2010) Acta D 66 393-408.

## Slide 157: Data collection parameters:
16 crystals
360° each, inverse beam
7235 eV photon energy
< 1 MGy per xtal
Australian Synchrotron MX1
  35 kGy/s into 100 μm x 100 μm

## Slide 158
RESOLUTION COMPLETENESS R-FACTOR I/SIGMA R-meas CC(1/2) Anomal SigAno Nano
LIMIT OF DATA observed Corr
9.17 99.1% 3.9% 257.47 3.9% 100.0* 91* 5.024 450
6.49 100.0% 5.2% 214.33 5.2% 100.0* 86* 3.836 882
5.30 100.0% 7.2% 165.13 7.2% 100.0* 76* 3.257 1175
4.59 100.0% 7.2% 175.42 7.3% 100.0* 67* 2.589 1403
4.10 99.9% 7.7% 174.13 7.7% 100.0* 59* 2.264 1594
3.74 99.9% 9.4% 143.09 9.4% 100.0* 49* 1.953 1783
3.47 100.0% 11.2% 120.17 11.2% 100.0* 39* 1.696 1942
3.24 100.0% 14.1% 91.14 14.1% 100.0* 30* 1.333 2103
3.06 99.9% 19.5% 65.79 19.5% 100.0* 23* 1.117 2214
2.90 99.9% 29.0% 44.85 29.1% 99.9* 17* 1.008 2369
2.77 99.9% 40.5% 32.58 40.6% 99.8* 11* 0.901 2493
2.65 99.9% 52.8% 25.16 52.9% 99.8* 10* 0.866 2605
2.54 100.0% 67.4% 19.47 67.6% 99.6* 2 0.804 2705
2.45 100.0% 88.9% 14.58 89.2% 99.2* 4 0.831 2859
2.37 100.0% 109.3% 9.97 109.7% 98.1* 5 0.829 2925
2.29 100.0% 138.2% 6.87 138.9% 96.1* 1 0.760 3037
2.22 100.0% 197.1% 4.03 198.6% 83.5* -1 0.721 3159
2.16 100.0% 227.3% 2.41 230.8% 46.9* -1 0.677 3224
2.10 61.2% 154.4% 1.28 163.6% 47.0* -2 0.660 1999
2.05 47.9% 170.1% 0.68 196.5% 25.7* 3 0.629 1578
total 93.3% 15.7% 54.30 15.8% 100.0* 12* 1.217 42499
140-fold multiplicity: 16 crystals, 360° each, inverse beam, 7235 eV

## Slide 159: 140-fold multiplicity
18 σ
Phased anomalous difference Fourier
16 σ

## Slide 160: 140-fold multiplicity
15σ = PO4
Phased anomalous difference Fourier

## Slide 161: 140-fold multiplicity
~2σ = Mg?
Phased anomalous difference Fourier

## Slide 162: 140-fold multiplicity
8.2σ = Mg?
DELFAN residual anomalous difference

## Slide 163: Discerning Na+ from Mg++
f’’ (electrons)
DELFAN peak height (σ)
Mg
Ne
Na
F
O
N

## Slide 164: 140-fold multiplicity
7.4σ = Na
DELFAN residual anomalous difference

## Slide 165: Suggested anomalous protocol:
2 wavelengths are better than 1
- (peak + inf)/2, and remote
MAD, not M-SAD!
360° in < 5 MGy
move detector
4X exposure
goto 2

## Slide 166: Summary
100 ADU/pixel
10 μm for lysozyme
~3% error per spot, 1%/MGy
7235 eV for S-SAD
“Attenu-wait”: dose slicing
http://bl831.als.lbl.gov/xtalsize.html
http://bl831.als.lbl.gov/~jamesh/mlfsom/
http://bl831.als.lbl.gov/~jamesh/powerpoint/ACMC_GBU_2013.ppt

## Slide 167: Summary
15 ADU/pixel on Rayonix HS 300
> 0.1 s exposures (HS 300)
10 μm for lysozyme
~3% error per spot, 1%/MGy
7235 eV for S-SAD on CCD
“Attenu-wait”: dose slicing
http://bl831.als.lbl.gov/xtalsize.html
http://bl831.als.lbl.gov/xtallife.html
http://bl831.als.lbl.gov/~jamesh/damage_rates.pdf
http://bl831.als.lbl.gov/~jamesh/powerpoint/CSHL_tipsNtrick_2017.pptx

## Slide 168: Why harvest at all?
Radiation damage
  10-100x reduced at ~100 K
Background scatter
    tray >> xtal
Timing
    nothing lasts forever
Shipping
    cryo samples are glass beads

## Slide 169: Why harvest at all?
Radiation damage
  10-100x reduced at ~100 K
Background scatter
    tray >> xtal
Timing
    nothing lasts forever
Shipping
    cryo samples are glass beads

## Slide 170: X-ray transmission of materials
Depth into material (μm)
transmission
glass
protein
crystal
water
air

## Slide 171: X-ray scattering of materials
Differential cross section
(scattered photons/solid angle/incident photon/meter thickness)
 7.5 3.8 2.5 1.9 1.5
d-spacing (Ǻ)
water
25% glycerol
paratone-N oil
plastic
glass

## Slide 172: Crowfoot Cell

## Slide 173: Crowfoot Cell

## Slide 174: Crowfoot Cell
J. D. Bernal and D. Crowfoot, Nature (London) 133, 794 (1934).

## Slide 175
pinhole
mirror
microscope
SBS
format
crystallization
tray
Notes: Inspired by the ESRF microcrystal camera, the superbend endstations are all equipped with zero-paralax optics. The user sees their sample from the point of view of the x-ray beam. Sample alignment is not only more stable but more intuitive than an off-axis camera. Placing the crystal into the microscope crosshairs gaurantees that it will be hit by the beam.

## Slide 176: Plate goniometers
Notes: The non-rotating “plate goniometer” (left) was developed last year. A new incarnation for taking rotation-camera data is now being developed (right). Either of these devices will accept most any crystallization tray that can be rotated 90 degrees to gravity and probe arbitrary points in each crystallization drop with a 50 um x-ray beam. Installing a plate goniometer takes ~30 minutes and the location ~300 mm downbeam from the “normal” protein crystallography sample mounting point makes this technology transferable to most any beamline.
Current status is that a 95% complete data set has been collected from a 50 um lysozyme crystal at room temperature, and trials are now underway to collect data sets from “real” crystals.

## Slide 177
Notes: non-rotating tray goniometer stage installed in ALS 8.3.1. Note that the “normal” cryo-pin goniometer on the left is still in place and a brass flight tube has been added to contain the air scatter from the main beam on its way to the tray goniometer.

## Slide 179
Fluidigm
Topaz® DC
http://www.fluidigm.com/home/fluidigm/docs/Datasheet_1.96DC.pdf
200 μm
Soliman, A. et al. (2011). Acta Cryst. D 67, 646-656.
250 μm
NatX-ray
CrystalQuick™ X
~50 μm tape
MiTeGen
MicroRT™
25 μm wall
25 μm wall
http://www.mitegen.com/products/micrort/micrort.shtml
PDMS
PDMS
http://www.mitegen.com/mic_catalog.php?c=insituplates
100 μm
MiTeGen
in-situ-1
~50 μm tape
in-situ diffraction systems
Márquez, J. & Cipriani, F. (2014) Structural Genomics (Chen, Y. W., ed) pp. 197-203.
Crystal DirectTM
~50 μm tape
25 μm
Notes: Next slide please

## Slide 180: What do we need?
Thin plastic = water = protein
Light low atomic number
Tight evaporation = nonisomorphism
Cheap $1-20/tray

## Slide 181: What we really need
GRAPHENE
Thin
Light
Tight
Cheap

## Slide 182: Why harvest at all?
Radiation damage
  10-100x reduced at ~100 K
Background scatter
    tray >> xtal
Timing
    nothing lasts forever
Shipping
    cryo samples are glass beads

## Slide 184
8000 Gy/s

## Slide 185: Example Data
lysozyme
50 μm beam
37 Gy/s (0.775 Å)
30s exposures at 20C
90° of data, 97% complete
I/ = 1.5 at 1.9 Å
Rmerge 18% (overall) 5% (low)
ΔB same as 20 min at 100K
B factor
-2
-4
-6
scale vs batch
2
1
0

## Slide 186: Why harvest at all?
Radiation damage
  10-100x reduced at ~100 K
Background scatter
    tray >> xtal
Timing
    nothing lasts forever
Shipping
    cryo samples are glass beads

## Slide 187: Why harvest at all?
Radiation damage
  10-100x reduced at ~100 K
Background scatter
    tray >> xtal
Timing
    nothing lasts forever
Shipping
    cryo samples are glass beads

## Slide 188: MiTeGen RT cell
http://www.mitegen.com/products/micrort/micrort.shtml

## Slide 189: Summary
assay, control and open mind
purity, purity, purity
oil is your friend
Warkentin method
try room temp!
http://bl831.als.lbl.gov/~jamesh/powerpoint/
CSHL_tipsNtricks_2015.pptx