## Slide 1: Data processing
How did we get here?
https://bl831.als.lbl.gov/~jamesh/powerpoint/DiffMet_2024.pptx

## Slide 2: Key advances to date
Understanding radiation damage
Sample preparation
Sample introduction
  robots, trays & serial
Background reduction
  Elimination of “read-out” noise
  Tighter beams: size, stability and speed
  detector calibration
Population of “fold space”

## Slide 3: Radiation damage

## Slide 4: Xtal5_t1 graph

## Slide 5: Dose-rate dependence of damage
dose rate (kGy/s)
maximum useful dose (MGy)

## Slide 6: Key advances to date
Understanding radiation damage
Sample preparation
Sample introduction
  robots, trays & serial
Background reduction
  Elimination of “read-out” noise
  Tighter beams: size, stability and speed

## Slide 7: Resolution
http://bl831.als.lbl.gov/~jamesh/movies/resolution.mpeg

## Slide 8: Resolution vs B factor
1 1.5 2 2.5 3 3.5 4
140
120
100
90
60
40
20
0
1970-1995
1970-2004
2013-2023
4*d^2+12
Reported resolution (Å)
Average atomic B factor

## Slide 9
0.8
0.7
0.6
0.5
0.4
0.3
0.2
0.1
0
Average outer-shell Rmerge
Year
Rmerge at the resolution limit in PDB

## Slide 10: Expected Rmerge as Iobs → 0

## Slide 11: Optimum resolution cutoff is:
Too optimistic: add nothing but noise
Too pessimistic: series-termination error
Happy medium?
Simulate:
  Random atoms, compute F2
  Add Gaussian noise, RMS = 1
  Truncate
  Subtract “right” map, RMS difference

## Slide 12
Chart: 
| 0.8 | 0.81 | 0.82 | 0.83 | 0.84 | 0.85 | 0.86 | 0.87 | 0.88 | 0.89 | 0.9 | 0.91 | 0.92 | 0.93 | 0.94 | 0.95 | 0.96 | 0.97 | 0.98 | 0.99 | 1 | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.0900000000000001 | 1.1000000000000001 | 1.1100000000000001 | 1.1200000000000001 | 1.1299999999999999 | 1.1399999999999999 | 1.1499999999999999 | 1.1599999999999999 | 1.17 | 1.18 | 1.19 | 1.2 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.29 | 1.3
total error | 0.0063854 | 0.006377 | 0.0063704 | 0.0063613 | 0.0063546 | 0.0063466 | 0.006339 | 0.0063331 | 0.0063281 | 0.0063218 | 0.0063169 | 0.0063122 | 0.0063076 | 0.0063035 | 0.0062986 | 0.0062952 | 0.006292 | 0.0062882 | 0.0062851 | 0.0062825 | 0.0062808 | 0.0062806 | 0.0062829 | 0.0062871 | 0.0063049 | 0.0063317 | 0.0063768 | 0.006453 | 0.00655 | 0.0067899 | 0.006935 | 0.0072772 | 0.0075335 | 0.0079267 | 0.0082267 | 0.0085917 | 0.0090308 | 0.0096403 | 0.010142 | 0.010578 | 0.011234 | 0.011798 | 0.012663 | 0.013292 | 0.013934 | 0.014784 | 0.015263 | 0.016206 | 0.01711 | 0.01817 | 0.018793
series termination | 0.0001167 | 0.00014118 | 0.00016693 | 0.00020738 | 0.00024052 | 0.00028835 | 0.00034393 | 0.00040287 | 0.00046367 | 0.00054497 | 0.00062312 | 0.00072092 | 0.00082139 | 0.00091787 | 0.0010631 | 0.0011963 | 0.0013389 | 0.0015127 | 0.0017075 | 0.0019087 | 0.0021401 | 0.0023446 | 0.0026241 | 0.0028107 | 0.0031805 | 0.0034765 | 0.0037849 | 0.0041158 | 0.0044212 | 0.0049665 | 0.0052371 | 0.0057965 | 0.0061838 | 0.0067416 | 0.0071474 | 0.0076234 | 0.0081753 | 0.0089132 | 0.0095026 | 0.010003 | 0.010739 | 0.011361 | 0.012297 | 0.012969 | 0.013647 | 0.014539 | 0.015036 | 0.016012 | 0.016942 | 0.018027 | 0.018662
noise in data | 0.0063844 | 0.0063755 | 0.0063682 | 0.0063579 | 0.0063501 | 0.0063401 | 0.0063297 | 0.0063203 | 0.0063111 | 0.0062983 | 0.0062861 | 0.0062709 | 0.0062539 | 0.0062363 | 0.0062082 | 0.0061804 | 0.0061479 | 0.0061034 | 0.0060487 | 0.0059854 | 0.0059049 | 0.0058264 | 0.0057085 | 0.0056237 | 0.0054437 | 0.0052917 | 0.0051318 | 0.0049697 | 0.0048324 | 0.0046297 | 0.0045459 | 0.0043996 | 0.0043026 | 0.0041691 | 0.0040733 | 0.0039623 | 0.0038365 | 0.0036725 | 0.0035436 | 0.0034402 | 0.0032968 | 0.0031835 | 0.0030235 | 0.002914 | 0.0028109 | 0.0026816 | 0.0026168 | 0.0024966 | 0.0023891 | 0.0022725 | 0.0022113
Resolution cutoff (Å)
Error in map (rms e-)
Optimal resolution cutoff

## Slide 13
Chart: 
| 0.8 | 0.81 | 0.82 | 0.83 | 0.84 | 0.85 | 0.86 | 0.87 | 0.88 | 0.89 | 0.9 | 0.91 | 0.92 | 0.93 | 0.94 | 0.95 | 0.96 | 0.97 | 0.98 | 0.99 | 1 | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.0900000000000001 | 1.1000000000000001 | 1.1100000000000001 | 1.1200000000000001 | 1.1299999999999999 | 1.1399999999999999 | 1.1499999999999999 | 1.1599999999999999 | 1.17 | 1.18 | 1.19 | 1.2 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.29 | 1.3
total error | 0.011234 | 0.011798 | 0.012663 | 0.013292 | 0.013934 | 0.014784 | 0.015263 | 0.016206 | 0.01711 | 0.01817 | 0.018793
CC1/2 | 0.00027543 | 0.0009316 | 0.0011407 | 0.00093939 | -0.00085547 | 8.519e-05 | 0.0017707 | 0.0018655 | 0.00044742 | -0.00083644 | -8.9753e-05 | -0.0012856 | -0.0022441 | -0.0011232 | -0.0006217 | -0.00063863 | 0.00039935 | 0.0020736 | 0.0023749 | 0.001991 | 0.0015254 | 0.0011874 | 0.0012966 | 0.0070442 | 0.010659 | 0.014306 | 0.021539 | 0.031659 | 0.044804 | 0.054608 | 0.085054 | 0.10512 | 0.15123 | 0.19219 | 0.24658 | 0.30221 | 0.374 | 0.43569 | 0.50225 | 0.58077 | 0.63948 | 0.71345 | 0.75573 | 0.79925 | 0.8353 | 0.86091 | 0.89448 | 0.91138 | 0.92903 | 0.94132 | 0.95425
CCtrue | -0.0018739 | -0.0009566 | 0.00078927 | -0.00047161 | 0.00086676 | 0.0013379 | 0.00021407 | 0.001888 | 0.0027656 | 0.0036656 | 0.0050632 | 0.0063693 | 0.0067999 | 0.010765 | 0.013871 | 0.01712 | 0.023754 | 0.030523 | 0.038896 | 0.047732 | 0.06188 | 0.078774 | 0.090571 | 0.11775 | 0.13935 | 0.1692 | 0.20286 | 0.24067 | 0.29172 | 0.32687 | 0.40014 | 0.43863 | 0.51218 | 0.56571 | 0.62646 | 0.68068 | 0.73761 | 0.77832 | 0.81873 | 0.8591 | 0.88409 | 0.9127 | 0.92782 | 0.94254 | 0.95429 | 0.96199 | 0.97176 | 0.97663 | 0.98146 | 0.98479 | 0.98825
Resolution cutoff (Å)
Correlation Coefficient
Optimal resolution cutoff

## Slide 14
Chart: 
| 0.8 | 0.81 | 0.82 | 0.83 | 0.84 | 0.85 | 0.86 | 0.87 | 0.88 | 0.89 | 0.9 | 0.91 | 0.92 | 0.93 | 0.94 | 0.95 | 0.96 | 0.97 | 0.98 | 0.99 | 1 | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.0900000000000001 | 1.1000000000000001 | 1.1100000000000001 | 1.1200000000000001 | 1.1299999999999999 | 1.1399999999999999 | 1.1499999999999999 | 1.1599999999999999 | 1.17 | 1.18 | 1.19 | 1.2 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.29 | 1.3
total error | 0.011234 | 0.011798 | 0.012663 | 0.013292 | 0.013934 | 0.014784 | 0.015263 | 0.016206 | 0.01711 | 0.01817 | 0.018793
CC1/2 | 0.00027543 | 0.0009316 | 0.0011407 | 0.00093939 | -0.00085547 | 8.519e-05 | 0.0017707 | 0.0018655 | 0.00044742 | -0.00083644 | -8.9753e-05 | -0.0012856 | -0.0022441 | -0.0011232 | -0.0006217 | -0.00063863 | 0.00039935 | 0.0020736 | 0.0023749 | 0.001991 | 0.0015254 | 0.0011874 | 0.0012966 | 0.0070442 | 0.010659 | 0.014306 | 0.021539 | 0.031659 | 0.044804 | 0.054608 | 0.085054 | 0.10512 | 0.15123 | 0.19219 | 0.24658 | 0.30221 | 0.374 | 0.43569 | 0.50225 | 0.58077 | 0.63948 | 0.71345 | 0.75573 | 0.79925 | 0.8353 | 0.86091 | 0.89448 | 0.91138 | 0.92903 | 0.94132 | 0.95425
CCtrue | -0.0018739 | -0.0009566 | 0.00078927 | -0.00047161 | 0.00086676 | 0.0013379 | 0.00021407 | 0.001888 | 0.0027656 | 0.0036656 | 0.0050632 | 0.0063693 | 0.0067999 | 0.010765 | 0.013871 | 0.01712 | 0.023754 | 0.030523 | 0.038896 | 0.047732 | 0.06188 | 0.078774 | 0.090571 | 0.11775 | 0.13935 | 0.1692 | 0.20286 | 0.24067 | 0.29172 | 0.32687 | 0.40014 | 0.43863 | 0.51218 | 0.56571 | 0.62646 | 0.68068 | 0.73761 | 0.77832 | 0.81873 | 0.8591 | 0.88409 | 0.9127 | 0.92782 | 0.94254 | 0.95429 | 0.96199 | 0.97176 | 0.97663 | 0.98146 | 0.98479 | 0.98825
CC* | 0.023467 | 0.043145 | 0.047737 | 0.043325 | 0 | 0.013052 | 0.059457 | 0.061026 | 0.029907 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.028256 | 0.064333 | 0.068837 | 0.063041 | 0.055192 | 0.048703 | 0.05089 | 0.11828 | 0.14523 | 0.16795 | 0.20535 | 0.24774 | 0.29286 | 0.32181 | 0.39595 | 0.43617 | 0.51257 | 0.56782 | 0.62898 | 0.68128 | 0.73783 | 0.77906 | 0.81772 | 0.8572 | 0.88323 | 0.91256 | 0.92783 | 0.94256 | 0.95408 | 0.9619 | 0.97175 | 0.97654 | 0.98143 | 0.98477 | 0.98823
Resolution cutoff (Å)
Correlation Coefficient
Optimal resolution cutoff

## Slide 15
Chart: 
log( intensity )
Wilson Plot
2.2 1.6 1.3 1.1 1.0 0.9 0.85 0.8 0.75 0.7
resolution (Å)
8
2

## Slide 16: Optimum resolution cutoff is:
0.0 Å
after straightening Wilson plot

## Slide 17: what is “partiality”?
F(h,k,l)
Ewald sphere
λ1*
λ2*
F(0,0,0)
F(h,k,l)

## Slide 18
Real XFEL still
Lyubimov et al. (2016).
eLife 5, e18740.

## Slide 19: Best-fit Simulation of XFEL still

## Slide 20
phiz
Simulation of XFEL still
phiy
phiz
phix
phix
phiy
VS
real

## Slide 21
Chart: 
| -52.098599999999998 | -52.0319 | -51.965299999999999 | -51.898600000000002 | -52.440399999999997 | -52.373699999999999 | -52.307099999999998 | -52.240400000000001 | -52.322600000000001 | -52.255899999999997 | -52.189300000000003 | -52.122599999999998
domain 1 | 0.18748 | 0.18794 | 0.1884 | 0.188
domain 2 | 0.18826 | 0.1884 | 0.18776 | 0.18688
domain 3 | 0.18712 | 0.18775 | 0.1884 | 0.18726
Angle about y axis (°)
Fraction pixels predicted
Fitting multiple clusters is stable!
Best of
262,140 trials

## Slide 22: 3-cluster mosaic simulation

## Slide 23
Real XFEL still
Lyubimov et al. (2016).
eLife 5, e18740.

## Slide 24: 1-cluster Simulation of XFEL still

## Slide 25
The true nature of “mosaic spread”
Lovelace JJ, Murphy CR, Pahl R, Brister K & Borgstahl GEO (2006)."Tracking reflections through cryogenic cooling with topography", J. Appl. Cryst. 39, 425-432.
Any function can be approximated by a family of Gaussians

## Slide 26
3 domains within 1°

## Slide 27
3 domains within 0.2°

## Slide 28: Texture sufficient for poor stats
texture | 1 cluster | 3 clusters
Rtrue* | 4% | 16%
500 images Results from iota: cctbx & prime
*Agreement with F value fed into simulation

## Slide 29: Texture sufficient for poor stats
texture | 1 cluster | 3 clusters
Rtrue* | 4% | 16%
Rmerge | 23% | 66%
500 images Results from iota: cctbx & prime
*Agreement with F value fed into simulation

## Slide 30: Texture sufficient for poor stats
texture | 1 cluster | 3 clusters
Rtrue* | 4% | 16%
Rmerge | 23% | 66%
I/σ | 32 (1.1) | 19 (1.2)
500 images Results from iota: cctbx & prime
*Agreement with F value fed into simulation

## Slide 31: Texture sufficient for poor stats
texture | 1 cluster | 3 clusters
Rtrue* | 4% | 16%
Rmerge | 23% | 66%
I/σ | 32 (1.1) | 19 (1.2)
CC1/2 | 99 (69) | 99 (58)
500 images Results from iota: cctbx & prime
*Agreement with F value fed into simulation

## Slide 32: Texture sufficient for poor stats
texture | 1 cluster | 3 clusters
Rtrue* | 4% | 16%
Rmerge | 23% | 66%
I/σ | 32 (1.1) | 19 (1.2)
CC1/2 | 99 (69) | 99 (58)
resolution | 1.55 | 1.94
500 images Results from iota: cctbx & prime
*Agreement with F value fed into simulation

## Slide 33: Key advances to date
Understanding radiation damage
Sample preparation
Sample introduction
  robots, trays & serial
Background reduction
  Elimination of “read-out” noise
  Tighter beams: size, stability and speed
  detector calibration

## Slide 34: 140-fold multiplicity
8.2σ = Na
DELFAN residual anomalous difference

## Slide 35
Pick-up tool mark
Bragg
glitch
oxygen
inclusions
“tree rings”
1% high
average
1% low
~3x105 photon/pixel
Pilatus: subtract smooth baseline

## Slide 36
RESOLUTION COMPLETENESS R-FACTOR I/SIGMA R-meas CC(1/2) Anomal SigAno Nano
LIMIT OF DATA observed Corr
9.17 99.1% 3.9% 257.47 3.9% 100.0* 91* 5.024 450
6.49 100.0% 5.2% 214.33 5.2% 100.0* 86* 3.836 882
5.30 100.0% 7.2% 165.13 7.2% 100.0* 76* 3.257 1175
4.59 100.0% 7.2% 175.42 7.3% 100.0* 67* 2.589 1403
4.10 99.9% 7.7% 174.13 7.7% 100.0* 59* 2.264 1594
3.74 99.9% 9.4% 143.09 9.4% 100.0* 49* 1.953 1783
3.47 100.0% 11.2% 120.17 11.2% 100.0* 39* 1.696 1942
3.24 100.0% 14.1% 91.14 14.1% 100.0* 30* 1.333 2103
3.06 99.9% 19.5% 65.79 19.5% 100.0* 23* 1.117 2214
2.90 99.9% 29.0% 44.85 29.1% 99.9* 17* 1.008 2369
2.77 99.9% 40.5% 32.58 40.6% 99.8* 11* 0.901 2493
2.65 99.9% 52.8% 25.16 52.9% 99.8* 10* 0.866 2605
2.54 100.0% 67.4% 19.47 67.6% 99.6* 2 0.804 2705
2.45 100.0% 88.9% 14.58 89.2% 99.2* 4 0.831 2859
2.37 100.0% 109.3% 9.97 109.7% 98.1* 5 0.829 2925
2.29 100.0% 138.2% 6.87 138.9% 96.1* 1 0.760 3037
2.22 100.0% 197.1% 4.03 198.6% 83.5* -1 0.721 3159
2.16 100.0% 227.3% 2.41 230.8% 46.9* -1 0.677 3224
2.10 61.2% 154.4% 1.28 163.6% 47.0* -2 0.660 1999
2.05 47.9% 170.1% 0.68 196.5% 25.7* 3 0.629 1578
total 93.3% 15.7% 54.30 15.8% 100.0* 12* 1.217 42499
140-fold multiplicity: 16 crystals, 360° each, inverse beam, 7235 eV

## Slide 37: Key advances to date
Understanding radiation damage
Sample preparation
Sample introduction
  robots, trays & serial
Background reduction
  Elimination of “read-out” noise
  Tighter beams: size, stability and speed
  detector calibration
Population of “fold space”

## Slide 38: MR simulation
Chart: 
Rmsd from perfect search model (Å)
corrupted model
Correlation coefficient to correct density

## Slide 39: Key advances to come
Firehose of testable hypotheses
Physiological temperature
  Time-resolved, complex sample environment
  Plates, jets, magnets, and lasers!
Serial comes of age
  stills should be better
  End of the era of beams missing crystals
Painting with X-rays

## Slide 40: “rastering” to find the crystal
Membrane
protein crystal
X-ray beam
ice

## Slide 41: “painting” for flat dose

## Slide 42: “painting” for virtual beam

## Slide 43: Painting with X-rays
0°
img_000045

## Slide 44: Painting with X-rays
0.2°
img_000100

## Slide 45: Painting with X-rays
30°
img_015000

## Slide 46: Painting with X-rays
90°
img_045000

## Slide 47: Painting with X-rays
360°
img_180000

## Slide 48: Painting with X-rays
1080°
img_540000

## Slide 49: Beam Shape Dichotomy
Flat-top beam
  Uniform dose in illuminated region
  Sharp edges: sensitive to drift
Gaussian beam
  Softer edges
  High dose contrast
Best of both?
  “dose flattening”

## Slide 50: How many images?
360° / 0.1° * ( 100 µm / 1 µm )2
= 36x106
28,000 photons/image

## Slide 51: Key advances to come
Firehose of testable hypotheses
Physiological temperature
  Time-resolved, complex sample environment
  Plates, jets, magnets, and lasers!
Serial comes of age
  stills should be better
  End of the era of beams missing crystals
Painting with X-rays
Weak-image limit

## Slide 52: Weak-image limit: is that a spot?

## Slide 53
Weak-image limit: is that a spot?

## Slide 54: Weak-image limit: is that a spot?

## Slide 55: Weak-image limit: is that a spot?

## Slide 56: Weak-image limit: is that a spot?

## Slide 57: Weak-image limit: is that a spot?

## Slide 58
photons
0
4
time
photons
0
4
photons
0
4
time
photons
0
4
Single Photons Carry Information
About Radiation Damage

## Slide 59: We should stop using lossy compression!
long exposures
  >1 photon/image = order of arrivals
Wide phi slice
  sum over angle
Large beams
  sum over profile area
Pink/divergent beam
  sum over lambda/angle
Background
  non-crystal stuff
Systematics
  Dose contrast, drift, read-out gap

## Slide 60: Lossless data rate
The 1.2e-5 scattered photons per incident photon can be derived form the cross section of oxygen (0.2 cm^2/g) and the rough density of protein crystals (1.2 g/cm^3), and the fact that about half of all scattered photons hit a flat detector: 0.2*1.2*1e-4/2 = 1.155e-5 That is, flux*thickness*1.2e-5 is the rate of photon arrivals and log2(pixels) is the number of bits required to store the location of one photon in the entropy-limited regime. Lossless compression cannot be better than this because individual photon arrivals are genuinely random.
bits
s
= log2(pixels) * flux * thick * 1.2x10-5
where: bits/s = entropy-limited compressed data rate
pixels = number of pixels in the detector
flux = photons/s onto the sample thick = sample thickness in microns
EG: 1e12 ph/s * 100um thick = 3.4 GBytes/s

## Slide 61: To discuss later…
Resolution cut-off?
Blue sausages?
Detector calibration?
Weak image limit?

## Slide 62: read noise = 2 photons

## Slide 63: read noise = 0 photons

## Slide 64: read noise = 0.07 photons

## Slide 65: read noise = 0.5 photons

## Slide 66: read noise = 2 photons

## Slide 67: read noise = 7.5 photons

## Slide 68: read noise = 47 photons

## Slide 69: read noise = 190 photons

## Slide 70: read noise = 750 photons

## Slide 71: read noise = 4700 photons

## Slide 72: read noise = 19000 photons

## Slide 73: read noise = 75000 photons

## Slide 74: read noise = 470000 photons

## Slide 75: read-out noise
Chart: 
| 9.9999999999999995E-21 | 3.3306200000000001E-2 | 3.3306200000000001 | 13.3225 | 29.9756 | 83.265600000000006 | 163.20099999999999 | 333.06200000000001 | 333.06200000000001 | 1332.25 | 1332.25 | 8326.56 | 18734.8 | 33306.199999999997 | 3330625 | 9.9999999999999995E-21 | 1.8765199999999999E-2 | 7.5061000000000003E-2 | 0.46913100000000002 | 1.87652 | 7.5061 | 46.9131 | 187.65199999999999 | 750.61 | 4691.3100000000004 | 18765.2 | 75061 | 469131
myoglobin (Fe) | 1 | 0.99913 | 1.0028 | 1.0031 | 0.99965 | 0.99844 | 0.98597 | 0.96258 | 0.96327 | 0.89241 | 0.89865 | 0.70825 | 0.56133 | 0.48597 | 0.16978
lysozyme (Gd) | 1 | 0.99555 | 1.0058 | 0.99738 | 0.98402 | 0.96673 | 0.9531 | 0.90359 | 0.71994 | 0.52947 | 0.34268 | 0.22243 | 0.11056
read-out noise (“extra” photons)
normalized peak height
Q315r
Q315
Mar
325HE
0.1 1 10 100 103 104 105 106
Pilatus

## Slide 76: Beam Flicker
Chart: 
time (seconds)
normalized flux
pinhole
removed

## Slide 77: Beam Flicker
Chart: 
| 9.9999999999999995E-21 | 1E-3 | 2E-3 | 4.0000000000000001E-3 | 8.0000000000000002E-3 | 1.6E-2 | 3.2000000000000001E-2 | 6.4000000000000001E-2 | 0.128 | 0.25600000000000001 | 0.51200000000000001 | 0.6 | 0.7 | 0.8 | 0.9 | 1 | 1.1000000000000001 | 1.2 | 1.3 | 1.4 | 1.5 | 2.0499999999999998 | 4.0999999999999996 | 8.19 | 16.399999999999999 | 0.15 | 0.15
CC | 0.5406 | 0.5502 | 0.5774 | 0.5402 | 0.5516 | 0.5745 | 0.5264 | 0.5498 | 0.5471 | 0.5191 | 0.4908 | 0.4819 | 0.5012 | 0.459 | 0.423 | 0.4922 | 0.4057 | 0.4178 | 0.4259 | 0.3138 | 0.3285 | 0.4044 | 0.2088 | 0.1058 | 0.0254
 | 0 | 1
flicker noise (%/√Hz)
CC to correct model

## Slide 78: Shutter Jitter
Chart: 
| 0.01 | 0.02 | 0.04 | 0.08 | 0.16 | 0.32 | 0.64 | 1.28 | 2.56 | 5.12 | 10.24 | 20.48 | 40.96 | 5.7000000000000002E-2 | 5.7000000000000002E-2
CC | 0.7083 | 0.7178 | 0.709 | 0.7164 | 0.7061 | 0.7004 | 0.705 | 0.6983 | 0.6576 | 0.5574 | 0.3127 | 0.0879 | 0.0216
 | 0 | 1
rms timing error (% exposure)
CC to correct model

## Slide 79: Shutter Jitter
open
closed

## Slide 80: Shutter Jitter
open
closed

## Slide 81: Shutter Jitter
open
closed

## Slide 82: Shutter Jitter
open
closed

## Slide 83: Shutter Jitter
open
closed

## Slide 84: Shutter Jitter
open
closed

## Slide 85: Shutter Jitter
open
closed

## Slide 86: Shutter Jitter
open
closed

## Slide 87: Shutter Jitter
open
closed

## Slide 88: Shutter Jitter
open
closed

## Slide 89: Shutter Jitter
open
closed
shutter jitter

## Slide 90: “Shutterless” full vs partial
time
intensity
1 ms

## Slide 91: Rough values of energy quanta
1 MeV 100 GJ/mol Medical radiation therapy
100 keV 10 GJ/mol Medical imaging
10 keV 1 GJ/mol X-ray crystallography
1 keV 100 MJ/mol S and P K-edges
100 eV 10 MJ/mol “water window”
10 eV 1 MJ/mol C≡C bond
1 eV 100 kJ/mol C-C bond, visible light
100 meV 10 kJ/mol hydrogen bond
10 meV 1 kJ/mol heat (~300 K)

## Slide 92: signal vs noise
“If you don’t have good data, then you have no data at all.”
-Sung-Hou Kim

## Slide 93: signal vs noise
easy
hard
impossible

## Slide 94: signal vs noise
easy
hard
impossible
threshold of “solvability”

## Slide 95: signal vs noise
“If you don’t have good data, then you must learn statistics.”
-James Holton

## Slide 96: Adding noise

## Slide 97: Adding noise
12 + 12 = 1.42

## Slide 98: Adding noise
12 + 12 = 1.42
32 + 12 = 3.22
σtotal2 = σ12 + σ22

## Slide 99: Adding noise
12 + 12 = 1.42
32 + 12 = 3.22
σtotal2 = σ12 + σ22

## Slide 100: Adding noise
12 + 12 = 1.42
32 + 12 = 3.22
σtotal2 = σ12 + σ22

## Slide 101: Adding noise
12 + 12 = 1.42
32 + 12 = 3.22
102 + 12 = 10.052

## Slide 102
The number of photons scattered
before crystal is dead
is independent
of flux & time
1 um3 = 105 photons (roomT)
1 um3 = 106 photons (synch)
1 um3 = 108 photons (XFEL)
Henderson, 1990; Gonzalez & Nave, 1994; Glaeser et al., 2000; Sliz et al., 2003; Leiros et al., 2006; Owen et al., 2006; Garman & McSweeney, 2006; Garman & Nave, 2009; Holton, 2009
(synch)

## Slide 103: Dose slicing
crystal’s useful life
N
photons
N
photons
N
photons
low
completeness
Can’t process?

## Slide 104
Same number of photons | 
fine | coarse
3.177 | 3.140
Damage correctable | Damage uncorrectable
resolution vs dose slice