## Slide 1: Acknowledgements
UCSF LBNL SLAC
ALS 8.3.1: TomAlberTron
NIH NIAID P50 AI150476
NIH NIGMS R01 GM124149
NIH NIGMS P30 GM124169
Integrated Diffraction Analysis Technologies (IDAT)
Plexxikon, Inc.
Synchrotron Radiation Structural Biology Resource (SLAC)
The Advanced Light Source is supported by the Director, Office of Science, Office of Basic Energy Sciences, Materials Sciences Division, of the US Department of Energy under contract No. DE-AC02-05CH11231 at Lawrence Berkeley National Laboratory.
Robert Stroud
James Fraser
Nevan Krogan
Paul Adams
Greg Hura
Aina Cohen
Mike Soltis
HARC Core 6: diffraction
Advanced Light Source

## Slide 2: Advanced Light Source
5.0.2
5.0.1
5.0.3
MBC 4.2.2
8.2.2
8.3.1
8.2.1
12.3.1 SIBYLS
MX & SAXS
Berkeley
Center for
Structural
Biology
ALS
Structural Biology
2021
UC
2.1
NCXT
2.0.1 Gemini
SF

## Slide 3: ALS 8.3.1 optics
ADSC Quantum 315r
Superbend
Plane
Parabolic
mirror
Torroidal
mirror
Si(111)
monochromator
Protein Crystal
pinhole
Scatter
guard
divergence
slits
MacDowell et. al. (2004) J. Sync. Rad. 11 447-455
2:1 demagnification cancels spherical aberrations
Pt/Rh mirrors 1.3 x 108 ph/s/μm2 ≈ 300 s/xtal
total power dissipation: 100 W
Notes: Most notable acheivements are high brightness with very little “wasted” light. The discovery made by Howard Padmore? At ALS showed that a 2:1 demagnification cancels nearly all of the spherical abberations in the parabolic/torroid system. 30-50% of the total x-ray intensity passes through the 100um pinhole. The flux through the 100um pinhole at the superbend beamlines is comparable to that of the 5.0.2 wiggler or an APS bending magnet, but only 150 Watts of power must be dissipated by the beamline optics. In contrast, wiggler beamlines must disipate kilowatts. The optics here are much simpler and not exotic in any way. The system is stable and robust. There is never a need to re-focus the beam at a new wavelength. Only a 2-minute automatic tuneup is recommended for very large energy changes (more than 1-2 keV).

## Slide 4: ALS robots
5.0.2
5.0.1
5.0.3
MBC 4.2.2
ACTOR
8.2.2
8.3.1
Cool Hand Luke
8.2.1
12.3.1 SIBYLS
ACTOR
Berkeley
Auto-
Mounter
ALS MX robots

## Slide 5: What is ALS Beamline 8.3.1 ?
Macromolecular Crystallography (MX)
  MAD, 1x1012 ph/s, Pilatus 6M 50 Hz
  Technology development
    Room temp, Elves, multi-crystal
  Staff: Holton, Meigs, Burnett, Classen
Landmark and highly productive facility
  1st SuperBend – $16M invested since 2001
  1826 structures, 720 pubs, 250 high-impact
  8% of all membrane protein structures
  1000+ distinct users from 60+ institutions
  76% utilization
George
Scott
Kathryn
Notes: ALS 8.3.1 is a very productive facility. Our focus is on maximizing efficiency of the user experience.
The MRPI funds are not just to maintain this facility, but to upgrade it. The new detector is 50x faster, more sensitive and more accurate than the existing detector, and an optics upgrade will match these capabilities.
We will exceed the modern standard of screening and data collection by offering completely automated mail-in raster-based screening
This will position ALS 8.3.1 as the centerpiece of structural biology for the UC system

## Slide 6: TomAlberTron’s User Community
1000+ users, 60+ institutions
PRT investment: $16 M since 2001
Return: 1826 PDB 720 pubs 250 hi-impact
$132+ M (NIH) $935+ M (Plexxikon)
Beam time utilization : 76%
Median visit-to-next-visit interval:
  11 ± 8 days

## Slide 7
Pilatus3 M 6M
Pixel Array
Detector
6x106 pixels
$0.15/each
1 ms readout
50 Hz images

## Slide 8: ALS 8.3.1 data collection history
images x 106

## Slide 9: Shutterless rastering now available
Membrane
protein crystal
X-ray beam
ice

## Slide 10
BLU-ICE w unattended data collection

## Slide 11: Pin compatibility
SLAC
ALS
SPINE

## Slide 12: SuperTong

## Slide 13: Sample pin compatibility
Caltech
SSRL
Yale
Oxford
SPring-8
SPINE
MiTeGen RT
Syrrx
H. magnetic
H. hybrid
H. copper
H. screw cap
H. colored
Hampton screw cap
Hampton magnetic

## Slide 14: “infinite capacity” sample carousel

## Slide 15: Carousel open

## Slide 18
pinhole
mirror
microscope
SBS
format
crystallization
tray
Notes: Inspired by the ESRF microcrystal camera, the superbend endstations are all equipped with zero-paralax optics. The user sees their sample from the point of view of the x-ray beam. Sample alignment is not only more stable but more intuitive than an off-axis camera. Placing the crystal into the microscope crosshairs gaurantees that it will be hit by the beam.

## Slide 19
NatX-ray
CrystalQuick™ X
250 μm
~50 μm tape
Fluidigm
Topaz® DC
200 μm
PDMS
PDMS
MiTeGen
in-situ-1
100 μm
~50 μm tape
in-situ diffraction trays
CrystalDirect
~50 μm tape
25 μm
Microlytic
Crystal Former™
1000 μm
Notes: Next slide please

## Slide 20: scattering/absorptoin “rules”:
1 μm crystal ≈ 1 μm water
≈ 1 μm plastic
≈ 0.1 μm glass
≈ 1000 μm air
Optimum: support thickness = crystal thickness

## Slide 21: Crystal lifetime vs pinhole
pinhole (µm) | beam size | Cryo (min) | Room T (s)
100 | 60x80 | 4 | 1.7
75 | 60x78 | 5 | 2.0
50 | 60x54 | 8 | 3.4
30 | 54x34 | 6 | 2.3
20 | 44x24 | 6 | 2.4
15 | 40x19 | 7 | 2.8
10 | 34x14 | 10 | 4.1

## Slide 22
Holton (2009) J. Synchrotron Rad. 16 133-42
Radiation Damage World Records
MGy reaction reference
~45 global damage Owen et al. (2006)
10/Å global damage Howells et al. (2009)
5 Se-Met Holton (2007)
4 Hg-S Ramagopal et al. (2004)
4 R-C-COOH Garman et al. (2015)
3 S-S Murray et al. (2002)
1 Br-RNA Olieric et al. (2007)
~1? Cl-C ???
0.5 Mn in PS II Yano et al. (2005)
0.06 putidaredoxin Corbett et al. (2007)
0.02 Fe in myoglobin Denisov et al. (2007)

## Slide 23
Holton & Frankel (2010) Acta D 66 393-408.
Will my experiment work?

## Slide 24: S-SAD merging 20 xtals
Knott et al. (2017). "Guide-bound CRISPR-Cas13a." Nature S&MB online

## Slide 25: ALS 8.3.1 capabilities
200s data sets in shutterless mode
Support for room temperature
in-situ screening from trays
15-100 μm beam sizes
True BLU-ICE interface
Excellent detector calibration
Infinite capacity sample carousel
Automated data backup
Remote operation

## Slide 26: Recent Highlights
Chart: 
6
8.3.1
6
cryoEM
2
SLS
Doudna
CRISPR PDBs
COVID protection:
nanobody aerosol
Schoof et al., Science 370 (2020)
UCSF-Oxford SARS-CoV-2
fragment screen
Schuller et al., Science Advances in press (2021)

## Slide 27
The Future of Light
APS-U shutdown 2022
APS-U shutdown 2023

## Slide 28: Summary
Leading in MX methods
Universal pin compatibility
In-situ screening of trays
Unattended data collection
Future: less light → more innovation
Questions? JMHolton@lbl.gov

## Slide 29: 8.3.1 Recent Improvements
Dectris Pilatus3 6M installation - 2018
  50 Hz framing, 1 ms dead time, previously 25 Hz
  0.8% calibration error, previously 2.5%
Complete control system update
  Shutterless data collection
  “Rastering” automated sample alignment
  Screening Tab workflow for efficient sample evaluation
  Raster centering for unattended operation
Overhaul of data archiving infrastructure
  LTO4 to LTO6, support USB3 user hard drives
First remote data collection – Apr 2017