## Slide 1: PowerPoint File available:
http://bl831.als.lbl.gov/
~jamesh/powerpoint/
IGBMC_SvN_2016.pptx

## Slide 2: Acknowledgements
UCSF LBNL SLAC
ALS 8.3.1 creator: Tom Alber
UC Multicampus Research Programs and Initiatives (MRPI)
UCSF Program for Breakthrough Biomedical Research (PBBR)
one-time NIH-DOE Inter-agency agreement (IAA)
Integrated Diffraction Analysis Technologies (IDAT)
Plexxikon, Inc.
M D Anderson CRC
Synchrotron Radiation Structural Biology Resource (SLAC)
The Advanced Light Source is supported by the Director, Office of Science, Office of Basic Energy Sciences, Materials Sciences Division, of the US Department of Energy under contract No. DE-AC02-05CH11231 at Lawrence Berkeley National Laboratory.
Robert Stroud James Fraser John Spence
Chris Nielsen Clemens Schulze-Briese Aina Cohen Ana Gonzalez

## Slide 3: Decisions, Decisions, Decisions
Exposure time
Number of images
Wavelengths
Beam: size, flux, divergence, bandpass
Inverse beam?
Kappa? - overlaps
Multiple crystals? - non-isomorphism
Run Strategy

## Slide 4: The “success rate” of structure determination
100 s/dataset
200 days/year
~150 beamlines
~26,000,000 datasets/year
9333 PDBs in 2015

## Slide 5: signal vs noise
“If you don’t have good data, then you have no data at all.”
-Sung-Hou Kim

## Slide 6: signal vs noise
easy
hard
impossible

## Slide 7: signal vs noise
easy
hard
impossible
threshold of “solvability”

## Slide 8: signal / noise = ?
CCD
Read-out
Photon
counting
Beam flicker
Shutter jitter
Sample vibration
Pile-up
Detector calibration
attenuation
partiality
Non-isomorphism
Radiation damage
Time
none sqrt proportional
none
1/sqrt
1/prop.

## Slide 9: signal vs noise
“If you don’t have good data, then you must learn statistics.”
-James Holton

## Slide 10: Adding noise

## Slide 11: Adding noise
12 + 12 = 1.42

## Slide 12: Adding noise
12 + 12 = 1.42
32 + 12 = 3.22
σtotal2 = σ12 + σ22

## Slide 13: Adding noise
12 + 12 = 1.42
32 + 12 = 3.22
σtotal2 = σ12 + σ22

## Slide 14: Adding noise
12 + 12 = 1.42
32 + 12 = 3.22
σtotal2 = σ12 + σ22

## Slide 15: Adding noise
12 + 12 = 1.42
32 + 12 = 3.22
102 + 12 = 10.052

## Slide 16: 100% SeMet incorporation
Trivial to solve

## Slide 17: 100% S -Met incorporation
Impossible to solve

## Slide 18: Phase Problem: Se vs S
1.0 - fraction Se
correlation coefficient

## Slide 19: Phase Problem: Se vs S
correlation coefficient
1.0 - fraction Se

## Slide 20: Phase Problem: Se vs S
correlation coefficient
1.0 - fraction Se

## Slide 21: Phase Problem: Se vs S
correlation coefficient
1.0 - fraction Se

## Slide 22: Phase Problem: Se vs S
correlation coefficient
1.0 - fraction Se

## Slide 23
29% Se 71% S

## Slide 24
28% Se 72% S

## Slide 25
21% Se 79% S

## Slide 26
11% Se 89% S

## Slide 27: MR simulation
Rmsd from perfect search model (Å)
corrupted model
Correlation coefficient to correct density

## Slide 28: The transitions are sharp!
How can we predict success/failure?
Know Thy Experiment

## Slide 29: Basic Principles
“Hell, there are NO RULES here - we're trying to accomplish something.”
Thomas A. Edison – inventor
“You’ve got to have an ASSAY.”
Arthur Kornberg – Nobel Laureate
“Control, control, you must learn CONTROL!”
Yoda – Jedi Master

## Slide 30: Where do photons go?
beamstop
Transmitted (98%)
Protein
1A x-rays

## Slide 31: Where do photons go?
beamstop
Transmitted (98%)
Protein
1A x-rays
2% total → error < 2%

## Slide 32: attenuation factor
Bouguer, P. (1729). Essai d'optique sur la gradation de la lumière.
Lambert, J. H. (1760). Photometria: sive De mensura et gradibus luminis, colorum et umbrae. E. Klett.
Beer, A. (1852)."Bestimmung der Absorption des rothen Lichts in farbigen Flüssigkeiten", Ann. Phys. Chem 86, 78-90.
A = = exp(-μt)
IT
Ibeam
t
μ

## Slide 33: attenuation factor
Bouguer, P. (1729). Essai d'optique sur la gradation de la lumière.
Lambert, J. H. (1760). Photometria: sive De mensura et gradibus luminis, colorum et umbrae. E. Klett.
Beer, A. (1852)."Bestimmung der Absorption des rothen Lichts in farbigen Flüssigkeiten", Ann. Phys. Chem 86, 78-90.
A = = exp(-μ(tin+ tout))
IT
Ibeam
t
tin
tout
A = = exp(-μt)
IT
Ibeam
μ

## Slide 34: attenuation factor
Bouguer, P. (1729). Essai d'optique sur la gradation de la lumière.
Lambert, J. H. (1760). Photometria: sive De mensura et gradibus luminis, colorum et umbrae. E. Klett.
Beer, A. (1852)."Bestimmung der Absorption des rothen Lichts in farbigen Flüssigkeiten", Ann. Phys. Chem 86, 78-90.
A = = exp(-μ(tin+ tout))
IT
Ibeam
tin
tout
tin
tout
tin
tout
μ

## Slide 35: attenuation factor
Bouguer, P. (1729). Essai d'optique sur la gradation de la lumière.
Lambert, J. H. (1760). Photometria: sive De mensura et gradibus luminis, colorum et umbrae. E. Klett.
Beer, A. (1852)."Bestimmung der Absorption des rothen Lichts in farbigen Flüssigkeiten", Ann. Phys. Chem 86, 78-90.
A = = exp[-μxtal(txi+ txo)
-μsolvent(tsi + tso)]
IT
Ibeam
μxtal
txi
txo
tsi
tso
txi
txo
tsi
tso
txi
txo
tsi
tso
μsolvent

## Slide 36: Decisions, Decisions, Decisions
Exposure time
Number of images
Wavelengths
Beam: size, flux, divergence, bandpass
Inverse beam?
Kappa? - overlaps
Multiple crystals? - non-isomorphism
Run Strategy

## Slide 37: Elastic scattering

## Slide 38: Elastic scattering

## Slide 39: Where do photons go?
beamstop
elastic scattering (6%)
Transmitted (98%)
Protein
1A x-rays

## Slide 40: Elastic scattering

## Slide 41: Inelastic scattering

## Slide 42: Where do photons go?
beamstop
elastic scattering (6%)
Transmitted (98%)
Protein
1A x-rays

## Slide 43: Where do photons go?
beamstop
elastic scattering (6%)
Transmitted (98%)
inelastic scattering (7%)
Protein
1A x-rays

## Slide 44: Where do photons go?
beamstop
elastic scattering (6%)
Transmitted (98%)
inelastic scattering (7%)
Protein
1A x-rays
Re-emitted (99%)
Absorbed (~0%)

## Slide 45: Photoelectric absorption

## Slide 46: Photoelectric absorption
e-
+

## Slide 47: Fluorescence
+

## Slide 48: Fluorescence
e-
+

## Slide 49: Metal identification

## Slide 50: Auger emission
+

## Slide 51: Auger emission
++
e-

## Slide 52: Where do photons go?
beamstop
elastic scattering (6%)
Transmitted (98%)
inelastic scattering (7%)
Protein
1A x-rays
Re-emitted (99%)
Absorbed (~0%)

## Slide 53: Where do photons go?
beamstop
elastic scattering (6%)
Transmitted (98%)
inelastic scattering (7%)
Photoelectric (87%)
Protein
1A x-rays
Re-emitted (99%)
Absorbed (~0%)

## Slide 54: Where do photons go?
beamstop
elastic scattering (6%)
Transmitted (98%)
inelastic scattering (7%)
Photoelectric (87%)
Protein
1A x-rays
Re-emitted (~0%)
Absorbed (99%)
Re-emitted (99%)
Absorbed (~0%)

## Slide 55: Where do photons go?
beamstop
elastic scattering (6%)
Transmitted (98%)
useful/absorbed energy: 7.3%
inelastic scattering (7%)
Photoelectric (87%)
Protein
1A x-rays
Re-emitted (~0%)
Absorbed (99%)
Re-emitted (99%)
Absorbed (~0%)

## Slide 56: Photoelectric absorption

## Slide 57: Photoelectric absorption
e-
+

## Slide 58: Secondary ionization
e-

## Slide 59: Secondary ionization
e-
e-
+

## Slide 60: Excitation
e-

## Slide 61: Excitation
e-

## Slide 62: Excitation
e-

## Slide 63: Excitation

## Slide 64: Electron attachment
e-

## Slide 65: Electron attachment
-

## Slide 66: Sample atoms

## Slide 67: Sample atoms

## Slide 68: Primary ionization
e-
+

## Slide 69: Secondary ionizations
e-
e-
e-
e-
e-
e-
+
+
+
+
+
+

## Slide 70: Ionization track
e-
e-
e-
e-
e-
e-
e-
e-
e-
+
+
+
+
+
+
+
+
+
e-
+

## Slide 71: Ionization track
e-
e-
e-
e-
e-
e-
e-
e-
e-
+
+
+
+
+
+
+
+
+
e-
+

## Slide 72: Ionization track

## Slide 73: Ionization track

## Slide 74: Ionization track
3 μm

## Slide 75: Ionization track
Linear Energy Transfer

## Slide 76: initial effects
e-
e-
e-
e-
e-
e-
e-
e-
e-
+
+
+
+
+
+
+
+
+
e-
+

## Slide 77: initial effects
e-
e-
e-
e-
e-
e-
e-
e-
e-
+
+
+
+
+
+
+
+
+
neutralization
e-
+

## Slide 78: initial effects
e-
e-
e-
e-
e-
e-
e-
e-
+
+
+
+
+
+
+
+
neutralization
e-
+

## Slide 79: initial effects
e-
e-
e-
e-
e-
e-
e-
+
+
+
+
+
+
neutralization
neutralization
+

## Slide 80: initial effects
e-
e-
e-
e-
e-
e-
e-
+
+
+
+
+
+
attachment
+

## Slide 81: initial effects
e-
e-
e-
e-
-
e-
e-
+
+
+
+
+
+
attachment
+

## Slide 82: initial effects
e-
e-
e-
e-
-
e-
e-
+
+
+
+
+
+
attachment
attachment
+

## Slide 83: initial effects
e-
e-
-
e-
e-
+
+
+
+
+
+
-
-
+

## Slide 84: initial effects
-
e-
+
+
+
+
+
-
-
-
-
+

## Slide 85: initial effects
-
e-
+
+
+
+
+
-
-
-
-
+
solvation

## Slide 86: initial effects
-
e-
+
+
+
+
+
-
-
-
-
solvation
+

## Slide 87: Solvated electron
e-

## Slide 88: Solvated electron
e-

## Slide 91: initial effects
-
e-
+
+
+
+
+
-
-
-
-
+

## Slide 92: initial effects
-
e-
+
+
+
+
+
-
-
-
-
+

## Slide 93: initial effects
-
e-
+
+
+
+
-
-
-
+

## Slide 94: initial effects
-
e-
+
+
-
+

## Slide 95: initial effects
-
e-
+
+

## Slide 96: Excitation transfer

## Slide 97: Excitation transfer
UV photon

## Slide 98: Excitation transfer

## Slide 99: Excitation transfer
Förster transfer
r-6

## Slide 100: Excitation transfer

## Slide 101: t = 0 s

## Slide 102: t ≈ 10-14 s
-
e-
+
+

## Slide 103: Timescales of radiation damage
Garret et. al. (2005) Chem. Rev. 105, 355-389

## Slide 104: Timescales of radiation damage
Garret et. al. (2005) Chem. Rev. 105, 355-389
LCLS
ALS
bunch

## Slide 105: Ionization track
Heterogeneous reactions
Homogeneous reactions

## Slide 106: Rough values of energy quanta
1 MeV 100 GJ/mol Medical radiation therapy
100 keV 10 GJ/mol Medical imaging
10 keV 1 GJ/mol X-ray crystallography
1 keV 100 MJ/mol S and P K-edges
100 eV 10 MJ/mol “water window”
10 eV 1 MJ/mol C≡C bond
1 eV 100 kJ/mol C-C bond, visible light
100 meV 10 kJ/mol hydrogen bond
10 meV 1 kJ/mol heat (~300 K)

## Slide 107: Energy Transfer Analogy
37 C

## Slide 108
Damage is done
by dose (MGy)
proportional to photons/area
not time
not heat
Sliz P, Harrison SC & Rosenbaum G (2003). Structure 11, 13-19.
Garman EF & McSweeney SM (2006). J. Sync. Rad. 14, 1-3.
Owen RL, Rudino-Pinera E & Garman EF (2006). PNAS 103, 4912-4917.
Leiros et al. (2006). Acta Cryst. D 62, 125-132.
Holton JM (2007). J. Synch Rad. 14, 51-72.

## Slide 109: what the is a MGy?
http://bl831.als.lbl.gov/
damage_rates.pdf
Holton J. M. (2009) J. Synchrotron Rad. 16 133-42
1 MGy = 106 J/kg

## Slide 110
How long will my crystal last?
synch line type flux beamsize flux density dose max xtal min site
ph/s μm ph/μm2/s rate lifetime lifetime
ALS 4.2.2 MAD 2.2e11 75x80 3.7e+07 27.3 kGy/s 18 m 73 s
ALS 5.0.1 mono 1.6e11 100 2.0e+07 10.4 kGy/s 48 m 3.2 m
ALS 5.0.2 MAD 8e11 100 1.0e+08 51.8 kGy/s 9.6 m 39 s
ALS 5.0.3 mono 1.7e11 100 2.2e+07 11 kGy/s 45 m 3 m
ALS 8.2.1 MAD 1.8e11 100 2.3e+07 11.7 kGy/s 43 m 2.9 m
ALS 8.2.2 MAD 2.3e11 100 2.9e+07 14.9 kGy/s 34 m 2.2 m
ALS 8.3.1 MAD 9e11 70 2.3e+08 119 kGy/s 4.2 m 17 s
ALS 8.3.1 typical 6e11 70 1.6e+08 115 kGy/s 4.3 m 17 s
ALS 12.3.1 MAD 1.8e11 100 2.3e+07 11.7 kGy/s 43 m 2.9 m
ALS 12.3.1 ML 4.0e12 100 5.1e+08 513 kGy/s 58 s 3.9 s
APS 8-BM MAD 1e11 200 2.5e+06 1.27 kGy/s 6.6 h 26 m
APS 14-BM-C mono 5.8e10 200 1.4e+06 738 Gy/s 11 h 45 m
APS 14-BM-D MAD 3.3e9 200 8.2e+04 42 Gy/s 8.3 d 13 h
APS 14-ID-B MAD 6.0e10 200 1.5e+06 763 Gy/s 11 h 44 m
APS 17-BM MAD 1.1e11 200 2.8e+06 1.4 kGy/s 6 h 24 m
APS 17-ID MAD 2.3e11 200 5.8e+06 2.93 kGy/s 2.8 h 11 m
APS 19-BM MAD 2.0e11 70x60 4.8e+07 24.2 kGy/s 21 m 83 s
APS 19-ID MAD 1.3e13 80x40 4.1e+09 2.07 MGy/s 15 s 0.97 s
APS 19-ID typical 5.5e11 100x100 5.5e+07 28 kGy/s 18 m 71 s
APS 22-BM MAD 7e12 80x40 2.2e+09 1.23 MGy/s 24 s 1.6 s
APS 22-ID MAD 7e12 80x40 2.2e+09 1.23 MGy/s 24 s 1.6 s
APS 22-ID typical 1.5e12 80 2.3e+08 119 kGy/s 4.2 m 17 s
APS 23-ID-B MAD 1e13 75x25 5.3e+09 3.01 MGy/s 10 s 0.66 s
APS 23-ID typical 1.5e12 80 2.3e+08 119 kGy/s 4.2 m 17 s
APS 24-ID-C MAD 1.3e13 20x60 1.1e+10 5.23 MGy/s 5.7 s 0.38 s
APS 24-ID-E MAD 0.5e13 20x100 2.5e+09 1.19 MGy/s 25 s 1.7 s
APS 31-ID MAD 2e12 70 4.1e+08 194 kGy/s 2.6 m 10 s

## Slide 111: resolution dependence of global damage
resolution (Å)
maximum tolerable dose (MGy)
1 2 3 5 7 10 20 40 70 100
1
10
100
103
Howells et al. (2009) J. Electron. Spectrosc. Relat. Phenom. 170 4-12

## Slide 112: resolution dependence of global damage
resolution (Å)
maximum tolerable dose (MGy)
1 2 3 5 7 10 20 40 70 100
1
10
100
103
Howells et al. (2009) J. Electron. Spectrosc. Relat. Phenom. 170 4-12
10 MGy/Å

## Slide 113: Riso vs dose
Riso (%)
change in dose (MGy)
data taken from Banumathi, et al. (2004) Acta Cryst. D 60, 1085-1093.

## Slide 114: Riso vs dose
Riso (%)
change in dose (MGy)
data taken from Banumathi, et al. (2004) Acta Cryst. D 60, 1085-1093.
Riso ≈ 0.7 %/MGy

## Slide 115: Dose-doubling concentration
Na | 44
Mg | 27
P | 9
S | 7
Cl | 5
K | 3.3
Ca | 2.7
Fe | 1.0
Cu | 0.67
Zn | 0.58
molar, at the Se edge based on RADDOSE
As | 0.44
Se | 0.42
Br | 2.9
I | 0.47
Gd | 0.21
Ta | 0.13
Pt | 0.17
Au | 0.16
Hg | 0.16
U | 0.23

## Slide 116: Dose-doubling concentration
Na | 44
Mg | 27
P | 9
S | 7
Cl | 5
K | 3.3
Ca | 2.7
Fe | 1.0
Cu | 0.67
Zn | 0.58
molar, at the Se edge based on RADDOSE
As | 0.44
Se | 0.42
Br | 2.9
I | 0.47
Gd | 0.21
Ta | 0.13
Pt | 0.17
Au | 0.16
Hg | 0.16
U | 0.23

## Slide 117
Holton (2009) J. Synchrotron Rad. 16 133-42
Specific Damage
World Records
MGy reaction reference
~45 global damage Owen et al. (2006)
5 Se-Met Holton (2007)
4 Hg-S Ramagopal et al. (2004)
3 S-S Murray et al. (2002)
1 Br-RNA Olieric et al. (2007)
? Cl-C ???
0.5 Mn in PS II Yano et al. (2005)
0.02 Fe in myoglobin Denisov et al. (2007)

## Slide 119
Holton & Frankel (2010) Acta D 66 393-408.

## Slide 120
The number of photons you get
before a given crystal dies
is independent
of data collection time
1 um3 = 106 photons
http://www.nist.gov/pml/data/xcom/

## Slide 121: Decisions, Decisions, Decisions
Exposure time
Number of images
Wavelengths
Beam: size, flux, divergence, bandpass
Inverse beam?
Kappa? - overlaps
Multiple crystals? - non-isomorphism
Run Strategy

## Slide 122: Dose slicing
crystal’s useful life
N
photons
N
photons
N
photons
unacceptable
damage
unacceptable
read noise

## Slide 123: What if…
You only have a few small crystals?
Should you:
  Collect 360° from each?
  Collect 10° from each at 36x exposure?
  Glue 36 xtals together, then collect 360° ?
  Glue together and do 12960° ?

## Slide 124
Resolution (Å)
CC vs right answer
“true” resolution limit

## Slide 125
processing procedure | Dose slice collection scenario |  |  | 
 | fine | coarse | glue | gluefine
XDS/XSCALE | 3.03874 | 3.02970 | 3.0270 | 3.02948
XDS/aimless | 3.13351 | 3.03531 | 3.02638 | 3.05292
XDS/noscale | 3.01449 | 3.03048 | 3.02576 | 3.31013
“true” resolution vs strategy
301,640,334 photons

## Slide 126
processing procedure | collection scenario |  |  | 
 | fine | coarse | glue | glue_fine
XDS/XSCALE | 3.03874 | 3.02970 | 3.0270 | 3.02948
XDS/aimless | 3.13351 | 3.03531 | 3.02638 | 3.05292
XDS/noscale | 3.01449 | 3.03048 | 3.02576 | 3.31013
MOSFLM-defaults | x | x | x | x
MOSFLM-cheat | 3.51254 | 3.04355 | 3.08244 | x
MOSFLM/noscale | 3.51311 | 3.04287 | 3.21664 | x
HKL2000-defaults | x | x | x | x
HKL2000-cheat | 3.04147 | 3.01415 | 3.01415 | 3.03948
DIALS-defaults | x | x | 3.32378 | x
DIALS-cheat | 2.92007 | x | 2.98204 | 2.92007
DIALS-noscale | 2.96392 | x | 3.26126 | 5.14852
d*Trek | ? | ? | ? | ?
EVAL15 | ? | ? | ? | ?

## Slide 127
Where:
IDL - average damage-limited intensity (photons/hkl) at a given resolution
105 - converting R from μm to m, re from m to Å, ρ from g/cm3 to kg/m3 and MGy to Gy
re - classical electron radius (2.818 x 10-15 m/electron)
h - Planck’s constant (6.626 x 10-34 J∙s)
c - speed of light (299792458 m/s)
fdecayed - fractional progress toward completely faded spots at end of data set
ρ - density of crystal (~1.2 g/cm3)
R - radius of the spherical crystal (μm)
λ - X-ray wavelength (Å)
fNH - the Nave & Hill (2005) dose capture fraction (1 for large crystals)
nASU - number of proteins in the asymmetric unit
Mr - molecular weight of the protein (Daltons or g/mol)
VM - Matthews’s coefficient (~2.4 Å3/Dalton)
H - Howells’s criterion (10 MGy/Å)
θ - Bragg angle
a2 - number-averaged squared structure factor per protein atom (electron2)
Ma - number-averaged atomic weight of a protein atom (~7.1 Daltons)
B - average (Wilson) temperature factor (Å2)
μ - attenuation coefficient of sphere material (m-1)
μen - mass energy-absorption coefficient of sphere material (m-1)
Self-calibrated damage limit
Holton & Frankel (2010) Acta D 66 393-408.
No flux
No symmetry

## Slide 128
Crystal diameter (micron)
Resolution (Å)
B = 0
B = 20
B = 60
B = 170
B = 330
B = 560
B = 850
Bigger is better, but not by much

## Slide 129
Home Lab Resolution (Å)
Synchrotron Resolution (Å)
Predicting resolution limits

## Slide 130
Synchrotron Resolution (Å)
XFEL Resolution (Å)
Predicting resolution limits

## Slide 131: B factor from image analysis
B = 500

## Slide 132: B factor from image analysis
B = 20

## Slide 133: Muybridge’s galloping horse (1878)
Real space
reciprocal
“Time-resolved” diffraction

## Slide 134: Muybridge’s galloping horse (1878)
Real space
reciprocal
Average electron density

## Slide 135: Muybridge’s galloping horse (1878)
Real space
reciprocal
Average electron density: 1 sigma contour

## Slide 136: lysozyme: real and reciprocal

## Slide 137: lysozyme: thermal motion

## Slide 138: Proteins move! Your crystal may be trying to tell you something…
Beernink, Endrizzi, Alber & Schachman (1999). PNAS USA 96, 5388-5393.

## Slide 139: At the beamline…
Resolution
  problem: background
  solution: use as few pixels as possible
Phases
  problem: fractional errors
  solution: use as many pixels as possible

## Slide 140: Resolution
http://bl831.als.lbl.gov/~jamesh/movies/resolution.mpeg

## Slide 141: Completeness: random deletion
http://bl831.als.lbl.gov/~jamesh/movies/completeness.mpeg

## Slide 142: Completeness: missing wedge
http://bl831.als.lbl.gov/~jamesh/movies/osc.mpeg

## Slide 143: Overloads
http://bl831.als.lbl.gov/~jamesh/movies/overloads.mpeg

## Slide 144: Resolution: low-angle cutoff
http://bl831.als.lbl.gov/~jamesh/movies/lo_cut.mpeg

## Slide 145: The truth about x-ray beams
Term units significance
Flux photons/s duration of experiment
Beam Size μm match to crystal
Divergence mrad spot size vs distance
Wavelength Å resolution and absorption
Dispersion Δλ/λ spot size
Flux density ph/s/area scattering/damage rate
Fluence ph/area scattering/damage

## Slide 146: The truth about x-ray beams
Term units significance
Flux photons/s duration of experiment
Beam Size μm match to crystal
Divergence mrad spot size vs distance
Wavelength Å resolution and absorption
Dispersion Δλ/λ spot size
Flux density ph/s/area scattering/damage rate
Fluence ph/area scattering/damage

## Slide 147: The truth about x-ray beams
quantity | units | home source | MX2
flux | Photons/second | 2.5 x 109 | 2 x 1012
exposure | seconds | 400 | 1

## Slide 148: Decisions, Decisions, Decisions
Exposure time
Number of images
Wavelengths
Beam: size, flux, divergence, bandpass
Inverse beam?
Kappa? - overlaps
Multiple crystals? - non-isomorphism
Run Strategy

## Slide 149: The truth about x-ray beams
quantity | units | home source | MX2
flux | Photons/second | 2.5 x 109 | 2 x 1012
exposure | seconds | 400 | 1
Dispersion | wavelength range / wavelength | 0.2% (Kα1 - Kα2) | 0.014% (Si111)
Divergence | milliRadian | 4.8 | 1.8 (h) 0.18 (v)
Beam size | microns | 100 | 24 (h) 12 (v)
Spectral brightness | Photons/s/mm2/ mR2/0.1%BW | 5.4 x 109 | 1.5 x 1017

## Slide 150: The truth about x-ray beams
quantity | units | home source | MX2
flux | Photons/second | 72 x 106 | 2 x 1012
exposure | time | 4 hours | 1 second
Dispersion | wavelength range / wavelength | 0.2% (Kα1 - Kα2) | 0.014% (Si111)
Divergence | milliRadian | 4.8 | 1.8 (h) 0.18 (v)
Beam size | microns | 17 | 24 (h) 12 (v)
Spectral brightness | Photons/s/mm2/ mR2/0.1%BW | 5.4 x 109 | 1.5 x 1017

## Slide 151: The truth about x-ray beams
quantity | units | home source | MX2
flux | Photons/second | 1 x 106 | 2 x 1012
exposure | time | 10 days | 1 second
Dispersion | wavelength range / wavelength | 0.2% (Kα1 - Kα2) | 0.014% (Si111)
Divergence | milliRadian | 0.6 | 1.8 (h) 0.18 (v)
Beam size | microns | 17 | 24 (h) 12 (v)
Spectral brightness | Photons/s/mm2/ mR2/0.1%BW | 5.4 x 109 | 1.5 x 1017

## Slide 152: The truth about x-ray beams
quantity | units | home source | MX2
flux | Photons/second | 8 x 104 | 2 x 1012
exposure | time | 5 months | 1 second
Dispersion | wavelength range / wavelength | 0.014% (Si111) | 0.014% (Si111)
Divergence | milliRadian | 0.6 | 1.8 (h) 0.18 (v)
Beam size | microns | 17 | 24 (h) 12 (v)
Spectral brightness | Photons/s/mm2/ mR2/0.1%BW | 5.4 x 109 | 1.5 x 1017

## Slide 153: The truth about x-ray beams
Term units significance
Flux photons/s duration of experiment
Beam Size μm match to crystal
Divergence mrad spot size vs distance
Wavelength Å resolution and absorption
Dispersion Δλ/λ spot size
Flux density ph/s/area scattering/damage rate
Fluence ph/area scattering/damage

## Slide 154: spectral dispersion
Ewald sphere
spindle axis
Φ circle
diffracted ray
(h,k,l)
d*
λ*
λ*

## Slide 155: spectral dispersion
Ewald sphere
spindle axis
Φ circle
diffracted ray
(h,k,l)
d*
λ’*
λ’*

## Slide 156: Si(111) vs multilayers
0.014%
Si(111)
1%
multilayer
4 Å
1.9 Å
resolution
spectral
dispersion

## Slide 157: Wide-bandpass MX?
Mono | Si(111) | multilayer | Difference (%) | good
Energy (eV) | 11111 | 12027.5 | +8.2 | Same
Scale | 1.0000 | 1.0964 | +9.6 | Yes
Rmerge (%) | 7.5 | 8.4 | +12 | No
Ranom (%) | 5.6 | 5.5 | -1.8 | Same
I/sd | 13.2 | 12.3 | -6.8 | Same
I/sd (1.5Å) | 2.3 | 1.8 | -21.7 | No
Ano Patt | 5.96 | 4.75 | -20.3 | Same
Ano Phased | 41 | 39.33 | -4.1 | Same
FOM | 0.238 | 0.192 | -19.3 | Same
CC | 0.555 | 0.647 | -6.1 | Yes
Rcryst (%) | 26.70 | 26.48 | -0.2 | Same
Ligand peak | 8.213 | 8.250 | +0.4 | Same

## Slide 158: Wide bandpass reveals mosaicity from a still
Snell EH, Weisgerber S, Helliwell JR, Holzer K & Schroer K (1995)."Improvements in lysozyme protein crystal perfection through microgravity growth", Acta crystallographica. Section D, Biological crystallography 51, 1099-1102.
low mosaic
high mosaic
Δradial = 2ηD/cos22θ

## Slide 159: Decisions, Decisions, Decisions
Exposure time
Number of images
Wavelengths
Beam: size, flux, divergence, bandpass
Inverse beam?
Kappa? - overlaps
Multiple crystals? - non-isomorphism
Run Strategy

## Slide 160: The truth about x-ray beams
Term units significance
Flux photons/s duration of experiment
Beam Size μm match to crystal
Divergence mrad spot size vs distance
Wavelength Å resolution and absorption
Dispersion Δλ/λ spot size
Flux density ph/s/area scattering/damage rate
Fluence ph/area scattering/damage

## Slide 161: beam divergence
Ewald sphere
spindle axis
Φ circle
diffracted ray
(h,k,l)
d*
λ*
λ*

## Slide 162: beam divergence
spindle axis
Φ circle
diffracted ray
(h,k,l)
d*
Ewald sphere
λ*
λ*

## Slide 163: Divergence Arndt & Wonacott (1977)

## Slide 164: divergence = 0 º

## Slide 165: divergence = 0.3 º

## Slide 166: The truth about x-ray beams
Term units significance
Flux photons/s duration of experiment
Beam Size μm match to crystal
Divergence mrad spot size vs distance
Wavelength Å resolution and absorption
Dispersion Δλ/λ spot size
Flux density ph/s/area scattering/damage rate
Fluence ph/area scattering/damage

## Slide 167: Air absorption
Distance from sample (mm)
transmittance

## Slide 168: The truth about x-ray beams
Term units significance
Flux photons/s duration of experiment
Beam Size μm match to crystal
Divergence mrad spot size vs distance
Wavelength Å resolution and absorption
Dispersion Δλ/λ spot size
Flux density ph/s/area scattering/damage rate
Fluence ph/area scattering/damage

## Slide 169: Decisions, Decisions, Decisions
Exposure time
Number of images
Wavelengths
Beam: size, flux, divergence, bandpass
Inverse beam?
Kappa? - overlaps
Multiple crystals? - non-isomorphism
Run Strategy

## Slide 170: beam size vs xtal size
  Put your crystal into the beam

## Slide 171: Put the “crystal” into the beam
Evans et al. (2011)."Macromolecular microcrystallography", Crystallography Reviews 17, 105-142.

## Slide 172: “crystal” = thing you want to shoot
Membrane Protein Expression Center © 2012

## Slide 173: beam size vs xtal size
  Put your crystal into the beam
  Shoot the whole crystal

## Slide 174: shoot the whole crystal

## Slide 175: shoot the whole crystal

## Slide 176: shoot the whole crystal

## Slide 177: How many crystals do you see?
Shoot the “crystal” (singular)

## Slide 178: mosaic spread
Ewald sphere
spindle axis
Φ circle
diffracted ray
(h,k,l)
λ*
λ*
d*
d*

## Slide 179: mosaic spread
Ewald sphere
spindle axis
Φ circle
diffracted rays
(h,k,l)
d*
d*

## Slide 180: mosaic spread = 0 º

## Slide 181: mosaic spread = 0.1º

## Slide 182: mosaic spread = 0.2º

## Slide 183: mosaic spread = 0.4º

## Slide 184: mosaic spread = 0.6º

## Slide 185: mosaic spread = 0.8º

## Slide 186: mosaic spread = 1.0º

## Slide 187: mosaic spread = 1.5º

## Slide 188: mosaic spread = 2.0º

## Slide 189: mosaic spread = 2.5º

## Slide 190: mosaic spread = 3.2º

## Slide 191: mosaic spread = 6.4º

## Slide 192: mosaic spread = 12.8º

## Slide 193
~0 MGy

## Slide 194
5 MGy

## Slide 195
11 MGy

## Slide 196
16 MGy

## Slide 197
22 MGy

## Slide 198
26 MGy

## Slide 199: shoot the whole crystal
Do not bend!

## Slide 200: platy crystals

## Slide 201: platy crystals

## Slide 202: MiTeGen MicroMesh grids
http://www.mitegen.com/products/micromeshes/micromeshes.shtml

## Slide 203: platy crystals

## Slide 204: platy crystals

## Slide 205: platy crystals

## Slide 206: beam size vs xtal size
  Put your crystal into the beam
  Shoot the whole crystal
  Shoot nothing but the crystal

## Slide 207: shoot nothing but the crystal

## Slide 208: shoot nothing but the crystal

## Slide 209
X-ray beam size (μm)
Resolution (Å)
too big
too small
just right

## Slide 210: Gaussian beam
50% of flux
the other 50%

## Slide 211: Collimated beam
100% of flux
0%

## Slide 212: shoot nothing but the crystal

## Slide 213: shoot nothing but the crystal

## Slide 214: shoot nothing but the crystal

## Slide 215: X-ray scattering “rules”:
1 μm crystal ≈ 1 μm water
≈ 1 μm plastic
≈ 0.1 μm glass
≈ 1000 μm air
~ 50 mm He
≈ 370 mm sat water vapor @ 4C
Optimum: support thickness = crystal thickness

## Slide 216: beam size vs xtal size
  Put your crystal into the beam
  Shoot the whole crystal
  Shoot nothing but the crystal
  Back off!

## Slide 217: Mosquito tips cost $0.15 each
So do Pilatus pixels!
We bought 6,224,001
Use them!
(they are reusable)

## Slide 218: Background scattering

## Slide 219: Background scattering
$100,000.00
$100,000.00
$100,000.00
$100,000.00
$100,000.00
$100,000.00
$100,000.00
$100,000.00
real estate is
expensive
use it!

## Slide 220: Fine Slicing
Pflugrath, J. W. (1999)."The finer things in X-ray diffraction data collection", Acta Cryst. D 55, 1718-1725.
background
background

## Slide 221: Pilatus pile-up for RT MX? same photons, different speeds
Sum of 193 shots with
193-fold attenuation

## Slide 222: Pilatus pile-up for RT MX? same photons, different speeds
1 shot with
no attenuation

## Slide 223: Pilatus pile-up for RT MX? same photons, different speeds
0.016° mosaic

## Slide 224: Pilatus pile-up for RT MX? same photons, different speeds
20 deg/s
0.016 deg mosaic

## Slide 225: Pilatus pile-up for RT MX? same photons, different speeds

## Slide 226: Classes of error in MX
Dependence on signal
Time
none sqrt proportional
none
1/sqrt
1/prop.
CCD
Read-out
Photon
counting
Beam flicker
Shutter jitter
Sample vibration
Pile-up
Detector calibration
attenuation
partiality
Non-isomorphism
Radiation damage

## Slide 227: Beam Flicker
time (seconds)
normalized flux through pinhole
pinhole
removed

## Slide 228: This is an Aperture!

## Slide 229: Classes of error in MX
Dependence on signal
Time
none sqrt proportional
none
1/sqrt
1/prop.
CCD
Read-out
Photon
counting
Beam flicker
Shutter jitter
Sample vibration
Pile-up
Detector calibration
attenuation
partiality
Non-isomorphism
Radiation damage

## Slide 230: Optimal exposure time (faint spots on CCD)
thr Optimal exposure time for data set (s)
tref exposure time of reference image (s)
bgref background level near weak spots on
reference image (ADU)
bg0 ADC offset of detector (ADU)
bghr optimal background level (via thr)
σ0 rms read-out noise (ADU)
gain ADU/photon
m multiplicity of data set (including partials)
adjust exposure
so this is ~100

## Slide 231: Optimal exposure time (faint spots on PAD)
thr Optimal exposure time for data set (s)
tro read-out time (s)
2.03 ms on PILATUS3 S
3 μs on EIGER

## Slide 232: Optimal exposure time (anomalous on PAD)
thr Optimal exposure time for data set (s)
tro read-out time (s)
2.03 ms on PILATUS3 S
3 μs on EIGER

## Slide 233: Decisions, Decisions, Decisions
Exposure time
Number of images
Wavelengths
Beam: size, flux, divergence, bandpass
Inverse beam?
Kappa? - overlaps
Multiple crystals? - non-isomorphism
Run Strategy

## Slide 234: How to avoid time-dependent error?
Attenu-wait!
Apply 10x attenuation
Expose 10x longer
Number of photons should be the same!
Damage will also be the same!

## Slide 235: beam size vs xtal size
  Put your crystal into the beam
  Shoot the whole crystal
  Shoot nothing but the crystal
  Back off!
  The crystal must rotate

## Slide 236: The crystal rotates!

## Slide 237: The crystal rotates!
Membrane Protein Expression Center © 2012

## Slide 238: Shoot nothing but the “crystal”
Evans et al. (2011)."Macromolecular microcrystallography", Crystallography Reviews 17, 105-142.

## Slide 239: Decisions, Decisions, Decisions
Exposure time
Number of images
Wavelengths
Beam: size, flux, divergence, bandpass
Inverse beam?
Kappa? - overlaps
Multiple crystals? - non-isomorphism
Run Strategy

## Slide 240: avoiding overlaps
c
c

## Slide 241: avoiding overlaps
c

## Slide 242
SmarAct’s
“SmarGon”
Commercialized
PRIGo

## Slide 243: At the beamline…
Resolution
  problem: background
  solution: use as few pixels as possible
Phases
  problem: fractional errors
  solution: use as many pixels as possible

## Slide 244: R-factor
http://bl831.als.lbl.gov/~jamesh/movies/rfactor.mpeg

## Slide 245: Figure of Merit
http://bl831.als.lbl.gov/~jamesh/movies/dephase.mpeg

## Slide 246
anomalous signal
Crick, F. H. C. & Magdoff, B. S. (1956) Acta Crystallogr. 9, 901-908.
Hendrickson, W. A. & Teeter, M. M. (1981) Nature 290, 107-113.
# sites
MW (Da)
ΔF
F
≈ 1.2 f”
√
World record!
ΔF/F = 0.5%
Wang, Dauter & Dauter (2006) Acta Cryst. D 62, 1475-1483.

## Slide 247: Fractional error
mult > (—)2
~3%
<ΔF/F>

## Slide 248
Holton & Frankel (2010) Acta D 66 393-408.

## Slide 249: Systematic error does not “average out”

## Slide 250: Can you count to 1,000,000 ?
= 0.1%
sqrt(1,000,000)
1,000,000
= 3%
sqrt(1,000)
1,000
> 1000 is a waste!
photon
spot
Theoretically:
In reality:
ISa ~ 33
Rmeas ≈ 0.1% ?
ISa = 1000
Rmeas = ≈ 3%

## Slide 251
Required signal-to-noise (I/σ)
Solve-able proteins (%)
Current
technology
Goal
Threshold of a revolution in phasing

## Slide 252
Source of error | realistic simulation | No SHSSS | Perfect detector
Photon counting | + | + | +
Shutter jitter | + | + | -
Beam flicker | + | + | -
Sample absorption | + | + | -
Radiation damage | + | + | -
Imperfect spindle | + | + | -
vignette | + | + | -
Corner correction | + | + | -
SHSSS | + | - | -
Rmeas (∞-10 Å) | 2.8% | 0.7% | 0.7%
I/σ asymptotic | 26.8 | 74.2 | 81.0
Threshold of a revolution in phasing
Holton et al (2014) "R-factor gap", FEBS Journal 281, 4046-4060.

## Slide 253
Spatial Heterogeneity in Sharp Spot Sensitivity

## Slide 254
Spatial Heterogeneity in Sharp Spot Sensitivity

## Slide 255
Spot centroid position (pixels)
Relative spot intensity
Spatial Heterogeneity in Sharp Spot Sensitivity

## Slide 256
Spot centroid position (pixels)
Relative spot intensity
Spatial Heterogeneity in Sharp Spot Sensitivity

## Slide 257: SHSSS! The dominant source of error for anomalous difference measurements
Spatial
Heterogeneity in
Sharp
Spot
Sensitivity
Flat
Field

## Slide 258: SHSSS! The dominant source of error for anomalous difference measurements
Spatial
Heterogeneity in
Sharp
Spot
Sensitivity
Flat
Field
90%
110%
x 1.11 =
x 0.91 =
100%
100%

## Slide 259: SHSSS! The dominant source of error for anomalous difference measurements
Spatial
Heterogeneity in
Sharp
Spot
Sensitivity
Sharp
Spot
20%
110%
x 1.11 =
x 0.91 =
100%
22%

## Slide 260: SHSSS! The dominant source of error for anomalous difference measurements
Spatial
Heterogeneity in
Sharp
Spot
Sensitivity
Sharp
Spot
90%
30%
x 1.11 =
x 0.91 =
27.3%
100%
Integral: 127.3

## Slide 261: SHSSS! The dominant source of error for anomalous difference measurements
Spatial
Heterogeneity in
Sharp
Spot
Sensitivity
Sharp
Spot
20%
110%
x 1.11 =
x 0.91 =
100%
22.2%
Integral: 122.2

## Slide 262: SHSSS! The dominant source of error for anomalous difference measurements
Spatial
Heterogeneity in
Sharp
Spot
Sensitivity
Sharp
Spot
90%
30%
x 1.11 =
x 0.91 =
27.3%
100%
Integral: 127.3

## Slide 263
Spatial Heterogeneity in Sharp Spot Sensitivity

## Slide 264
down
Spatial Heterogeneity in Sharp Spot Sensitivity

## Slide 265
down
up
Spatial Heterogeneity in Sharp Spot Sensitivity

## Slide 266
down
up
Rseparate
Spatial Heterogeneity in Sharp Spot Sensitivity

## Slide 267
odd
even
Rmixed
Spatial Heterogeneity in Sharp Spot Sensitivity

## Slide 268
separate:
2.5%
Spatial Heterogeneity in Sharp Spot Sensitivity

## Slide 269
separate:
mixed:
2.5%
0.9%
Spatial Heterogeneity in Sharp Spot Sensitivity

## Slide 273
separate:
mixed:
2.5%
0.9%
2.5%2-0.9%2 = 2.3%2
Spatial Heterogeneity in Sharp Spot Sensitivity

## Slide 274
3.5
3.0
2.5
2.0
1.5
1.0
0.5
SHSSS
Systematic component of Rsseparate (%)
distance between spots (mm)
0.1 1 10 100
CCD detector
anomalous mates are always on
different modules
different
modules
Pilatus
SN001
Spatial Heterogeneity in Sharp Spot Sensitivity
Notes: Comparision of SHSSS measured by the detector-shift experiment on three different detectors:
The Q315R in ALS 8.3.1
The first Pilatus detector at the Swiss Light Soruce
A newer Pilatus detector at Stanford
Improvements in calibration technology have finally made
Pilatus competitive with calibration on CCDs

## Slide 275: Pilatus: 1-module shift
Notes: Insulin diffraction recorded at Swiss light source detector centered on beam

## Slide 276
Pilatus: 1-module shift
Notes: Same insulin pattern, but with detector raised by one module width. This image can then be “digitally shifted” to overlay the same spots on top of one another, but recorded by different pixels.

## Slide 277
Pilatus: 1-module shift
Notes: Subtraction of equivalent pixels after shifting the detector
Note that some modules are mis-aligned relative to a perfect grid. These sub-pixel shifts can be resolved with “triangle binning”

## Slide 278
Pilatus: 1-module shift - aligned
Notes: Subtracted image after correcting for sub-pixel shifts

## Slide 279
3% high
average
3% low
Pilatus: 1-module shift - ratio
Notes: Ratio image: difference divided by average, again after correcting for sub-pixel shifts
Note: range of calibration error from module to module is RMS 3%

## Slide 280
3.5
3.0
2.5
2.0
1.5
1.0
0.5
SHSSS
Systematic component of Rsseparate (%)
distance between spots (mm)
0.1 1 10 100
CCD detector
anomalous mates are always on
different modules
different
modules
Pilatus
SN113
Pilatus
SN001
Notes: Comparision of SHSSS measured by the detector-shift experiment on three different detectors:
The Q315R in ALS 8.3.1
The first Pilatus detector at the Swiss Light Soruce
A newer Pilatus detector at Stanford
Improvements in calibration technology have finally made
Pilatus competitive with calibration on CCDs

## Slide 281
Pick-up tool mark
Bragg
glitch
oxygen
inclusions
“tree rings”
1% high
average
1% low
~3x105 photon/pixel
Pilatus: subtract smooth baseline

## Slide 282
N
N
N
θ
λ
detection
event
incoming
photon
PAD: need more than flood-field

## Slide 283
thickness
width
N
N
N
Bragg
glitch
oxygen
inclusion
θ
λ
detection
event
incoming
photon
PAD: need more than flood-field

## Slide 284: CCD calibration: 7235 eV

## Slide 285: CCD calibration: 7247 eV

## Slide 286: Gadox calibration vs energy
photon energy (keV)
Relative absorption depth
same = good!
bad!

## Slide 287: Data collection parameters:
16 crystals
360° each, inverse beam
7235 eV photon energy
< 1 MGy per xtal
Australian Synchrotron MX1
  35 kGy/s into 100 μm x 100 μm

## Slide 288
RESOLUTION COMPLETENESS R-FACTOR I/SIGMA R-meas CC(1/2) Anomal SigAno Nano
LIMIT OF DATA observed Corr
9.17 99.1% 3.9% 257.47 3.9% 100.0* 91* 5.024 450
6.49 100.0% 5.2% 214.33 5.2% 100.0* 86* 3.836 882
5.30 100.0% 7.2% 165.13 7.2% 100.0* 76* 3.257 1175
4.59 100.0% 7.2% 175.42 7.3% 100.0* 67* 2.589 1403
4.10 99.9% 7.7% 174.13 7.7% 100.0* 59* 2.264 1594
3.74 99.9% 9.4% 143.09 9.4% 100.0* 49* 1.953 1783
3.47 100.0% 11.2% 120.17 11.2% 100.0* 39* 1.696 1942
3.24 100.0% 14.1% 91.14 14.1% 100.0* 30* 1.333 2103
3.06 99.9% 19.5% 65.79 19.5% 100.0* 23* 1.117 2214
2.90 99.9% 29.0% 44.85 29.1% 99.9* 17* 1.008 2369
2.77 99.9% 40.5% 32.58 40.6% 99.8* 11* 0.901 2493
2.65 99.9% 52.8% 25.16 52.9% 99.8* 10* 0.866 2605
2.54 100.0% 67.4% 19.47 67.6% 99.6* 2 0.804 2705
2.45 100.0% 88.9% 14.58 89.2% 99.2* 4 0.831 2859
2.37 100.0% 109.3% 9.97 109.7% 98.1* 5 0.829 2925
2.29 100.0% 138.2% 6.87 138.9% 96.1* 1 0.760 3037
2.22 100.0% 197.1% 4.03 198.6% 83.5* -1 0.721 3159
2.16 100.0% 227.3% 2.41 230.8% 46.9* -1 0.677 3224
2.10 61.2% 154.4% 1.28 163.6% 47.0* -2 0.660 1999
2.05 47.9% 170.1% 0.68 196.5% 25.7* 3 0.629 1578
total 93.3% 15.7% 54.30 15.8% 100.0* 12* 1.217 42499
140-fold multiplicity: 16 crystals, 360° each, inverse beam, 7235 eV

## Slide 289: 140-fold multiplicity
18 σ
Phased anomalous difference Fourier
16 σ

## Slide 290: 140-fold multiplicity
15σ = PO4
Phased anomalous difference Fourier

## Slide 291: 140-fold multiplicity
~2σ = Mg?
Phased anomalous difference Fourier

## Slide 292: 140-fold multiplicity
8.2σ = Mg or Na
DELFAN residual anomalous difference

## Slide 293: non-isomorphism
Why doesn’t everyone do this?

## Slide 294: Non-isomorphism in lysozyme
RH: 84.2%
vs 71.9%
Riso = 44.5%
3aw6
3aw7
Δcell = 0.7 %
RMSD = 0.18 Å

## Slide 295
h,k,l | structure factors (F) |  |  | 
 | #1 | #2 | #3 | #4
5,3,4 | 523.7 | 559.8 | 579.9 | 603.2
5,4,4 | 168.2 | 166.6 | 177.2 | 196.1
5,5,4 | 34.9 | 26.4 | 19.2 | 17.3
6,1,4 | 305.7 | 301.1 | 298.1 | 296.3
6,2,4 | 353.0 | 353.9 | 356.2 | 366.9
6,3,4 | 300.9 | 285.3 | 273.5 | 259.4
6,3,5 | 223.8 | 226.3 | 234.6 | 251.4
… | … | … | … | …
X-ray Data Sets
h,k,l | 100% | 17 % | 13 % | 9 %
 | 1st | 2nd | 3rd | 4th
5,3,4 | -1.46 | -0.02 | 1.84 | -0.72
5,4,4 | -0.88 | -0.29 | -0.34 | 0.15
5,5,4 | -0.42 | -0.65 | 1.02 | 1.47
6,1,4 | -0.90 | -0.85 | 1.44 | -0.40
6,2,4 | -1.20 | -0.37 | 0.67 | 0.01
6,3,4 | -0.75 | 0.31 | 0.48 | 0.00
6,3,5 | -0.75 | -0.85 | 0.72 | -0.82
… | … | … | … | …
data
set
Singular value decomp.
SVD
vector
value
Singular values & vectors
Correlation
Coefficients
-0.68
CC2
CC3
CC1
CC1 CC2 CC3
data set #1
Data sets
Positioned
in
“Correlation
Space”
-0.24
data set #2
data set #3
0.32 …

## Slide 297
Correlation to 2nd & 3rd singular vectors
Structure factor (electrons)

## Slide 298
Correlation to 2nd & 3rd singular vectors
Structure factor (electrons)

## Slide 299
Correlation to 2nd & 3rd singular vectors
Structure factor (electrons)

## Slide 300: Bragg’s “minimum wavelength principle”
Perutz (1985). “Early Days of Cryst…” Methods in Enzymology, Vol. 114, 3-18.
Bragg & Perutz (1952)."external form haemoglobin I", Acta Cryst. 5, 277-283.

## Slide 301: DMMULTI – fake data 4 deg rotation: 8 “xtals”: before
CC = 0.02

## Slide 302
CC = 0.8
DMMULTI – fake data 4 deg rotation: 8 “xtals”: after

## Slide 303: Suggested anomalous protocol:
2 wavelengths are better than 1
- (peak + inf)/2, and remote
MAD, not M-SAD!
360° in < 5 MGy
move detector
4X exposure
goto 2

## Slide 304: Decisions, Decisions, Decisions
Exposure time
Number of images
Wavelengths
Beam: size, flux, divergence, bandpass
Inverse beam?
Kappa? - overlaps
Multiple crystals? - non-isomorphism
Run Strategy

## Slide 305: Optimal exposure time (CCD)
thr Optimal exposure time for data set (s)
tref exposure time of reference image (s)
bgref background level near weak spots on
reference image (ADU)
bg0 ADC offset of detector (ADU)
bghr optimal background level (via thr)
σ0 rms read-out noise (ADU)
gain ADU/photon
m multiplicity of data set (including partials)
adjust exposure
so this is ~100

## Slide 306: Optimal exposure time (PAD)
thr Optimal exposure time for data set (s)
tro read-out time (s)
2.03 ms on PILATUS3 S
3 μs on EIGER

## Slide 307: Decisions, Decisions, Decisions
Exposure time
Number of images
Wavelengths
Beam: size, flux, divergence, bandpass
Inverse beam?
Kappa? - overlaps
Multiple crystals? - non-isomorphism
Run Strategy

## Slide 309: Decisions, Decisions, Decisions
Exposure time
Number of images
Wavelengths
Beam: size, flux, divergence, bandpass
Inverse beam?
Kappa? - overlaps
Multiple crystals? - non-isomorphism
Run Strategy

## Slide 310: Gadox calibration vs energy
photon energy (keV)
Relative absorption depth
same = good!
bad!

## Slide 311: attenuation factor
Bouguer, P. (1729). Essai d'optique sur la gradation de la lumière.
Lambert, J. H. (1760). Photometria: sive De mensura et gradibus luminis, colorum et umbrae. E. Klett.
Beer, A. (1852)."Bestimmung der Absorption des rothen Lichts in farbigen Flüssigkeiten", Ann. Phys. Chem 86, 78-90.
A = = exp[-μxtal(txi+ txo)
-μsolvent(tsi + tso)]
IT
Ibeam
μxtal
txi
txo
tsi
tso
txi
txo
tsi
tso
txi
txo
tsi
tso
μsolvent

## Slide 312: Decisions, Decisions, Decisions
Exposure time
Number of images
Wavelengths
Beam: size, flux, divergence, bandpass
Inverse beam?
Kappa? - overlaps
Multiple crystals? - non-isomorphism
Run Strategy

## Slide 313
X-ray beam size (μm)
Resolution (Å)
too big
too small
just right

## Slide 314: Decisions, Decisions, Decisions
Exposure time
Number of images
Wavelengths
Beam: size, flux, divergence, bandpass
Inverse beam?
Kappa? - overlaps
Multiple crystals? - non-isomorphism
Run Strategy

## Slide 315: Systematic error does not “average out”

## Slide 316: Decisions, Decisions, Decisions
Exposure time
Number of images
Wavelengths
Beam: size, flux, divergence, bandpass
Inverse beam?
Kappa? - overlaps
Multiple crystals? - non-isomorphism
Run Strategy

## Slide 317
SmarAct’s
“SmarGon”
Commercialized
PRIGo

## Slide 318: Decisions, Decisions, Decisions
Exposure time
Number of images
Wavelengths
Beam: size, flux, divergence, bandpass
Inverse beam?
Kappa? - overlaps
Multiple crystals? - non-isomorphism
Run Strategy

## Slide 319
Correlation to 2nd & 3rd singular vectors
Structure factor (electrons)

## Slide 320: Decisions, Decisions, Decisions
Exposure time
Number of images
Wavelengths
Beam: size, flux, divergence, bandpass
Inverse beam?
Kappa? - overlaps
Multiple crystals? - non-isomorphism
Run Strategy

## Slide 321: URL Summary
100 ADU/pixel
10 μm for lysozyme
~3% error per spot, 1%/MGy
7235 eV for S-SAD on CCDs
“Attenu-wait” & dose slicing
http://bl831.als.lbl.gov/~jamesh/powerpoint/
IGBMC_SvN_2016.pptx
http://bl831.als.lbl.gov/xtalsize.html
http://bl831.als.lbl.gov/xtallife.html