## Slide 1: ALS BL 8.3.1
First Light!
June 8 2001
Tom-A-Tron
L
B
E
R

## Slide 2: Measured success rate: ALS 8.3.1
images x 106
1250
1000
750
500
250
0
PDB entries
Notes: Of the 1275 PDB deposits involving ALS 8.3.1, 1038 of them mention a data collection date. The timeline between collection and deposition for each PDB entry is shown in blue. For example, in April 2004 only 46 PDBs had been deposited, but the data for 252 had already been collected. The red line indicates the cumulative number of images collected at a given date. The kink in the red line in mid-2003 is coincident with an optics upgrade that tripled the flux. In 2014 another flux upgrade brought another factor of 4.5.

## Slide 3: The 100 structure/year club
PDBs | Sync | BL | Flux
2829 | APS | 19-ID | 1.3E+13
2110 | NSLS | X29A | 2.9E+11
1830 | APS | 22-ID | 7.0E+12
1651 | SLS | X06SA | 4.0E+12
1517 | SLS | X10SA | 2.5E+12
1475 | ESRF | ID29 | 1.0E+13
1447 | SSRF | BL17U | 4.0E+12
1384 | ESRF | ID14-4 | 5.0E+12
PDBs | Sync | BL | Flux
1384 | ESRF | ID14-4 | 5.0E+12
1362 | ESRF | ID23-1 | 4.0E+12
1285 | SSRL | BL11-1 | 4.0E+11
1282 | APS | 24-ID-C | 1.3E+13
1153 | ESRF | ID14-1 | 1.0E+11
1137 | SSRL | BL9-2 | 7.0E+11
1122 | ALS | 8.3.1 | 1.8E+11
1042 | APS | 17-ID | 7.0E+11
1026 | NSLS | X25 | 2.4E+11
USA
Europe
Decommissioned
8% of all membrane
protein structures!

## Slide 4: BL 8.3.1 User Science Highlights
Rosenberg OS, Dovala D, Li X, Connolly L, Bendebury A, Finer-Moore J, Holton J, Cheng Y, Stroud RM & Cox JS (2015)."Substrates Control Multimerization and Activation of the Multi-Domain ATPase Motor of Type VII Secretion", Cell 161, 501

## Slide 5: BL 8.3.1 User Science Highlights
Lang PT, Holton JM, Fraser JS & Alber T (2014)."Protein structural ensembles are revealed by redefining X-ray electron density noise", Proc. Nat. Acad. Sci. USA 111, 237.

## Slide 6: BL 8.3.1 User Science Highlights
Zhang C, Spevak W, Zhang Y, Burton EA, Ma Y, Habets G, et al. (2015) “RAF inhibitors that evade paradoxical MAPK pathway activation” Nature. 526, 583-6.

## Slide 7: 8.3.1 beam time distribution 2010
26.5%
Staff/Development
9%
Chart: 
| General User | Staff | UCSF | UCB | Plexxicon | MD Anderson
Sales | 25 | 9 | 26.5 | 26.5 | 8 | 5
General User
Program
25%
26.5%
26.5%
UC Berkeley
8%
5%
9%
Plexxikon

## Slide 8: BL8.3.1 Beam Time Allocation (%)
Year | 2010 | 2011 | 2012 | 2013 | 2014 | 2015
Gen. Users | 25 | 25 | 25 | 25 | 25 | 25
R&D | 9 | 9 | 9 | 9 | 9 | 9
Plexxikon | 10 | 8 | 8 | 6 | 6 | 6
MD Anderson | 10 | 5 | 3 | 5 | 8 | 8
UCSF | 23 | 26.5 | 27.5 | 27.5 | 26 | 18.5
UCB | 23 | 26.5 | 27.5 | 27.5 | 26 | 9.7
Total | 100 | 100 | 100 | 100 | 100 | 100

## Slide 9: BL8.3.1 PRT member contributions ($k)
Year | 2010 | 2011 | 2012 | 2013 | 2014 | 2015
Gen. Users | 0 | 0 | 0 | 0 | 0 | 0
R&D | 0 | 0 | 0 | 0 | 0 | 0
Plexxikon | 250 | 250 | 250 | 125 | 125 | 125
MD Anderson | 116 | 75 | 75 | 130 | 130 | 130
UCSF | 180 | 180 | 180 | 180 | 180 | 967
UCB | 0 | 0 | 0 | 0 | 0 | 248
Total | 546 | 505 | 505 | 435 | 435 | 1585

## Slide 10: BL8.3.1 Non-salary direct costs ($k)
Year | 2010 | 2011 | 2012 | 2013 | 2014 | 2015
general Purchases | 89 | 24 | 66 | 13 | 16 | 24
Cryojet | 37 |  |  |  |  | 
New Computers | 30 |  |  |  |  | 
New M1 mirror |  |  |  |  | 105 | 
M1 mirror mount |  |  |  |  | 60 | 
Mono access |  |  |  |  | 5 | 

## Slide 11: BL8.3.1 Staff Effort
Year | 2010 | 2011 | 2012 | 2013 | 2014 | 2015
sci: James Holton | 0.9 | 0.9 | 0.9 | 0.9 | 0.6 | 0.75
tec: George Meigs | 0.9 | 0.9 | 0.9 | 0.9 | 1.0 | 1.0
adm: Jane Tanamachi | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2
Total | 2.00 | 2.00 | 2.00 | 2.00 | 1.80 | 1.95

## Slide 12: BL8.3.1 Publication Record
Year | 2010 | 2011 | 2012 | 2013 | 2014 | 2015
Refereed | 48 | 38 | 47 | 48 | 36 | 15
High Impact | 20 | 16 | 16 | 16 | 16 | 5

## Slide 13: BL8.3.1 Financial Support: 2016
Award | dates | source
$2,578,000 | 01/01/15 – 12/31/18 | MRPI: University of California Office of the President (not renewable, no overhead)
$357,492 | 12/15/14 – 12/15/16 | UCSF-PBBR: Program for Breakthrough Biomedical Research (not renewable)
$750,000 | 08/22/14 – 09/30/19 | National Institutes of Health Inter-Agency Agreement (non renewable)
$625,000 | 05/01/12 – 04/30/16 | Plexxikon Inc.
$130,000 | 08/12/15 – 08/11/16 | M. D. Anderson Cancer Research Center
Planned: |  | 
$6,567,804 | 04/01/16 – 03/31/21 | NIH P41 – X-ray Structural Biology Resource
$500,000 | 05/01/16 – 04/30/21 | Plexxikon Inc.
$620,000 | 08/12/16 – 08/11/18 | M. D. Anderson Cancer Research Center

## Slide 14: BL8.3.1 PRT member contributions ($k)
Year | 2010 | 2011 | 2012 | 2013 | 2014 | 2015
Gen. Users | 0 | 0 | 0 | 0 | 0 | 0
R&D | 0 | 0 | 0 | 0 | 0 | 0
Plexxikon | 250 | 250 | 250 | 125 | 125 | 125
MD Anderson | 116 | 75 | 75 | 130 | 130 | 130
UCSF | 180 | 180 | 180 | 180 | 180 | 967
UCB | 0 | 0 | 0 | 0 | 0 | 248
Total | 546 | 505 | 505 | 435 | 435 | 1585

## Slide 15: BL8.3.1 Beam Time Allocation (%)
Year | 2010 | 2011 | 2012 | 2013 | 2014 | 2015
Gen. Users | 25 | 25 | 25 | 25 | 25 | 25
R&D | 9 | 9 | 9 | 9 | 9 | 9
Plexxikon | 10 | 8 | 8 | 6 | 6 | 6
MD Anderson | 10 | 5 | 3 | 5 | 8 | 8
UCSF | 23 | 26.5 | 27.5 | 27.5 | 26 | 18.5
UCB | 23 | 26.5 | 27.5 | 27.5 | 26 | 9.7
Total | 100 | 100 | 100 | 100 | 100 | 100

## Slide 16: 8.3.1 beam time distribution: 2016
26.5%
Staff/Development
9%
Chart: 
| General User | Staff | UCSF | UCB | UCLA | UCI | UCD | UCSD | Plexxikon | MD Anderson
Sales | 25 | 9 | 18.5 | 9.7 | 6.2 | 6.2 | 6.2 | 6.2 | 5 | 8
General User
Program
25%
9.7%
18.5%
UC Berkeley
6%
9%
Plexxikon
5%
6.2%
6.2%
6.2%
6.2%
UC Davis
UC Santa Cruz
UCLA
UC Irvine

## Slide 17: Grand Challenges in Structural Biology for the 21st Century
Phase Problem
native elements - universally
Amplitude Problem
    Improving resolution
R-factor Gap
    Publishing / being right

## Slide 18: Grand Challenges in Structural Biology for the 21st Century
Phase Problem
detector calibration
Amplitude Problem
    disorder
R-factor Gap
    Limited representation

## Slide 19: Grand Challenges in Structural Biology for the 21st Century
Phase Problem
detector calibration
Amplitude Problem
    disorder
R-factor Gap
    Limited representation

## Slide 20
Pick-up tool mark
Bragg
glitch
oxygen
inclusions
“tree rings”
1% high
average
1% low
~3x105 photon/pixel
Pilatus: subtract smooth baseline

## Slide 21
Source of error | realistic simulation | No SHSSS | Perfect detector
Photon counting | + | + | +
Shutter jitter | + | + | -
Beam flicker | + | + | -
Sample absorption | + | + | -
Radiation damage | + | + | -
Imperfect spindle | + | + | -
vignette | + | + | -
Corner correction | + | + | -
SHSSS | + | - | -
Rmeas (∞-10 Å) | 2.8% | 0.7% | 0.7%
I/σ asymptotic | 26.8 | 74.2 | 81.0
SHSSS: necessary and sufficient
Holton et al (2014) "R-factor gap", FEBS Journal 281, 4046-4060.

## Slide 22
Required signal-to-noise (I/σ)
Solve-able proteins (%)
Current
technology
Goal
Threshold of a revolution in phasing

## Slide 23: ALS 8.3.1: The 6th detector

## Slide 24: HPAD vs dmPAD
N
N
N
N
N
N
Q
Q
Q
trim
trim
trim
trim
trim
trim
gain
gain
gain
= k
= k
= k
θ
θ
λ
λ
thick
width

## Slide 25: Grand Challenges in Structural Biology
Phase Problem
Calibrate it out boost the signal average over it
Amplitude Problem
R-factor Gap
http://bl831.als.lbl.gov/~jamesh/powerpoint/Janelia_thresholds_2015.pptx

## Slide 26: Grand Challenges in Structural Biology
Phase Problem
Calibrate it out boost the signal average over it
Amplitude Problem
R-factor Gap
http://bl831.als.lbl.gov/~jamesh/powerpoint/Janelia_thresholds_2015.pptx

## Slide 27
Photon Energy (eV)
Anomalous f” (electrons)
2472.7 2473.5
Bohic et al. Anal. Chem. 2008; 80(24):9557. doi: 10.1021/ac801817k

## Slide 28: S-MAD Sample size Dilemma

## Slide 29
Ewald sphere 2
diffracted ray
λ*
λ*
θ
1
Ewald sphere
λ*
(h,k,l)
diffracted ray
λ*
θ
d*
(-h,-k,-l)
d*

## Slide 30
h,k,l
-h,-k,-l
Detector
Detector
Detector
Detector
λ = 5 Å
sample
injector
Si(111)
52.87deg
Si(111)
Si(111)
Si(111)
Si(111)
Si(111)
2 Multilayer mirrors
d=2nm, W/B4C
KB Horiz focus
KB vertical focus
~ 1m
Notes: Proposed 5-crystal monochromator for colliding-beam crystallography. The generation of two opposing Ewald spheres means that for every h,k,l index reflected from the protein crystal, the Friedel mate is reflected in the opposite direction, allowing for very accurate anomalous difference measurements. The wavelength need not be set precisely at the sulfur or phosphorous edge, but merely at a wavelength where the f” of those elements are strong. For fs-class pulses, the distances above will have to be precise to withing ~3 microns, and for accurate reproduction of the partiality of the Friedel mates, the angular precision of the colliding beams must be within ~1% of the sample crystal’s rocking width, or ~100 microrads.

## Slide 31: “sweet spot” for sample size ?
2 Å
2sin(90°)
3 Å
2sin(90°)
S
P
Mg
Na
C
F
N
O
Cl
Ar
K
Ca
Ne
Se
4 Å
2sin(90°)

## Slide 32: ALS-U: Small, soft beams

## Slide 33: ALS-U: Small, soft beams
Diamond
($100, reusable)

## Slide 34: ALS-U: Small, soft beams
Diamond
($100, reusable)
oil ?

## Slide 35: ALS-U: Small, soft beams
Diamond
($100, reusable)
oil ?

## Slide 36: ALS-U: Small, soft beams
oil ?
λ=2d sinθ
2.5 Å data
with 5 Å X-rays
Only analytic absorption correction:
International Tables for Crystallography, Vol. C, 2nd ed., chapter 6.3

## Slide 37: Grand Challenges in Structural Biology
Phase Problem
Calibrate it out boost the signal average over it
Amplitude Problem
R-factor Gap
http://bl831.als.lbl.gov/~jamesh/powerpoint/Janelia_thresholds_2015.pptx

## Slide 38: Averaging with systematic error

## Slide 39: 140-fold multiplicity
SUBSET OF INTENSITY DATA WITH SIGNAL/NOISE >= -3.0 AS FUNCTION OF RESOLUTION
RESOLUTION NUMBER OF REFLECTIONS COMPLETENESS R-FACTOR R-FACTOR COMPARED I/SIGMA R-meas CC(1/2) Anomal SigAno Nano
LIMIT OBSERVED UNIQUE POSSIBLE OF DATA observed expected Corr
9.17 188919 1061 1071 99.1% 3.9% 4.8% 188919 257.47 3.9% 100.0* 91* 5.024 450
6.49 356827 1933 1933 100.0% 5.2% 5.5% 356827 214.33 5.2% 100.0* 86* 3.836 882
5.30 466503 2515 2515 100.0% 7.2% 7.0% 466503 165.13 7.2% 100.0* 76* 3.257 1175
4.59 548405 2974 2974 100.0% 7.2% 6.6% 548405 175.42 7.3% 100.0* 67* 2.589 1403
4.10 615117 3353 3355 99.9% 7.7% 6.7% 615117 174.13 7.7% 100.0* 59* 2.264 1594
3.74 684947 3734 3737 99.9% 9.4% 8.3% 684947 143.09 9.4% 100.0* 49* 1.953 1783
3.47 743557 4051 4051 100.0% 11.2% 10.1% 743557 120.17 11.2% 100.0* 39* 1.696 1942
3.24 799722 4366 4366 100.0% 14.1% 13.9% 799722 91.14 14.1% 100.0* 30* 1.333 2103
3.06 839252 4598 4603 99.9% 19.5% 20.2% 839252 65.79 19.5% 100.0* 23* 1.117 2214
2.90 891398 4904 4908 99.9% 29.0% 31.7% 891398 44.85 29.1% 99.9* 17* 1.008 2369
2.77 931622 5148 5151 99.9% 40.5% 44.8% 931622 32.58 40.6% 99.8* 11* 0.901 2493
2.65 969629 5384 5387 99.9% 52.8% 58.8% 969629 25.16 52.9% 99.8* 10* 0.866 2605
2.54 997213 5574 5574 100.0% 67.4% 76.0% 997213 19.47 67.6% 99.6* 2 0.804 2705
2.45 1031222 5889 5889 100.0% 88.9% 101.2% 1031222 14.58 89.2% 99.2* 4 0.831 2859
2.37 788617 6008 6008 100.0% 109.3% 125.5% 788617 9.97 109.7% 98.1* 5 0.829 2925
2.29 614145 6252 6252 100.0% 138.2% 161.4% 614145 6.87 138.9% 96.1* 1 0.760 3037
2.22 470283 6481 6481 100.0% 197.1% 231.7% 470283 4.03 198.6% 83.5* -1 0.721 3159
2.16 259836 6629 6631 100.0% 227.3% 268.7% 259831 2.41 230.8% 46.9* -1 0.677 3224
2.10 36259 4169 6807 61.2% 154.4% 169.4% 36229 1.28 163.6% 47.0* -2 0.660 1999
2.05 13567 3372 7037 47.9% 170.1% 187.0% 13503 0.68 196.5% 25.7* 3 0.629 1578
total 12247040 88395 94730 93.3% 15.7% 16.4% 12246941 54.30 15.8% 100.0* 12* 1.217 42499
16 crystals, 360° each, < 1 MGy, inverse beam
7235 eV ADSC Q210r Australian Synchrotron MX1
I/SIGMA
257.47

## Slide 40: 140-fold multiplicity
18 σ
Phased anomalous difference Fourier
16 σ

## Slide 41: 140-fold multiplicity
8.2σ = Na
DELFAN residual anomalous difference
10 σ is enough for phasing!
Bunkoczi et al. (2015)."weak, single-wavelength anomalous", Nat. Meth. 12, 127-130.
Not
Lysozyme!
f” = 0.15

## Slide 42: Non-isomorphism in lysozyme
RH: 84.2%
vs 71.9%
Riso = 44.5%
3aw6
3aw7
Δcell = 0.7 %
RMSD = 0.18 Å

## Slide 43
h,k,l | structure factors (F) |  |  | 
 | #1 | #2 | #3 | #4
5,3,4 | 523.7 | 559.8 | 579.9 | 603.2
5,4,4 | 168.2 | 166.6 | 177.2 | 196.1
5,5,4 | 34.9 | 26.4 | 19.2 | 17.3
6,1,4 | 305.7 | 301.1 | 298.1 | 296.3
6,2,4 | 353.0 | 353.9 | 356.2 | 366.9
6,3,4 | 300.9 | 285.3 | 273.5 | 259.4
6,3,5 | 223.8 | 226.3 | 234.6 | 251.4
… | … | … | … | …
X-ray Data Sets
h,k,l | 100% | 17 % | 13 % | 9 %
 | 1st | 2nd | 3rd | 4th
5,3,4 | -1.46 | -0.02 | 1.84 | -0.72
5,4,4 | -0.88 | -0.29 | -0.34 | 0.15
5,5,4 | -0.42 | -0.65 | 1.02 | 1.47
6,1,4 | -0.90 | -0.85 | 1.44 | -0.40
6,2,4 | -1.20 | -0.37 | 0.67 | 0.01
6,3,4 | -0.75 | 0.31 | 0.48 | 0.00
6,3,5 | -0.75 | -0.85 | 0.72 | -0.82
… | … | … | … | …
data
set
Singular value decomp.
SVD
vector
value
Singular values & vectors
Correlation
Coefficients
-0.68
CC2
CC3
CC1
CC1 CC2 CC3
data set #1
Data sets
Positioned
in
“Correlation
Space”
-0.24
data set #2
data set #3
0.32 …

## Slide 45
Correlation to 2nd & 3rd singular vectors
Structure factor (electrons)

## Slide 46
Correlation to 2nd & 3rd singular vectors
Structure factor (electrons)

## Slide 47
Correlation to 2nd & 3rd singular vectors
Structure factor (electrons)

## Slide 48: Bragg’s “minimum wavelength principle”
Perutz (1985). “Early Days of Cryst…” Methods in Enzymology, Vol. 114, 3-18.
Bragg & Perutz (1952)."external form haemoglobin I", Acta Cryst. 5, 277-283.

## Slide 49: DMMULTI – fake data 4 deg rotation: 8 “xtals”: before
CC = 0.02

## Slide 50
CC = 0.8
DMMULTI – fake data 4 deg rotation: 8 “xtals”: after

## Slide 51: Grand Challenges in Structural Biology
Phase Problem
detector calibration
Amplitude Problem
    disorder
R-factor Gap
    Limited representation

## Slide 52: The Amplitude Problem
10 Å

## Slide 53: lysozyme: real and reciprocal

## Slide 54: lysozyme: thermal motion

## Slide 55: Muybridge’s galloping horse (1878)
Real space
reciprocal
“Time-resolved” diffraction

## Slide 56: Muybridge’s galloping horse (1878)
Real space
reciprocal
Average electron density

## Slide 57: B factor from image analysis
B = 500

## Slide 58: B factor from image analysis
B = 20

## Slide 59
Crystal diameter (micron)
Resolution (Å)
B = 0
B = 20
B = 60
B = 170
B = 330
B = 560
B = 850
Bigger is better, but not by much

## Slide 60
Synchrotron Resolution (Å)
XFEL Resolution (Å)
Predicting resolution limits

## Slide 61
Dissecting the Amplitude Problem
purify
protein
grow
crystal
cryo
MX
micro
beam
XFEL
traditional
approach
AUORA
auto optical
re-alignment
B factor
low

## Slide 62: Automatic and Unattended Optical Re-Alignment (AUORA)
beamline
microscope
reference
image

## Slide 63: Automatic and Unattended Optical Re-Alignment (AUORA)
search

## Slide 64: Automatic and Unattended Optical Re-Alignment (AUORA)
search

## Slide 65: Automatic and Unattended Optical Re-Alignment (AUORA)
search

## Slide 66: Automatic and Unattended Optical Re-Alignment (AUORA)
search

## Slide 67: Automatic and Unattended Optical Re-Alignment (AUORA)
search

## Slide 68: Automatic and Unattended Optical Re-Alignment (AUORA)
match

## Slide 69: Automatic and Unattended Optical Re-Alignment (AUORA)
move

## Slide 70: Automatic and Unattended Optical Re-Alignment (AUORA)
aligned
4.5 s Acquire 360°
3.0 s Algorithm
5.0 s Motor moves

## Slide 93
pinhole
prism
microscope
Styrofoam™ backlight
backstop
Zero-parallax simplifies alignment

## Slide 94: Zero-parallax simplifies alignment
visible light

## Slide 95: Zero-parallax simplifies alignment
X-ray

## Slide 96
MX
area
detector
beam center
3%
2%
1%
0
crystal in loop
Cu foil
X-ray excited
fluorescent
X-ray source
x-ray beam
Cone-Beam Online X-ray
Absorption Radiography
(CBOXAR)

## Slide 97: Cone-Beam Online X-ray Absorption Radiograph (CBOXAR)

## Slide 98: Flexible Targeting Platform
Use optical image to annotate targets
Correlate with other “maps”
  CBOXAR
  raster image scores
Compatible with:
  Ultraviolet, Infrared, SONICC, polarization, fluorescence, XRF, VISEX, …
AUORA: tie it all together

## Slide 99
Dissecting the Amplitude Problem
purify
protein
grow
crystal
cryo
MX
micro
beam
XFEL
traditional
approach
AUORA
auto optical
re-alignment
CBOXAR
& raster
B factor
low

## Slide 100: “rastering” to find the crystal
Membrane
protein crystal
X-ray beam
ice

## Slide 101: Why not just “raster” everything?
Rastering Needs:
Small beam
Very high flux
fast detector
Rapid image analysis
Small search area

## Slide 102: The 100 structure/year club
PDBs | Sync | BL | Flux
2829 | APS | 19-ID | 1.3E+13
2110 | NSLS | X29A | 2.9E+11
1830 | APS | 22-ID | 7.0E+12
1651 | SLS | X06SA | 4.0E+12
1517 | SLS | X10SA | 2.5E+12
1475 | ESRF | ID29 | 1.0E+13
1447 | SSRF | BL17U | 4.0E+12
1384 | ESRF | ID14-4 | 5.0E+12
PDBs | Sync | BL | Flux
1384 | ESRF | ID14-4 | 5.0E+12
1362 | ESRF | ID23-1 | 4.0E+12
1285 | SSRL | BL11-1 | 4.0E+11
1282 | APS | 24-ID-C | 1.3E+13
1153 | ESRF | ID14-1 | 1.0E+11
1137 | SSRL | BL9-2 | 7.0E+11
1122 | ALS | 8.3.1 | 1.8E+11
1042 | APS | 17-ID | 7.0E+11
1026 | NSLS | X25 | 2.4E+11
USA
Europe
Decommissioned

## Slide 103: “The flattest thing we have ever seen”
Old mirror: 800 nRad
New mirror: < 200 nRad
Dec 17 2014
Flux at
460%
- LBNL Metrology Group

## Slide 104: “8-shooter” adjustable beam size
rotary
pinholes:
300
100
75
50
30
25
10
“small”

## Slide 105: Low cost, portable rastering solution
Sample crystal
In cryo-loop
(not to scale)
New
CMOS detector
($70,000)
20 frames/s
photon/noise = 2
Existing MX detector
0.4 frames/s
Photon/noise = 0.4
Existing
Retractable
Front cover
http://researchdetectors.com/cmos_1m.html
Collaboration: 8.3.1 and MBC beamlines

## Slide 106: Rastering is robust, but slow
Damage & diffraction = flux / area
Gy/s = ph/s /2000 / um2
Example: flux= 1012 ph/s, 1 mm2 beam = 500 Gy/s
100% sample life → 30 MGy / 500 Gy/s = 17 hours
10% sample life → 3 MGy / 500 Gy/s = 1.7 hours
1% sample life → 300 kGy / 500 Gy/s = 10 min
Holton (2009)."beginner's guide rad dam", J. Synch. Rad. 16, 133-142.

## Slide 107: Rastering is robust, but slow
Damage & diffraction = flux / area
Gy/s = ph/s /2000 / um2
Example: 10% life, 1 mm2 sample, flux= 1012 ph/s
10x10 μm beam: 0.6 s
100x100 μm beam: 60s x 10x10 shots =
1% life & 1013 ph/s → 1 min
Holton (2009)."beginner's guide rad dam", J. Synch. Rad. 16, 133-142.
1% life → 10 min
1.7 hours
x 100x100 shots =
1.7 hours

## Slide 108: Rastering is robust, but slow
Damage & diffraction = flux / area
time/mm2 is fixed by flux
Example: ALS 8.3.1 – 200 kGy/s = 15 s for 10% lifedose
50 μm hole: 20x20 shots = 210 s/mm2
30 μm hole: 33x33 shots = 320 s/mm2
10 μm hole: 100x100 shots = 2730 s/mm2
10 μm hole + multilayer: 70 s/mm2
Holton (2009)."beginner's guide rad dam", J. Synch. Rad. 16, 133-142.

## Slide 109: Rastering is robust, but slow
Damage & diffraction = flux / area
time/mm2 is fixed by flux
Example: ALS 8.3.1 – 200 kGy/s ; 1% lifedose
50 μm hole: 20x20 shots = 21 s/mm2
30 μm hole: 33x33 shots = 32 s/mm2
10 μm hole: 100x100 shots = 273 s/mm2
10 μm hole + multilayer: 7 s/mm2
Holton (2009)."beginner's guide rad dam", J. Synch. Rad. 16, 133-142.

## Slide 110: Rastering is robust, but slow
Damage & diffraction = flux / area
time/mm2 is fixed by flux
Example: ALS 8.3.1 – 200 kGy/s ; 1% lifedose
50 μm hole: 20x20 shots = 21 s/mm2 20 Hz
30 μm hole: 33x33 shots = 32 s/mm2 35 Hz
10 μm hole: 100x100 shots = 273 s/mm2 37 Hz
10 μm hole + multilayer: 7 s/mm2 2 kHz
Holton (2009)."beginner's guide rad dam", J. Synch. Rad. 16, 133-142.

## Slide 111: Si(111) vs multilayers
0.014%
Si(111)
1%
multilayer
4 Å
1.9 Å
resolution
spectral
dispersion

## Slide 112: Narrow bandpass does not improve anomalous signal
Chart: Photon energy (eV)
| 12656 | 12657 | 12658 | 12658.1 | 12658.2 | 12658.6 | 12658.9 | 12659 | 12659.1 | 12659.2 | 12659.3 | 12659.4 | 12659.6 | 12659.7 | 12659.9 | 12660.4 | 12660.5 | 12660.6 | 12661 | 12656 | 12657 | 12658 | 12658.1 | 12658.2 | 12658.6 | 12658.9 | 12659 | 12659.1 | 12659.2 | 12659.3 | 12659.4 | 12659.6 | 12659.7 | 12659.9 | 12660.4 | 12660.5 | 12660.6 | 12661
I(+) | 159900 | 171800 | 190300 | 193800 | 195000 | 228800 | 232700 | 240200 | 253200 | 241400 | 260200 | 267800 | 243300 | 256400 | 251000 | 260200 | 233100 | 234600 | 247000
I(-) | 131500 | 100900 | 81560 | 76100 | 83960 | 69860 | 50120 | 60980 | 74270 | 66830 | 85140 | 76810 | 82330 | 69980 | 103800 | 106000 | 89630 | 108800 | 110400
vertical
divergence
(240 μRad)
100 μm
pinhole
Si(333)
Si(333)
Al
absorber
1rb5
h,k,l = 5,5,0

## Slide 113: ALS robots
5.0.2
5.0.1
5.0.3
MBC 4.2.2
ACTOR
8.2.2
8.3.1
Cool Hand Luke
8.2.1
12.3.1 SIBYLS
ACTOR
Berkeley
Auto-
Mounter
ALS MX robots
Stanford
Auto-
Mounter

## Slide 114: Pin compatibility
SLAC
ALS
SPINE

## Slide 115: SuperTong

## Slide 116: Hampton Pin
SuperTong

## Slide 117: Syrrx Pin
SuperTong

## Slide 118: plastic Pin
SuperTong

## Slide 119: Yale Pin
SuperTong

## Slide 120: SuperTong
SuperTong

## Slide 121: “infinite capacity” sample carousel

## Slide 122: Carousel open

## Slide 123
Dissecting the Amplitude Problem
purify
protein
grow
crystal
cryo
MX
micro
beam
traditional
approach
high
in-situ MX
Control for handling
B factor
XFEL
AUORA
auto optical
re-alignment
CBOXAR
& raster
low

## Slide 124
Notes: non-rotating tray goniometer stage installed in ALS 8.3.1. Note that the “normal” cryo-pin goniometer on the left is still in place and a brass flight tube has been added to contain the air scatter from the main beam on its way to the tray goniometer.

## Slide 125
Dissecting the Amplitude Problem
purify
protein
grow
crystal
cryo
MX
micro
beam
+
traditional
approach
high
in-situ MX
Control for handling
XFEL
Rapid RT
high multiplicity
RT MX
Control for cryo
B factor
AUORA
auto optical
re-alignment
CBOXAR
& raster
low

## Slide 126
Dissecting the Amplitude Problem
purify
protein
grow
crystal
cryo
MX
micro
beam
design
new clone /
conditions
+
-
traditional
approach
high
low
XFP
xtal packing
interfaces
in-situ MX
Control for handling
RT MX
Control for cryo
B factor
XFEL
XFEL
Rapid RT
high multiplicity
AUORA
auto optical
re-alignment
CBOXAR
& raster

## Slide 127
Dissecting the Amplitude Problem
purify
protein
grow
crystal
cryo
MX
micro
beam
design
new clone /
conditions
+
-
traditional
approach
high
low
SAXS
order/disorder
In solution
XFP
xtal packing
interfaces
in-situ MX
Control for handling
RT MX
Control for cryo
B factor
XFEL
XFEL
Rapid RT
high multiplicity
AUORA
auto optical
re-alignment
CBOXAR
& raster

## Slide 128: Grand Challenges in Structural Biology
Phase Problem
detector calibration
Amplitude Problem
    disorder
R-factor Gap
    Limited representation

## Slide 129: “R-factor Gap” for macromolecules
MLFSOM
Elves
Rmerge = 6%
Rwork = 7%
Rfree = 8%
single-conformer PDB file
1H87; conf “A”
Holton et al. (2014) "The R-factor gap", FEBS J. 281, 4046-4060.

## Slide 130: The R-factor Gap
Holton et al. (2014) "The R-factor gap", FEBS J. 281, 4046-4060.

## Slide 131: “R-factor Gap” for macromolecules
MLFSOM
Elves
Rmerge = 6%
Rwork = 17%
Rfree = 20%
multi-conformer PDB file
1H87
Holton et al. (2014) "The R-factor gap", FEBS J. 281, 4046-4060.

## Slide 132: “R-factor Gap” for macromolecules
MLFSOM
Elves
Rmerge = 6%
Rwork = 7%
Rfree = 8%
multi-conformer PDB file
1H87
Holton et al. (2014) "The R-factor gap", FEBS J. 281, 4046-4060.

## Slide 133: The R-factor Gap
Last “right”
atom added
Holton et al. (2014) "The R-factor gap", FEBS J. 281, 4046-4060.

## Slide 134: The R-factor Gap
MX data quality is not limiting
Hypotheses:
  Hidden systematic error?
  Model-building problem?
  Representation problem?

## Slide 135: nearBragg program
http://bl831.als.lbl.gov/~jamesh/nearBragg/
“assumption-free” total scattering
no Fourier Transform
no unit cells
no “mosaicity”
arbitrary “atoms”
arbitrary “source”
coherent or not

## Slide 136: Simulation of diffuse scatter
http://bl831.als.lbl.gov/example_data_sets/Illuin/lyso_DS

## Slide 137: fastBragg program
http://bl831.als.lbl.gov/~jamesh/fastBragg/
Optimized for nanocrystals
Square or round
Takes h,k,l and F
Supports 1 unit cell
no “mosaicity”
arbitrary “source”
arbitrary “phi” steps

## Slide 138: Simulation of XFEL stills
real

## Slide 139: Simulation of XFEL stills
fake

## Slide 140: Simulation of XFEL stills
phiz
phiy
phiz
phix
phix
phiy
VS
real

## Slide 141: Is it Real? Or is it MLFSOM
simulated
real

## Slide 142
Simulated | statistic | Real
5.4% | Rmerge | 6.2%
18.9 | I/sd | 16.2
1.9 | I/sd (1.4 Ǻ) | 1.6
1.7 4.0 0.020 | SDCORR | 1.0 2.2 0.065
36.8 | PADFPH | 31.46
3.871 | mlphare f” | 3.476
0.178 | FOM | 0.192
0.545 | FOMDM | 0.664
0.6270 | CC(1H87) | 0.6090
0.070 | Rwork | 0.184
0.080 | Rfree | 0.231

## Slide 143: MLFSOM: now available in “Pilatus”

## Slide 144: Pilatus pile-up for RT MX? same photons, different speeds
1 shot with
no attenuation

## Slide 145: Pilatus pile-up for RT MX? same photons, different speeds
Sum of 193 shots with
193-fold attenuation

## Slide 146: Global Challenge: Se phasing
Trivial to solve
100% SeMet incorporation
MLFSOM
simulation

## Slide 147: Global Challenge: Se phasing
Impossible to solve
100% S -Met incorporation
MLFSOM
simulation

## Slide 148: Global Challenge: Se phasing
fraction Sulfur
correlation coefficient

## Slide 149: Global Challenge: Se phasing
fraction Sulfur
correlation coefficient

## Slide 150: Global Challenge: Se phasing
fraction Sulfur
correlation coefficient

## Slide 151: Global Challenge: Se phasing
fraction Sulfur
correlation coefficient

## Slide 152: Global Challenge: Se phasing
fraction Sulfur
correlation coefficient

## Slide 153: Global Challenge → new software

## Slide 154: New software → benefits users

## Slide 155
Holton & Frankel (2010) Acta D 66 393-408.

## Slide 158: Grand Challenges in Structural Biology
Phase Problem
Calibrate it out boost the signal average over it
Non-isomorphism phasing (NIP)
Amplitude Problem
Disorder > everything Dissect sample prep
“Meta-beamline”
R-factor Gap
1 electron =chemistry validate dynamics
2D data vs 4D models
http://bl831.als.lbl.gov/~jamesh/powerpoint/MOU_831_2016.pptx

## Slide 159: Concluding remarks
8.3.1 finances stabilizing
Need new scientist
Poised for major hardware upgrades
Pilatus, multilayer, de-focus
Focus on biggest problems
Phases, amplitudes, systematics