## Slide 1: Maximum Data Quality Workshop
Made possible by:
UC Office of the President, Multicampus Research Programs and Initiatives (MRPI) grant MR‐15‐328599
Update on the Instrument

## Slide 2: MRPI Technical Highlights
Dectris Pilatus3 6M installation - May 2016
  25 Hz framing, 2 ms dead time, previously 0.3 Hz
  0.8% calibration error, previously 2.5%
  Repaired and re-installed 6 times, 10% replaced
Complete control system update
  Shutterless data collection
  “Rastering” automated sample alignment
  Re-write 64k lines of code
Overhaul of data archiving infrastructure
  LTO4 to LTO6, support USB3 user hard drives
First remote data collection – Apr 2017

## Slide 3
Pilatus3 S 6M
Pixel Array
Detector
6x106 pixels
$0.15/each
Single-photon
counting
25 Hz images

## Slide 4: ALS 8.3.1 data collection history
images x 106

## Slide 5: Shutterless rastering now available
Membrane
protein crystal
X-ray beam
ice

## Slide 6: ALS robots
5.0.2
5.0.1
5.0.3
MBC 4.2.2
ACTOR
8.2.2
8.3.1
Cool Hand Luke
8.2.1
12.3.1 SIBYLS
ACTOR
Berkeley
Auto-
Mounter
ALS MX robots
Stanford
Auto-
Mounter

## Slide 7
als-enable.lbl.gov

## Slide 8
ber
structural
bio
portal.org

## Slide 9: Pin compatibility
SLAC
ALS
SPINE

## Slide 10: SuperTong

## Slide 11: What if we didn’t have to choose?
Caltech
SSRL
Yale
Oxford
SPring-8
SPINE
MiTeGen RT
Syrrx
H. magnetic
H. hybrid
H. copper
H. screw cap
H. colored
Hampton screw cap
Hampton magnetic

## Slide 12: “infinite capacity” sample carousel

## Slide 13: Carousel open

## Slide 14
pinhole
mirror
microscope
SBS
format
crystallization
tray
Notes: Inspired by the ESRF microcrystal camera, the superbend endstations are all equipped with zero-paralax optics. The user sees their sample from the point of view of the x-ray beam. Sample alignment is not only more stable but more intuitive than an off-axis camera. Placing the crystal into the microscope crosshairs gaurantees that it will be hit by the beam.

## Slide 15
Notes: non-rotating tray goniometer stage installed in ALS 8.3.1. Note that the “normal” cryo-pin goniometer on the left is still in place and a brass flight tube has been added to contain the air scatter from the main beam on its way to the tray goniometer.

## Slide 16
NatX-ray
CrystalQuick™ X
250 μm
~50 μm tape
Fluidigm
Topaz® DC
200 μm
PDMS
PDMS
MiTeGen
in-situ-1
100 μm
~50 μm tape
in-situ diffraction trays
CrystalDirect
~50 μm tape
25 μm
Microlytic
Crystal Former™
1000 μm
Notes: Next slide please

## Slide 17: scattering/absorptoin “rules”:
1 μm crystal ≈ 1 μm water
≈ 1 μm plastic
≈ 0.1 μm glass
≈ 1000 μm air
Optimum: support thickness = crystal thickness

## Slide 18: Match crystal to beam
X-ray beam size (μm)
Resolution (Å)
too big
too small
just right

## Slide 19: “8-shooter” adjustable beam size
rotary
pinholes:
300
100
75
50
30
25
15

## Slide 20: Crystal lifetime vs pinhole
pinhole (µm) | beam size | Cryo (min) | Room T (s)
100 | 60x80 | 4 | 1.7
75 | 60x78 | 5 | 2.0
50 | 60x54 | 8 | 3.4
30 | 54x34 | 6 | 2.3
20 | 44x24 | 6 | 2.4
15 | 40x19 | 7 | 2.8
10 | 34x14 | 10 | 4.1

## Slide 21: Optimal exposure time (anomalous on a PAD)
time
intensity
2.03 ms

## Slide 22: Optimal exposure time (anomalous on a PAD)
tmad Optimal exposure time (s)
tro read-out time (s)
2.03 ms on Pilatus3 S
0.95 ms on Pilatus3 X
Pilatus S
>0.2 s/frame
Upgrade:
>0.1 s/frame
8.3.1 is too
hot for Pilatus
Attenuate for MAD

## Slide 23: % exposure_time.com
The last tune up was at 11111.0 eV (1.11587 A) convergance: 2 x 0.35
with Dose Rate 47280 Gy/s
Based on a Dose Rate of 4.73e+04 Grays/Sec at 2 mRad with slits 100.00% flux, the Dose Rate now at 2.0000 mRad, or 100.00% flux, is 4.73e+04 Grays/Sec
What is the current, or expected dose rate (Gy/s) [4.73e+04] : -->
Expected 1/2 life dose of typical protein crystals
Native 2e+7
MAD 5e+6
S SAD 3e+6
Br/I/RNA 1e+6
RT 2e+5
What is the Max total dose YOU want: [2e+7] -->
Total time for dataset must not exceed 422.8 seconds
How many frames will be in this data set [1800]: -->
Based on these parameters, the exposure time can be up to 0.235 seconds/frame, or 4.73e+04 Gy/s

## Slide 24
1% high
average
1% low
Relative Pixel
Calibration
ALS 8.3.1
Pilatus3 6M
0.9793 Å x-rays
RMS = 0.8%

## Slide 25: S-SAD merging 20 xtals
Knott et al. (2017). "Guide-bound CRISPR-Cas13a." Nature S&MB online

## Slide 26
Chart: 
| 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10
F+ | 52.3965 | 57.5371 | 48.523 | 47.5964 | 34.153 | 48.9106 | 70.3415 | 52.7871 | 59.3327 | 50.0622
F- | 46.3965 | 51.5371 | 42.523 | 41.5964 | 28.153 | 42.9106 | 64.3415 | 46.7871 | 53.3327 | 44.0622
observation
intensity
Systematic vs random error

## Slide 27: ALS 8.3.1 capabilities
200s data sets in shutterless mode
Support for room temperature
in-situ screening from trays
15-100 μm beam sizes
True BLU-ICE interface
Excellent detector calibration
Infinite capacity sample carousel
Automated data backup
Remote operation

## Slide 28: How do I get time on ALS 8.3.1?
Check for available shifts on the schedule:
http://bl831.als.lbl.gov/
Contact Jane Tanamachi!
JTanamachi@LBL.GOV
510-495-2404
No proposals, no reviews, no kidding

## Slide 29: BL8.3.1 online schedule

## Slide 30: ALS 8.3.1 beam time allocation
PRT Member | allocation | Details
General Users | 25% | Mandated by DOE
Staff | 9% | development & testing
Plexxikon | 5% | 5-year contract just renewed
UC | 35.5% total | MRPI ends Dec 2018
UCSF | 18.5% | Sandler, ETAC & NIH IAA
UCB | 7.1% | Same as other UCs
Total | 600 shifts | 200 days/year

## Slide 31: Please Acknowledge!
UC Office of the President, Multicampus Research Programs and Initiatives (MRPI) grant MR‐15‐328599

## Slide 32: Please email !
Has MRPI support led to any additional extramural funds being brought into UC?
Describe in lay terms one major accomplishment or success of your project to date.
JTanamachi@LBL.GOV

## Slide 33: non-isomorphism

## Slide 34
0.5% cell change → 15%
Crick & Magdoff (1956) “theory method isomorphous replacement”, Acta Cryst. 9, 901-8.
Same structure, different cell

## Slide 35
Same structure, different cell
0.5% cell change → 15%
Crick & Magdoff (1956) “theory method isomorphous replacement”, Acta Cryst. 9, 901-8.

## Slide 36
Same cell, rotated protein
0.5° rotation → 16%
Crick & Magdoff (1956) “theory method isomorphous replacement”, Acta Cryst. 9, 901-8.

## Slide 37
Same cell, rotated protein
0.5° rotation → 16%
Crick & Magdoff (1956) “theory method isomorphous replacement”, Acta Cryst. 9, 901-8.

## Slide 38
Uniform stretch
5% change → 0%

## Slide 39
5% change → 0%
Uniform stretch

## Slide 40: Non-isomorphism in lysozyme
RH: 84.2%
vs 71.9%
Riso = 44.5%
3aw6
3aw7
Δcell = 0.7 %
RMSD = 1.70 Å
RMSD = 0.18 Å

## Slide 41: Hygrostatic gradients
loop
jet
high
water activity
low

## Slide 42: 3aw6

## Slide 43: 3aw6 3aw7

## Slide 44: 3aw6, 3aw7, “bent” 3aw7

## Slide 45: 3aw6, “bent” 3aw7

## Slide 46: 3aw6

## Slide 47: “bent” 3aw7

## Slide 48: 3aw7

## Slide 49: “map bender” procedure
Convert to fractional coordinate shifts

## Slide 51: Basis function 1 0 0

## Slide 52: “map bender” procedure
Convert to fractional coordinate shifts
Basis functions: spatial sin waves
  Repeats with unit cell
  Low-order h,k,l indices first
Fit function to Δx vs x,y,z across P1 cell
Repeat with Δy, Δz
Apply to PDB file and interpolate map grid
Usually 3-5 orders comparable to LSQ

## Slide 53: Photosystem O=O bond?
Suga, M., et al., site of O=O bond formation in PSII
caught by XFEL. Nature, 2017. 543: p. 131.

## Slide 54: Fo – Fc maps
5ws5 vs 5ws6

## Slide 55
Fo – Fo maps
5ws5 vs 5ws6