## Slide 1: What is “MRPI”?
Multi-campus Research Programs and Initiatives
  Offered by: UC Office of Research and Graduate Studies
Wide range of fields
  Nanotech, Arts, Climate, Gender Studies, Sustainability, Health
ALS 8.3.1 MRPI Project Goals:
  Upgrade ALS 8.3.1 to latest technology
  Open up access to all UC structural biology labs
  Transition to sustainable funding sources
Notes: Brief description of what MRPI is. The University of California seeks to support collaboration between campuses in ways that benefit the entire system the most.
Funding was very limited, and it is extremely remarkable that we got this and our budget was not cut.

## Slide 2: What is ALS Beamline 8.3.1 ?
Macromolecular Crystallography (MX)
  MAD, 1x1012 ph/s, Pilatus 6M
  Technology development
    Room temp, Elves, multi-crystal
  Staff: Holton, Meigs, Tanamachi, Classen*
Landmark and highly productive facility
  1st ALS SuperBend beamline – constructed 2001
  1 of 10 in USA with >1000 PDBs/10 years running
  8% of all membrane protein structures
  996 distinct users from 57 institutions
  1407 structures, 636 pubs, 227 high-impact
45% MRPI
20% MRPI
25% * 6 month
Notes: ALS 8.3.1 is a very productive facility. Our focus is on maximizing efficiency of the user experience.
The MRPI funds are not just to maintain this facility, but to upgrade it. The new detector is 50x faster, more sensitive and more accurate than the existing detector, and an optics upgrade will match these capabilities.
We will exceed the modern standard of screening and data collection by offering completely automated mail-in raster-based screening
This will position ALS 8.3.1 as the centerpiece of structural biology for the UC system

## Slide 5: Maximum Data Quality Workshop
Made possible by:
UC Office of the President, Multicampus Research Programs and Initiatives (MRPI) grant MR‐15‐328599
Update on the Instrument

## Slide 6: It’s Official

## Slide 7
Apr 20 2016

## Slide 9: New Pilatus3 6M at ALS 8.3.1
Apr 20 2016

## Slide 10: New Pilatus3 6M at ALS 8.3.1
13:41 9-22-2016 Plexxkon’s first Pilatus shift

## Slide 11: MRPI Technical Highlights
Dectris Pilatus3 6M installation - May 2016
  25 Hz framing, 2 ms dead time, previously 0.3 Hz
  0.8% calibration error, previously 2.5%
  Repaired and re-installed 6 times, 10% replaced
Complete control system update: true-BLU-ICE
  Shutterless data collection
  “Rastering” automated sample alignment
  Re-write 64k lines of code
Overhaul of data archiving infrastructure
  LTO4 to LTO6, support USB3 user hard drives
First remote data collection – Apr 2017

## Slide 12: ALS 8.3.1 data collection history
images x 106

## Slide 13: Shutterless rastering now available
Membrane
protein crystal
X-ray beam
ice

## Slide 14
BLU-ICE “screening tab” almost integrated

## Slide 15: Pin compatibility
SLAC
ALS
SPINE

## Slide 16: What if we didn’t have to choose?
Caltech
SSRL
Yale
Oxford
SPring-8
SPINE
MiTeGen RT
Syrrx
H. magnetic
H. hybrid
H. copper
H. screw cap
H. colored
Hampton screw cap
Hampton magnetic

## Slide 17: SuperTong

## Slide 18: “infinite capacity” sample carousel

## Slide 19: Carousel open

## Slide 20
pinhole
mirror
microscope
SBS
format
crystallization
tray
Notes: Inspired by the ESRF microcrystal camera, the superbend endstations are all equipped with zero-paralax optics. The user sees their sample from the point of view of the x-ray beam. Sample alignment is not only more stable but more intuitive than an off-axis camera. Placing the crystal into the microscope crosshairs gaurantees that it will be hit by the beam.

## Slide 21
Notes: non-rotating tray goniometer stage installed in ALS 8.3.1. Note that the “normal” cryo-pin goniometer on the left is still in place and a brass flight tube has been added to contain the air scatter from the main beam on its way to the tray goniometer.

## Slide 22
NatX-ray
CrystalQuick™ X
250 μm
~50 μm tape
Fluidigm
Topaz® DC
200 μm
PDMS
PDMS
MiTeGen
in-situ-1
100 μm
~50 μm tape
in-situ diffraction trays
CrystalDirect
~50 μm tape
25 μm
Microlytic
Crystal Former™
1000 μm
Notes: Next slide please

## Slide 23: scattering/absorptoin “rules”:
1 μm crystal ≈ 1 μm water
≈ 1 μm plastic
≈ 0.1 μm glass
≈ 1000 μm air
Optimum: support thickness = crystal thickness

## Slide 24: Match crystal to beam
X-ray beam size (μm)
Resolution (Å)
too big
too small
just right

## Slide 25: “8-shooter” adjustable beam size
rotary
pinholes:
300
100
75
50
30
25
15

## Slide 26: Crystal lifetime vs pinhole
pinhole (µm) | beam size | Cryo (min) | Room T (s)
100 | 60x80 | 4 | 1.7
75 | 60x78 | 5 | 2.0
50 | 60x54 | 8 | 3.4
30 | 54x34 | 6 | 2.3
20 | 44x24 | 6 | 2.4
15 | 40x19 | 7 | 2.8
10 | 34x14 | 10 | 4.1

## Slide 27: PAD full vs partial
iteration
Summed intensity

## Slide 28: Optimal exposure time (anomalous on a PAD)
time
intensity
2.03 ms

## Slide 29: Optimal exposure time (anomalous on a PAD)
tmad Optimal exposure time (s)
tro read-out time (s)
2.03 ms on Pilatus3 S
0.95 ms on Pilatus3 X
Pilatus S
>0.2 s/frame
Upgrade:
>0.1 s/frame
8.3.1 is too
hot for Pilatus
Attenuate for MAD

## Slide 30: % exposure_time.com
The last tune up was at 11111.0 eV (1.11587 A) convergance: 2 x 0.35
with Dose Rate 47280 Gy/s
Based on a Dose Rate of 4.73e+04 Grays/Sec at 2 mRad with slits 100.00% flux, the Dose Rate now at 2.0000 mRad, or 100.00% flux, is 4.73e+04 Grays/Sec
What is the current, or expected dose rate (Gy/s) [4.73e+04] : -->
Expected 1/2 life dose of typical protein crystals
Native 2e+7
MAD 5e+6
S SAD 3e+6
Br/I/RNA 1e+6
RT 2e+5
What is the Max total dose YOU want: [2e+7] -->
Total time for dataset must not exceed 422.8 seconds
How many frames will be in this data set [1800]: -->
Based on these parameters, the exposure time can be up to 0.235 seconds/frame, or 4.73e+04 Gy/s

## Slide 31
1% high
average
1% low
Relative Pixel
Calibration
ALS 8.3.1
Pilatus3 6M
0.9793 Å x-rays
RMS = 0.8%

## Slide 32
Required signal-to-noise (I/σ)
Solve-able proteins (%)
CCD
old Pilatus
Goal
Threshold of a revolution in phasing
8.3.1’s Pilatus

## Slide 33: S-SAD merging 20 xtals
Knott et al. (2017). "Guide-bound CRISPR-Cas13a." Nature S&MB online

## Slide 34
Chart: 
| 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10
F+ | 52.3965 | 57.5371 | 48.523 | 47.5964 | 34.153 | 48.9106 | 70.3415 | 52.7871 | 59.3327 | 50.0622
F- | 46.3965 | 51.5371 | 42.523 | 41.5964 | 28.153 | 42.9106 | 64.3415 | 46.7871 | 53.3327 | 44.0622
observation
intensity
Systematic vs random error

## Slide 35: ALS 8.3.1 capabilities
200s data sets in shutterless mode
Support for room temperature
in-situ screening from trays
15-100 μm beam sizes
True BLU-ICE interface
Excellent detector calibration
Infinite capacity sample carousel
Automated data backup
Remote operation

## Slide 36: How do I get time on ALS 8.3.1?
Check for available shifts on the schedule:
http://bl831.als.lbl.gov/
Contact Jane Tanamachi!
JTanamachi@LBL.GOV
510-495-2404
No proposals, no reviews, no kidding

## Slide 37: BL8.3.1 online schedule

## Slide 38: ALS 8.3.1 beam time allocation 2017-18
PRT Member | allocation | Details
General Users | 25% | Mandated by DOE
Staff | 9% | development & testing
Plexxikon | 5% | 5-year contract
UC | 35.5% total | MRPI
UCSF | 18.5% | Sandler, ETAC & NIH IAA
UCB | 7.1% | Same as other UCs
Total | 600 shifts | 200 days/year

## Slide 39: Please Acknowledge!
UC Office of the President, Multicampus Research Programs and Initiatives (MRPI) grant MR‐15‐328599
and NIGMS:
P30 GM124169
R01 GM124149
http://bl831.als.lbl.gov/~jamesh/powerpoint/
MRPI_intro_2017.pptx
MRPI_raddam_2017.pptx

## Slide 40: Please email !
Has MRPI support led to any additional extramural funds being brought into UC?
Describe in lay terms one major accomplishment or success of your project to date.
JTanamachi@LBL.GOV