## Slide 1: Maximum Data Quality Workshop
Made possible by:
UC Office of the President, Multicampus Research Programs and Initiatives (MRPI) grant MR‐15‐328599
RADIATION DAMAGE!
What you need to know

## Slide 2
Damage is done
by dose (MGy)
proportional to photons/area
not time
not heat
Sliz P, Harrison SC & Rosenbaum G (2003). Structure 11, 13-19.
Garman EF & McSweeney SM (2006). J. Sync. Rad. 14, 1-3.
Owen RL, Rudino-Pinera E & Garman EF (2006). PNAS 103, 4912-4917.
Leiros et al. (2006). Acta Cryst. D 62, 125-132.
Holton JM (2007). J. Synch Rad. 14, 51-72.

## Slide 3
Holton (2009) J. Synchrotron Rad. 16 133-42
Radiation Damage World Records
MGy reaction reference
~45 global damage Owen et al. (2006)
10/Å global damage Howells et al. (2009)
5 Se-Met Holton (2007)
4 Hg-S Ramagopal et al. (2004)
4 R-C-COOH Garman et al. (2015)
3 S-S Murray et al. (2002)
1 Br-RNA Olieric et al. (2007)
~1? Cl-C ???
0.5 Mn in PS II Yano et al. (2005)
0.06 putidaredoxin Corbett et al. (2007)
0.02 Fe in myoglobin Denisov et al. (2007)

## Slide 4: what the is a MGy?
http://bl831.als.lbl.gov/
damage_rates.pdf
Holton J. M. (2009) J. Synchrotron Rad. 16 133-42

## Slide 5
How long will my crystal last?
synch line type flux beamsize flux density dose max xtal min site
ph/s μm ph/μm2/s rate lifetime lifetime
ALS 4.2.2 MAD 1e12 75x80 1.7e+08 124 kGy/s 4 m 16 s
ALS 5.0.1 mono 2e11 100 2.5e+07 13 kGy/s 39 m 2.6 m
ALS 5.0.2 MAD 1.5e12 100 1.5e+08 76.3 kGy/s 6.6 m 26 s
ALS 5.0.3 mono 3e11 100 3.8e+07 19.4 kGy/s 26 m 1.7 m
ALS 8.2.1 MAD 3.5e11 100 4.5e+07 22.7 kGy/s 22 m 88 s
ALS 8.2.2 MAD 3.5e11 100 4.5e+07 22.7 kGy/s 22 m 88 s
ALS 8.3.1 MAD 9e11 70 2.3e+08 119 kGy/s 4.2 m 17 s
ALS 12.3.1 MAD 2e11 65x90 3.4e+07 17.4 kGy/s 29 m 1.9 m
ALS 12.3.1 ML 4.0e13 65x90 6.8e+09 6.89 MGy/s 4.4 s 0.29 s
SSRL 1-5 MAD 1.7e10 200 4.2e+05 202 Gy/s 41 h 2.8 h
SSRL 7-1 mono 2.6e11 200 6.5e+06 3.09 kGy/s 2.7 h 11 m
SSRL 9-1 mono 3.9e10 200 9.8e+05 463 Gy/s 18 h 72 m
SSRL 9-2 MAD 4.8e11 200 1.2e+07 5.7 kGy/s 88 m 5.8 m
SSRL 11-1 MAD 3.9e11 200 9.8e+06 4.63 kGy/s 1.8 h 7.2 m
SSRL 11-3 mono 2.6e10 200 6.5e+05 302 Gy/s 28 h 1.8 h
SSRL 12-2 MAD 4e12 90x5 8.9e+09 5.01 MGy/s 6 s 0.4 s
SSRL 12-1 MAD 4e12 5x5 1.6e+11 90 MGy/s 0.35 s 22 ms
SSRL 12-1 ML 3e14 5x5 1.1e+13 6.4 GGy/s 5 ms 300 us
APS 24-ID-C MAD 1.3e13 20x60 1.1e+10 5.23 MGy/s 5.7 s 0.38 s
APS "typical" MAD 1.5e12 80 2.3e+08 119 kGy/s 4.2 m 17 s
synch line type flux beamsize flux density dose max xtal min site
ph/s μm ph/μm2/s rate lifetime lifetime

## Slide 7
Holton & Frankel (2010) Acta D 66 393-408.

## Slide 8
Where:
IDL - average damage-limited intensity (photons/hkl) at a given resolution
105 - converting R from μm to m, re from m to Å, ρ from g/cm3 to kg/m3 and MGy to Gy
re - classical electron radius (2.818 x 10-15 m/electron)
h - Planck’s constant (6.626 x 10-34 J∙s)
c - speed of light (299792458 m/s)
fdecayed - fractional progress toward completely faded spots at end of data set
ρ - density of crystal (~1.2 g/cm3)
R - radius of the spherical crystal (μm)
λ - X-ray wavelength (Å)
fNH - the Nave & Hill (2005) dose capture fraction (1 for large crystals)
nASU - number of proteins in the asymmetric unit
Mr - molecular weight of the protein (Daltons or g/mol)
VM - Matthews’s coefficient (~2.4 Å3/Dalton)
H - Howells’s criterion (10 MGy/Å)
θ - Bragg angle
a2 - number-averaged squared structure factor per protein atom (electron2)
Ma - number-averaged atomic weight of a protein atom (~7.1 Daltons)
B - average (Wilson) temperature factor (Å2)
μ - attenuation coefficient of sphere material (m-1)
μen - mass energy-absorption coefficient of sphere material (m-1)
Self-calibrated damage limit
Holton & Frankel (2010) Acta D 66 393-408.
No flux
No symmetry

## Slide 9
Crystal diameter (micron)
Resolution (Å)
B = 0
B = 20
B = 60
B = 170
B = 330
B = 560
B = 850
Bigger is better, but not by much

## Slide 10: B factor from image analysis
B = 500

## Slide 11: B factor from image analysis
B = 20

## Slide 12: Can radiation damage be “outrun”?

## Slide 13: Dose-rate effect at Room Temp?
Warkentin M, Badeau R, Hopkins JB, Mulichak AM, Keefe LJ & Thorne RE (2012)."Global radiation damage at 300 and 260 K with dose rates approaching 1 MGy s-1", Acta Cryst. D 68, 124-133.
Owen RL, Axford D, Nettleship JE, Owens RJ, Robinson JI, Morgan AW, Dore AS, Lebon G, Tate CG, Fry EE, Ren J, Stuart DI & Evans G (2012)."Outrunning free radicals in room-temperature macromolecular crystallography", Acta Cryst. D 68, 810-818.

## Slide 14: Dose-rate dependence of damage
dose rate (kGy/s)
maximum useful dose (MGy)

## Slide 15: Pilatus pile-up for RT MX? same photons, different speeds
Sum of 193 shots with
193-fold attenuation
SLOW

## Slide 16: Pilatus pile-up for RT MX? same photons, different speeds
1 shot with
no attenuation
FAST

## Slide 17: Pilatus pile-up for RT MX? same photons, different speeds

## Slide 18: Pilatus pile-up for RT MX? same photons, different speeds
0.016° mosaic

## Slide 19: Pilatus pile-up for RT MX? same photons, different speeds
20 deg/s
0.016 deg mosaic

## Slide 20: Pilatus pile-up for RT MX? same photons, different speeds

## Slide 21: Do we expect this with protein?
Mosaic spread (η) < 0.02°
20°/s rotation
50x50x50 μm lysozyme crystal
Istill = 0.07*(F/mosaic)2
F=130 → 3e6 photons/s/spot
Max Eiger count rate = 3e6 photons/s/spot
35% of Fs > 130 (lysozyme @ 2.0 Å)
45% under-counting expected

## Slide 22: What is a “streak camera”?

## Slide 23: What is a “streak camera”?
0.1 mm/pixel
= 4 ms/pixel
25 mm/s

## Slide 26: Individual spots at 768 kGy/s
Intensity (photons)
Dose (kGy)
0
Same xtal, different spots

## Slide 27: Double-tap: no dark progression at 675 kGy/s
Intensity (photons)
Dose (kGy)

## Slide 28: Xtal5_t1 graph

## Slide 29: How much H2 are we making?
Meents et al. (2010) claimed:
"400 molecules of H2 for every 12.5 keV photon“
= 3.3x10-7 mol/J
200 kGy to 100 um3 = 0.24 J
= 1.2x10-8 mol H2
100 um3 of water = 6.7x10-8 mol
→ 18% of water has reacted!

## Slide 30
The number of photons scattered
before crystal is dead
is independent
of flux & time
1 um3 = 105 photons (roomT)
1 um3 = 106 photons (synch)
1 um3 = 108 photons (XFEL)
Henderson, 1990; Gonzalez & Nave, 1994; Glaeser et al., 2000; Sliz et al., 2003; Leiros et al., 2006; Owen et al., 2006; Garman & McSweeney, 2006; Garman & Nave, 2009; Holton, 2009
(synch)

## Slide 31: Dose slicing
crystal’s useful life
N
photons
N
photons
N
photons
unacceptable
completeness
unacceptable
read noise

## Slide 32: Dose slicing: is that a spot?

## Slide 33: Dose slicing: is that a spot?

## Slide 34: Dose slicing: is that a spot?

## Slide 35: Dose slicing: is that a spot?

## Slide 36: Dose slicing: is that a spot?

## Slide 37: Dose slicing: is that a spot?

## Slide 38: Adding images
% cat << EOF > MERGE2CBF.INP
NAME_TEMPLATE_OF_DATA_FRAMES= /data/you/weak_?????.cbf
DATA_RANGE= 1 10
NAME_TEMPLATE_OF_OUTPUT_FRAMES=/data/you/sum_???.cbf
NUMBER_OF_DATA_FRAMES_COVERED_BY_EACH_OUTPUT_FRAME= 10
EOF
% merge2cbf
Which is better?:
% xds_runme.com /data/you/weak_?????.cbf
% xds_runme.com /data/you/sum_?????.cbf

## Slide 39: What if?
You only have a few small crystals…
Should you:
  Collect 360° from each?
  Collect 10° from each at 36x exposure?
  Glue 36 xtals together, then collect 360° ?
  Glue, and do 12960° faint exposures?

## Slide 40
Chart: 
Resolution (Å)
CC vs right answer
“true” resolution limit

## Slide 41
Dose slice collection scenario |  |  | 
fine | coarse | glue | Glue & fine
3.17700 | 3.14057 | 3.17604 | 3.12086
“true” resolution vs strategy
301,640,334 photons/xtal

## Slide 42
processing procedure | resolution limit |  |  | 
 | fine | coarse | glue | glue&fine
XDS/XSCALE | 3.17700 | 3.14057 | 3.17604 | 3.12086
XDS/aimless | 3.16251 | 3.15922 | 3.16675 | 3.12162
XDS/noscale | 3.27289 | 3.15617 | 3.38929 | 3.12162
MOSFLM-defaults | x | x | x | x
MOSFLM-cheat | 3.19755 | 3.18166 | 3.20354 | 3.19755
HKL2000-defaults | x | x | x | x
HKL2000-expert | 3.21803 | 3.12646 | 3.11551 | 3.21803
HKL2000-cheat | 3.16119 | 3.10338 | 3.10338 | 3.16300
DIALS-defaults | x | x | 3.42125 | x
DIALS-expert | 3.10296 | 3.09547 | 3.07653 | 3.08465
d*Trek | ? | ? | 4.34076 | ?
EVAL15 | 3.14352 | 3.10383 | 3.09936 | 3.14105
EVAL15-sadabs | 3.13920 | 3.10797 | 3.10171 | 3.13700

## Slide 43: Optimum exposure time?
~1 photon/pixel
  XDS, DIALS
~10 photon/pixel
HKL2000, MOSFLM
~30 photon/pixel
CCD detector
adjust exposure
so this is ~100

## Slide 44: Adding images
% cat << EOF > MERGE2CBF.INP
NAME_TEMPLATE_OF_DATA_FRAMES= /data/you/weak_?????.cbf
DATA_RANGE= 1 10
NAME_TEMPLATE_OF_OUTPUT_FRAMES=/data/you/sum_???.cbf
NUMBER_OF_DATA_FRAMES_COVERED_BY_EACH_OUTPUT_FRAME= 10
EOF
% merge2cbf
Which is better?:
% xds_runme.com /data/you/weak_?????.cbf
% xds_runme.com /data/you/sum_?????.cbf
“lossy”
compression

## Slide 45: But my reviewer says…
Rmerge in high-resolution bin is >100% !
Answer: this is expected

## Slide 46: Expected Rmerge as Iobs → 0

## Slide 47
Average value
Averaging Gaussian error
Number of samples averaged

## Slide 48
Average value
Averaging Gauss/Gauss error
Number of samples averaged

## Slide 49
Average outer-shell Rmerge
Rmerge at the resolution limit in PDB
Year

## Slide 50: Take-home lesson:
R factors are
undefined
as Iobs → 0
Report as “ – “ in outer bin

## Slide 51: Optimum resolution cutoff is:
Too optimistic: add nothing but noise
Too pessimistic: series-termination error
Happy medium?
Simulate:
  Random atoms, compute F2
  Add Gaussian noise, RMS = 1
  Truncate
  Subtract “right” map, RMS difference

## Slide 52
Chart: 
| 0.8 | 0.81 | 0.82 | 0.83 | 0.84 | 0.85 | 0.86 | 0.87 | 0.88 | 0.89 | 0.9 | 0.91 | 0.92 | 0.93 | 0.94 | 0.95 | 0.96 | 0.97 | 0.98 | 0.99 | 1 | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.0900000000000001 | 1.1000000000000001 | 1.1100000000000001 | 1.1200000000000001 | 1.1299999999999999 | 1.1399999999999999 | 1.1499999999999999 | 1.1599999999999999 | 1.17 | 1.18 | 1.19 | 1.2 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.29 | 1.3
total error | 0.0063854 | 0.006377 | 0.0063704 | 0.0063613 | 0.0063546 | 0.0063466 | 0.006339 | 0.0063331 | 0.0063281 | 0.0063218 | 0.0063169 | 0.0063122 | 0.0063076 | 0.0063035 | 0.0062986 | 0.0062952 | 0.006292 | 0.0062882 | 0.0062851 | 0.0062825 | 0.0062808 | 0.0062806 | 0.0062829 | 0.0062871 | 0.0063049 | 0.0063317 | 0.0063768 | 0.006453 | 0.00655 | 0.0067899 | 0.006935 | 0.0072772 | 0.0075335 | 0.0079267 | 0.0082267 | 0.0085917 | 0.0090308 | 0.0096403 | 0.010142 | 0.010578 | 0.011234 | 0.011798 | 0.012663 | 0.013292 | 0.013934 | 0.014784 | 0.015263 | 0.016206 | 0.01711 | 0.01817 | 0.018793
series termination | 0.0001167 | 0.00014118 | 0.00016693 | 0.00020738 | 0.00024052 | 0.00028835 | 0.00034393 | 0.00040287 | 0.00046367 | 0.00054497 | 0.00062312 | 0.00072092 | 0.00082139 | 0.00091787 | 0.0010631 | 0.0011963 | 0.0013389 | 0.0015127 | 0.0017075 | 0.0019087 | 0.0021401 | 0.0023446 | 0.0026241 | 0.0028107 | 0.0031805 | 0.0034765 | 0.0037849 | 0.0041158 | 0.0044212 | 0.0049665 | 0.0052371 | 0.0057965 | 0.0061838 | 0.0067416 | 0.0071474 | 0.0076234 | 0.0081753 | 0.0089132 | 0.0095026 | 0.010003 | 0.010739 | 0.011361 | 0.012297 | 0.012969 | 0.013647 | 0.014539 | 0.015036 | 0.016012 | 0.016942 | 0.018027 | 0.018662
noise in data | 0.0063844 | 0.0063755 | 0.0063682 | 0.0063579 | 0.0063501 | 0.0063401 | 0.0063297 | 0.0063203 | 0.0063111 | 0.0062983 | 0.0062861 | 0.0062709 | 0.0062539 | 0.0062363 | 0.0062082 | 0.0061804 | 0.0061479 | 0.0061034 | 0.0060487 | 0.0059854 | 0.0059049 | 0.0058264 | 0.0057085 | 0.0056237 | 0.0054437 | 0.0052917 | 0.0051318 | 0.0049697 | 0.0048324 | 0.0046297 | 0.0045459 | 0.0043996 | 0.0043026 | 0.0041691 | 0.0040733 | 0.0039623 | 0.0038365 | 0.0036725 | 0.0035436 | 0.0034402 | 0.0032968 | 0.0031835 | 0.0030235 | 0.002914 | 0.0028109 | 0.0026816 | 0.0026168 | 0.0024966 | 0.0023891 | 0.0022725 | 0.0022113
Resolution cutoff (Å)
Error in map (rms e-)
Optimal resolution cutoff

## Slide 53
Chart: 
| 0.8 | 0.81 | 0.82 | 0.83 | 0.84 | 0.85 | 0.86 | 0.87 | 0.88 | 0.89 | 0.9 | 0.91 | 0.92 | 0.93 | 0.94 | 0.95 | 0.96 | 0.97 | 0.98 | 0.99 | 1 | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.0900000000000001 | 1.1000000000000001 | 1.1100000000000001 | 1.1200000000000001 | 1.1299999999999999 | 1.1399999999999999 | 1.1499999999999999 | 1.1599999999999999 | 1.17 | 1.18 | 1.19 | 1.2 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.29 | 1.3
total error | 0.011234 | 0.011798 | 0.012663 | 0.013292 | 0.013934 | 0.014784 | 0.015263 | 0.016206 | 0.01711 | 0.01817 | 0.018793
CC1/2 | 0.00027543 | 0.0009316 | 0.0011407 | 0.00093939 | -0.00085547 | 8.519e-05 | 0.0017707 | 0.0018655 | 0.00044742 | -0.00083644 | -8.9753e-05 | -0.0012856 | -0.0022441 | -0.0011232 | -0.0006217 | -0.00063863 | 0.00039935 | 0.0020736 | 0.0023749 | 0.001991 | 0.0015254 | 0.0011874 | 0.0012966 | 0.0070442 | 0.010659 | 0.014306 | 0.021539 | 0.031659 | 0.044804 | 0.054608 | 0.085054 | 0.10512 | 0.15123 | 0.19219 | 0.24658 | 0.30221 | 0.374 | 0.43569 | 0.50225 | 0.58077 | 0.63948 | 0.71345 | 0.75573 | 0.79925 | 0.8353 | 0.86091 | 0.89448 | 0.91138 | 0.92903 | 0.94132 | 0.95425
CCtrue | -0.0018739 | -0.0009566 | 0.00078927 | -0.00047161 | 0.00086676 | 0.0013379 | 0.00021407 | 0.001888 | 0.0027656 | 0.0036656 | 0.0050632 | 0.0063693 | 0.0067999 | 0.010765 | 0.013871 | 0.01712 | 0.023754 | 0.030523 | 0.038896 | 0.047732 | 0.06188 | 0.078774 | 0.090571 | 0.11775 | 0.13935 | 0.1692 | 0.20286 | 0.24067 | 0.29172 | 0.32687 | 0.40014 | 0.43863 | 0.51218 | 0.56571 | 0.62646 | 0.68068 | 0.73761 | 0.77832 | 0.81873 | 0.8591 | 0.88409 | 0.9127 | 0.92782 | 0.94254 | 0.95429 | 0.96199 | 0.97176 | 0.97663 | 0.98146 | 0.98479 | 0.98825
Resolution cutoff (Å)
Correlation Coefficient
Optimal resolution cutoff

## Slide 54
Chart: 
| 0.8 | 0.81 | 0.82 | 0.83 | 0.84 | 0.85 | 0.86 | 0.87 | 0.88 | 0.89 | 0.9 | 0.91 | 0.92 | 0.93 | 0.94 | 0.95 | 0.96 | 0.97 | 0.98 | 0.99 | 1 | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.0900000000000001 | 1.1000000000000001 | 1.1100000000000001 | 1.1200000000000001 | 1.1299999999999999 | 1.1399999999999999 | 1.1499999999999999 | 1.1599999999999999 | 1.17 | 1.18 | 1.19 | 1.2 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.29 | 1.3
total error | 0.011234 | 0.011798 | 0.012663 | 0.013292 | 0.013934 | 0.014784 | 0.015263 | 0.016206 | 0.01711 | 0.01817 | 0.018793
CC1/2 | 0.00027543 | 0.0009316 | 0.0011407 | 0.00093939 | -0.00085547 | 8.519e-05 | 0.0017707 | 0.0018655 | 0.00044742 | -0.00083644 | -8.9753e-05 | -0.0012856 | -0.0022441 | -0.0011232 | -0.0006217 | -0.00063863 | 0.00039935 | 0.0020736 | 0.0023749 | 0.001991 | 0.0015254 | 0.0011874 | 0.0012966 | 0.0070442 | 0.010659 | 0.014306 | 0.021539 | 0.031659 | 0.044804 | 0.054608 | 0.085054 | 0.10512 | 0.15123 | 0.19219 | 0.24658 | 0.30221 | 0.374 | 0.43569 | 0.50225 | 0.58077 | 0.63948 | 0.71345 | 0.75573 | 0.79925 | 0.8353 | 0.86091 | 0.89448 | 0.91138 | 0.92903 | 0.94132 | 0.95425
CCtrue | -0.0018739 | -0.0009566 | 0.00078927 | -0.00047161 | 0.00086676 | 0.0013379 | 0.00021407 | 0.001888 | 0.0027656 | 0.0036656 | 0.0050632 | 0.0063693 | 0.0067999 | 0.010765 | 0.013871 | 0.01712 | 0.023754 | 0.030523 | 0.038896 | 0.047732 | 0.06188 | 0.078774 | 0.090571 | 0.11775 | 0.13935 | 0.1692 | 0.20286 | 0.24067 | 0.29172 | 0.32687 | 0.40014 | 0.43863 | 0.51218 | 0.56571 | 0.62646 | 0.68068 | 0.73761 | 0.77832 | 0.81873 | 0.8591 | 0.88409 | 0.9127 | 0.92782 | 0.94254 | 0.95429 | 0.96199 | 0.97176 | 0.97663 | 0.98146 | 0.98479 | 0.98825
CC* | 0.023467 | 0.043145 | 0.047737 | 0.043325 | 0 | 0.013052 | 0.059457 | 0.061026 | 0.029907 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.028256 | 0.064333 | 0.068837 | 0.063041 | 0.055192 | 0.048703 | 0.05089 | 0.11828 | 0.14523 | 0.16795 | 0.20535 | 0.24774 | 0.29286 | 0.32181 | 0.39595 | 0.43617 | 0.51257 | 0.56782 | 0.62898 | 0.68128 | 0.73783 | 0.77906 | 0.81772 | 0.8572 | 0.88323 | 0.91256 | 0.92783 | 0.94256 | 0.95408 | 0.9619 | 0.97175 | 0.97654 | 0.98143 | 0.98477 | 0.98823
Resolution cutoff (Å)
Correlation Coefficient
Optimal resolution cutoff

## Slide 55
log( intensity )
Wilson Plot
2.2 1.6 1.3 1.1 1.0 0.9 0.85 0.8 0.75 0.7
resolution (Å)
8
2

## Slide 56: Optimum resolution cutoff is:
0.0 Å

## Slide 57
Electron density (e-/Å)
optical definition of resolution
Position (Å)

## Slide 58
Electron density (e-/Å)
optical definition of resolution
Position (Å)

## Slide 59
Electron density (e-/Å)
optical definition of resolution
Position (Å)

## Slide 60
Electron density (e-/Å)
optical definition of resolution
Position (Å)

## Slide 61
Electron density (e-/Å)
optical definition of resolution
Position (Å)

## Slide 62
Electron density (e-/Å)
optical definition of resolution
Position (Å)

## Slide 63
Electron density (e-/Å)
optical definition of resolution
Position (Å)

## Slide 64
Electron density (e-/Å)
optical definition of resolution
Position (Å)

## Slide 65
Electron density (e-/Å)
optical definition of resolution
FWHM
0.849 FWHM
Position (Å)

## Slide 66
Electron density (e-/Å)
optical definition of resolution
map sharpening
Position (Å)

## Slide 67
Electron density (e-/Å)
optical definition of resolution
realistic
map sharpening
Position (Å)

## Slide 68
Electron density (e-/Å)
optical definition of resolution
Position (Å)
84%
98%
99.9%

## Slide 69
Atomic Resolution (Å)
“3σ resolution” of model from map
refined atomic B factor (8πux2)
RMS map error (e-/Å3)

## Slide 70: B factors are in the PDB file
ATOM 122 N LEU A 13 -3.244 25.808 19.998 1.00 16.96 N
ATOM 123 CA LEU A 13 -2.877 25.448 21.355 1.00 15.29 C
ATOM 124 C LEU A 13 -2.792 23.966 21.561 1.00 17.54 C
ATOM 125 O LEU A 13 -1.814 23.493 22.143 1.00 16.35 O
ATOM 126 CB LEU A 13 -3.907 26.164 22.268 1.00 18.72 C
ATOM 127 CG LEU A 13 -3.577 25.982 23.738 1.00 21.19 C
ATOM 128 CD1 LEU A 13 -2.283 26.820 24.019 1.00 19.43 C
ATOM 129 CD2 LEU A 13 -4.702 26.474 24.639 1.00 24.65 C
ATOM 130 N SER A 14 -3.677 23.149 20.979 1.00 15.96 N
ATOM 131 CA SER A 14 -3.646 21.711 21.061 1.00 18.26 C
ATOM 132 C SER A 14 -2.373 21.203 20.360 1.00 18.71 C
ATOM 133 O SER A 14 -1.747 20.315 20.930 1.00 17.47 O
ATOM 134 CB SER A 14 -4.875 21.077 20.419 1.00 17.62 C
ATOM 135 OG ASER A 14 -4.825 19.665 20.388 0.50 20.89 O
ATOM 136 OG BSER A 14 -6.027 21.408 21.164 0.50 18.67 O
ATOM 137 N LYS A 15 -2.045 21.772 19.215 1.00 18.03 N
ATOM 138 CA LYS A 15 -0.799 21.361 18.555 1.00 18.12 C
ATOM 139 C LYS A 15 0.446 21.707 19.351 1.00 18.81 C
ATOM 140 O LYS A 15 1.400 20.948 19.411 1.00 17.77 O
ATOM 141 CB LYS A 15 -0.700 22.034 17.177 1.00 14.49 C
ATOM 142 CG LYS A 15 -1.727 21.368 16.256 1.00 16.12 C
ATOM 143 CD LYS A 15 -1.663 22.147 14.936 1.00 19.40 C
ATOM 144 CE ALYS A 15 -2.725 21.614 13.986 0.50 17.42 C
ATOM 145 CE BLYS A 15 -1.750 21.211 13.750 0.50 17.01 C
ATOM 146 NZ ALYS A 15 -2.346 21.674 12.559 0.50 18.61 N
ATOM 147 NZ BLYS A 15 -3.052 20.513 13.741 0.50 18.76 N

## Slide 71
Number of columns, rows, sections ............... 160 132 264
Map mode ........................................ 2
Start and stop points on columns, rows, sections 0 159 0 131 0 263
Grid sampling on x, y, z ........................ 132 160 264
Cell dimensions ................................. 43.0000 52.6100 89.1200 90.0000 90.0000 90.0000
Fast, medium, slow axes ......................... Y X Z
Minimum density ................................. -0.09104
Maximum density ................................. 0.35006
Mean density .................................... 0.00000
Rms deviation from mean density ................. 0.01954
Space-group ..................................... 19
Number of titles ................................ 1
RMS map error
is in the FFT log!
echo labin F1=FOFCWT PHI=PHFOFCWT | fft hklin refine_out.mtz

## Slide 72: What is my “resolution” ?
Where do I cut my data?
What do I give to refinement?
What can I claim in my paper?
CC1/2→ 0
Rmerge meaningless as Iobs → 0
Karplus-Diederichs paired refinement
report “R1”
“Straighten” Wilson Plot for maps?
Retain historical relevance? I/σ = 3
Report “3σ resolution” for model?
σ(F) only matters for phasing

## Slide 73: How low can you dose?
Stay below World Record
Look up dose rate
  http://bl831.als.lbl.gov/damage_rates.pdf
Know your limits
  10 MGy/Å reso, 5 = SeMet, 20 kGy?
Can’t outrun damage
Long exposures → loose info
Resolution: optimum cutoff is 0 Å
R-factor is undefined at reso limit
http://bl831.als.lbl.gov/~jamesh/
powerpoint/MRPI_raddam_2017.pptx