## Slide 1: Anything is possible … with the right tools.

## Slide 2: Advanced Light Source
200 m circ. | 12 sectors
200 staff | 44 beamlines

## Slide 3: The truth about x-ray beams
Term units significance
Flux photons/s duration of experiment
Beam Size μm match to sample
Divergence mrad spot size vs distance
Wavelength Å resolution and absorption
Dispersion Δλ/λ spot size
Flux density ph/s/area scattering/damage rate
Fluence ph/area scattering/damage

## Slide 4: The truth about x-ray beams
quantity | units | home source | ALS 8.3.1
flux | Photons/second | 2.5 x 109 | 3 x 1011
exposure | seconds | 120 | 1

## Slide 5: The truth about x-ray beams
quantity | units | home source | ALS 8.3.1
flux | Photons/second | 2.5 x 109 | 3 x 1011
exposure | seconds | 120 | 1
Dispersion | wavelength range / wavelength | 0.2% (Kα1 - Kα2) | 0.014% (Si111)
Divergence | milliRadian | 4.8 | 2 (h) 0.3 (v)
Beam size | microns | 100 | 60 (h) 100 (v)
Spectral brightness | Photons/s/mm2/ mR2/0.1%BW | 5.4 x 109 | 6 x 1014

## Slide 6: The truth about x-ray beams
quantity | units | home source | ALS 8.3.1
flux | Photons/second | 1.8 x 108 | 3 x 1011
exposure | seconds | 28 min | 1 second
Dispersion | wavelength range / wavelength | 0.014% (Si111) | 0.014% (Si111)
Divergence | milliRadian | 4.8 | 2 (h) 0.3 (v)
Beam size | microns | 100 | 60 (h) 100 (v)
Spectral brightness | Photons/s/mm2/ mR2/0.1%BW | 5.4 x 109 | 6 x 1014

## Slide 7: The truth about x-ray beams
quantity | units | home source | ALS 8.3.1
flux | Photons/second | 3 x 107 | 3 x 1011
exposure | seconds | 2.8 hours | 1 second
Dispersion | wavelength range / wavelength | 0.014% (Si111) | 0.014% (Si111)
Divergence | milliRadian | 0.8 | 2 (h) 0.3 (v)
Beam size | microns | 100 | 60 (h) 100 (v)
Spectral brightness | Photons/s/mm2/ mR2/0.1%BW | 5.4 x 109 | 6 x 1014

## Slide 8: The truth about x-ray beams
quantity | units | home source | ALS 8.3.1
flux | Photons/second | 2 x 107 | 3 x 1011
exposure | seconds | 4 hours | 1 second
Dispersion | wavelength range / wavelength | 0.014% (Si111) | 0.014% (Si111)
Divergence | milliRadian | 0.8 | 2 (h) 0.3 (v)
Beam size | microns | 77 | 60 (h) 100 (v)
Spectral brightness | Photons/s/mm2/ mR2/0.1%BW | 5.4 x 109 | 6 x 1014

## Slide 9: The truth about x-ray beams
quantity | units | home source | ALS 8.3.1
flux | Photons/second | 2 x 107 | 3 x 1011
exposure | time | 13 years | 1 hour
Dispersion | wavelength range / wavelength | 0.014% (Si111) | 0.014% (Si111)
Divergence | milliRadian | 0.8 | 2 (h) 0.3 (v)
Beam size | microns | 77 | 60 (h) 100 (v)
Spectral brightness | Photons/s/mm2/ mR2/0.1%BW | 5.4 x 109 | 6 x 1014

## Slide 10: you must choose, but choose wisely!
Caltech
SSRL
Yale
Oxford
SPring-8
SPINE
MiTeGen RT
Syrrx
H. magnetic
H. hybrid
H. copper
H. screw cap
H. colored
Hampton screw cap
Hampton magnetic

## Slide 11: Touch screen

## Slide 12: Zero-parallax microscope
pinhole
mirror
microscope
Styrofoam™ backlight
backstop
Notes: Next slide please

## Slide 13: Zero-parallax microscope
pinhole
mirror
microscope
Styrofoam™ backlight
backstop
Notes: Inspired by the ESRF microcrystal camera, the superbend endstations are all equipped with zero-paralax optics. The user sees their sample from the point of view of the x-ray beam. Sample alignment is not only more stable but more intuitive than an off-axis camera. Placing the crystal into the microscope crosshairs gaurantees that it will be hit by the beam.

## Slide 14: Zero-parallax microscope
pinhole
mirror
microscope
Styrofoam™ backlight
backstop
Notes: Inspired by the ESRF microcrystal camera, the superbend endstations are all equipped with zero-paralax optics. The user sees their sample from the point of view of the x-ray beam. Sample alignment is not only more stable but more intuitive than an off-axis camera. Placing the crystal into the microscope crosshairs gaurantees that it will be hit by the beam.

## Slide 17: Zero-parallax microscope
pinhole
mirror
microscope
Styrofoam™ backlight
backstop
Notes: Inspired by the ESRF microcrystal camera, the superbend endstations are all equipped with zero-paralax optics. The user sees their sample from the point of view of the x-ray beam. Sample alignment is not only more stable but more intuitive than an off-axis camera. Placing the crystal into the microscope crosshairs gaurantees that it will be hit by the beam.

## Slide 18: Zero-parallax microscope
pinhole
mirror
microscope
Styrofoam™ backlight
backstop
Notes: Inspired by the ESRF microcrystal camera, the superbend endstations are all equipped with zero-paralax optics. The user sees their sample from the point of view of the x-ray beam. Sample alignment is not only more stable but more intuitive than an off-axis camera. Placing the crystal into the microscope crosshairs gaurantees that it will be hit by the beam.

## Slide 19: Semi-transparent beam stop
Phosphor attenuator
gold
gold
ZnS

## Slide 20: Semi-transparent beam stop
Phosphor attenuator
gold
gold
ZnS

## Slide 21: Semi-transparent beam stop
Phosphor attenuator
gold
gold
ZnS

## Slide 22: Traffic light

## Slide 23: Beamline pager

## Slide 24: Foam dewars

## Slide 25: Foam dewars

## Slide 26: Ice-B-gone

## Slide 27: X-ray rules of thumb: absorption and background scattering
1 μm crystal ≈ 1 μm water
≈ 1 μm plastic
≈ 0.1 μm glass
≈ 1000 μm air

## Slide 28: What we really need
GRAPHENE
Thin
Light
Tight
Cheap

## Slide 29: Crystallography
a tool for structural biology

## Slide 30: Crystallography
a tool for structural biology

## Slide 31: Geography

## Slide 32: X-rays “see” electrons

## Slide 33: How big is an atom?
C
1 Ångstrom (Å)
1 nanometer (10-9 m)
N
O
S
H
U

## Slide 34: How big is an atom?
U
C
N
O
S
H
as seen by X-rays

## Slide 35: Where do photons go?
beamstop
Transmitted (98%)
Protein
1A x-rays

## Slide 36: Elastic scattering

## Slide 37: Elastic scattering

## Slide 38: Where do photons go?
beamstop
elastic scattering (6%)
Transmitted (98%)
Protein
1A x-rays

## Slide 39: Elastic scattering

## Slide 40: Inelastic scattering

## Slide 41: Where do photons go?
beamstop
elastic scattering (6%)
Transmitted (98%)
Protein
1A x-rays

## Slide 42: Where do photons go?
beamstop
elastic scattering (6%)
Transmitted (98%)
inelastic scattering (7%)
Protein
1A x-rays

## Slide 43: Where do photons go?
beamstop
elastic scattering (6%)
Transmitted (98%)
inelastic scattering (7%)
Protein
1A x-rays
Re-emitted (99%)
Absorbed (~0%)

## Slide 44: Photoelectric absorption

## Slide 45: Photoelectric absorption
e-
+

## Slide 46: Fluorescence
+

## Slide 47: Fluorescence
e-
+

## Slide 48: Metal identification

## Slide 49: Where do photons go?
beamstop
elastic scattering (6%)
Transmitted (98%)
inelastic scattering (7%)
Protein
1A x-rays
Re-emitted (99%)
Absorbed (~0%)

## Slide 50: Where do photons go?
beamstop
elastic scattering (6%)
Transmitted (98%)
inelastic scattering (7%)
Photoelectric (87%)
Protein
1A x-rays
Re-emitted (99%)
Absorbed (~0%)

## Slide 51: Where do photons go?
beamstop
elastic scattering (6%)
Transmitted (98%)
inelastic scattering (7%)
Photoelectric (87%)
Protein
1A x-rays
Re-emitted (~0%)
Absorbed (99%)
Re-emitted (99%)
Absorbed (~0%)

## Slide 52: Where do photons go?
beamstop
elastic scattering (6%)
Transmitted (98%)
useful/absorbed energy: 7.3%
inelastic scattering (7%)
Photoelectric (87%)
Protein
1A x-rays
Re-emitted (~0%)
Absorbed (99%)
Re-emitted (99%)
Absorbed (~0%)

## Slide 53
detector
sample
detector
x-ray beam
scattering

## Slide 54: nearBragg program
http://bl831.als.lbl.gov/~jamesh/nearBragg/
“assumption-free” total scattering
no Fourier Transform
no unit cells
no “mosaicity”
arbitrary “atoms”
arbitrary “source”
coherent or not

## Slide 56
sample
detector
x-ray beam
scattering

## Slide 57
sample
detector
x-ray beam
scattering

## Slide 58
sample
detector
x-ray beam
scattering

## Slide 59
to source
to detector
d
d∙sin(θ)
θ
atom #1
atom #2
Bragg’s Law
nλ = 2d sin(θ)

## Slide 60
to source
to detector
d
d∙sin(θ)
θ
atom #1
atom #2
Bragg’s Law
nλ = 2d sin(θ)

## Slide 62
scattering from a structure
sample
detector

## Slide 63
forward Fourier Transform
inverse Fourier Transform
no phase

## Slide 64
scattering from a structure
colored by phase
sample
detector
inverse Fourier Transform

## Slide 65
scattering from a lattice
colored by phase
sample
detector

## Slide 66
scattering from a lattice
colored by phase
sample
detector

## Slide 67
scattering from a lattice
colored by phase
sample
detector

## Slide 68
scattering from a crystal structure
colored by phase
sample
detector

## Slide 69: Promise of Single-molecule Imaging
R. Neutze et al., Nature 406, pp. 752-757 (2000)

## Slide 70: lysozyme: real and reciprocal

## Slide 71: lysozyme: thermal motion

## Slide 72: Snapshot from single virus particle
TEM
2 keV LCLS 200 fs Mimi virus single-shot.
200 nm
Reconstructed image
Resolution 20nm
Seibert, et al. (2011). Nature 470, 78-81.
Notes: what limits single-particle resolution ?
get 100 times more photons/shot if correctly focussed beam
get 100 times more with submicron beam focus on CXI chamber
Problem – scattering goes from N^2 to N (Wilson statistics).
So use a-priori information – modelling Axel Brummer
Make beam smaller than jet to elliminate streak
72

## Slide 73: Single-molecule imaging result
200 nm
Resolution ~2 mm
2.5 eV 0.1 s Michelin P255/70R18 LTX A/S 112T BSW
to what
resolution
is this
interesting?
Notes: what limits single-particle resolution ?
get 100 times more photons/shot if correctly focussed beam
get 100 times more with submicron beam focus on CXI chamber
Problem – scattering goes from N^2 to N (Wilson statistics).
So use a-priori information – modelling Axel Brummer
Make beam smaller than jet to elliminate streak
73

## Slide 74: Muybridge’s galloping horse (1878)

## Slide 75: Muybridge’s multi-camera

## Slide 76: not enough signal

## Slide 77: brighter light

## Slide 78: even brighter

## Slide 79: very bright light

## Slide 80: a “crystal” of horses

## Slide 81: realistic “crystal” of horses

## Slide 82: average structure: galloping horse

## Slide 83: Muybridge’s galloping horse (1878)
Real space
reciprocal
“Time-resolved” diffraction

## Slide 84: Muybridge’s galloping horse (1878)
Real space
reciprocal
Diffract and Destroy

## Slide 85: Muybridge’s galloping horse (1878)
Real space
reciprocal
Average electron density

## Slide 86: Supporting a model with data

## Slide 87: That’ll be in 20 years What can we do now?

## Slide 88: Molecular Dynamics Simulation
1aho
Scorpion toxin
0.96 Å resolution
64 residues
Solvent: H20 + acetate
Cerutti et al. (2010).J. Phys. Chem. B 114, 12811-12824.
using real
crystal’s lattice

## Slide 89: 30 conformers from 24,000

## Slide 90: Electron density from 24,000 conformers

## Slide 91: Electron density from 24,000 conformers

## Slide 92: Molecular Dynamics vs Observation
Fsim
Fcalc
Rcryst= 0.137
refined_vs_Fsim.pdb
1aho 64-residue scorpion toxin in water to 1.0 Å resolution

## Slide 93: Molecular Dynamics vs Observation
Fobs
1aho.cif
1aho.pdb
Fsim
Fcalc
Fcalc
Rcryst= 0.137
Rcryst= 0.116
Rfree = 0.69
refined_vs_Fsim.pdb
LSQ rmsd = 0.43Å
rmsd = 1.05 Å
1aho 64-residue scorpion toxin in water to 1.0 Å resolution
Rfree = 0.48 to 4 Å
Rfree =

## Slide 94: Molecular Dynamics vs Observation
RMSD
1.05 Å

## Slide 95: Molecular Dynamics vs Observation
RMSD
0.45 Å
aligned

## Slide 96: Molecular Dynamics vs Observation
Sodium acetate trihydrate
5 atoms
3 waters
0.9 Å resolution
C2/c
Rvault = 0.03

## Slide 97: Molecular Dynamics vs Observation
fav8
8 residues
4 waters
1.0 Å resolution
P1

## Slide 98: Molecular Dynamics vs Observation
Fobs
fav8.fcf
fav8.cif
Fsim
Fcalc
Fcalc
Rcryst= 0.13
Rcryst= 0.16
Rfree= 0.23
refined_vs_Fsim.pdb
LSQ rmsd = 0.091Å & 0.15 Å
rmsd = 0.20 Å
“fav8” 8-residue aromatic peptide with 4 waters to 1.0 Å resolution
R1 = 0.041
With 4σ cutoff!
Rfree = 0.20 to 2Å

## Slide 99: MD-predicted water structure
Simulated
density
Janowski et al (2013) JACS 135, 7938-7948

## Slide 100: MD-predicted water structure
Simulated
density
2mFo-DFc
Janowski et al (2013) JACS 135, 7938-7948

## Slide 101: MD-predicted Val rotamer
Janowski et al (2013) JACS 135, 7938-7948

## Slide 102: MD-predicted Val rotamer
confs | Rwork | Rfree
one | 4.11% | 5.84%
A & B | 3.89% | 5.53%
Janowski et al (2013) JACS 135, 7938-7948

## Slide 103: Molecular Dynamics Simulation
1aho
Scorpion toxin
0.96 Å resolution
64 residues
Solvent: H20 + acetate
Cerutti et al. (2010).J. Phys. Chem. B 114, 12811-12824.
using real
crystal’s lattice

## Slide 104: MLFSOM simulation

## Slide 105: Simulated diffraction image MLFSOM

## Slide 106: Simulated diffraction image MLFSOM
simulated
real

## Slide 107: Simulated diffraction image (2008)
simulated
real

## Slide 108
Simulated | statistic | Real
5.4% | Rmerge | 6.2%
18.9 | I/sd | 16.2
1.9 | I/sd (1.4 Ǻ) | 1.6
1.7 4.0 0.020 | SDCORR | 1.0 2.2 0.065
36.8 | PADFPH | 31.46
3.871 | mlphare f” | 3.476
0.178 | FOM | 0.192
0.545 | FOMDM | 0.664
0.6270 | CC(1H87) | 0.6090
0.170 | Rcryst | 0.184
0.205 | Rfree | 0.231

## Slide 109: Electron density from 24,000 conformers

## Slide 110: 2Fsim-Fcalc and Fsim-Fcalc maps
1.0 Å data
Rcryst= 0.137
Rfree = 0.159
1 “sigma” 2Fsim-Fc
2.5 “sigma” Fsim-Fc

## Slide 111: 2Fsim-Fcalc and (Fsim Φsim) - (Fcalc Φcalc) maps
1 “sigma” 2Fsim-Fc
Fright – (2Fsim-Fc)
1.0 Å data
Rcryst= 0.137
Rfree = 0.159

## Slide 112: Regular model with real data!
1 “sigma” 2Fsim-Fc
2.5 “sigma” Fsim-Fc
1.0 Å data
Rcryst= 0.116
Rfree = 0.135

## Slide 113: How many conformers are there?

## Slide 114: How many conformers are there?
40 explain density to 3% error

## Slide 115: How many conformers are there?
8 explain density to 3% error

## Slide 116: The Future of Structural Biology
Will be noisy
Averaging destroys resolution
Multi-crystal data – isomorphism
structures of the future must be 4D
There is “life” below 1 σ

## Slide 117: PowerPoint File available:
http://bl831.als.lbl.gov/
~jamesh/powerpoint/
NanoGM_Xray_2014.pptx

## Slide 118: Acknowledgements
8.3.1 PRT: Tom Alber
National Institutes of Health (STTR)
Center for Structure of Membrane Proteins
Center for HIV Accessory and Regulatory Complexes
Membrane Protein Expression Center II (MPEC)
Center for HIV Accessory and Regulatory Complexes (HARC)
W. M. Keck Foundation
Plexxikon, Inc.
M D Anderson CRC
Alberta Synchrotron Institute
University of California Berkeley
University of California San Francisco
National Science Foundation
University of California Campus-Laboratory Collaboration Grant
Henry Wheeler
The Advanced Light Source is supported by the Director, Office of Science, Office of Basic Energy Sciences, Materials Sciences Division, of the US Department of Energy under contract No. DEAC02- 05CH11231 at Lawrence Berkeley National Laboratory.

## Slide 119: Major Phasing techniques
Molecular Replacement
Multiple Isomorphous Replacement
Multiwavelength Anomalous Diffraction
Single-wavelength Anomalous Diffraction

## Slide 120
inverse Fourier Transform
no phase

## Slide 121
inverse Fourier Transform
no phase

## Slide 122: Major Phasing techniques
Molecular Replacement
Multiple Isomorphous Replacement
Multiwavelength Anomalous Diffraction
Single-wavelength Anomalous Diffraction

## Slide 123
sample
x-ray beam
anomalous scattering
detector

## Slide 124
sample
detector
x-ray beam
anomalous scattering

## Slide 125
sample
x-ray beam
anomalous scattering
detector

## Slide 126
sample
detector
x-ray beam
anomalous scattering

## Slide 127
sample
detector
x-ray beam
anomalous scattering

## Slide 128: Major Phasing techniques
Molecular Replacement
Multiple Isomorphous Replacement
Multiwavelength Anomalous Diffraction
Single-wavelength Anomalous Diffraction

## Slide 129: Molecular Replacement
?
correct structure and intensities
http://www.ysbl.york.ac.uk/~cowtan/fourier/coeff.html

## Slide 130: Molecular Replacement
use something similar as a starting model

## Slide 131: Model Building
current model is missing something

## Slide 132: Model Building
phases from model

## Slide 133: Model Building
missing bits show up in “difference map”

## Slide 134: Model Building
missing bits show up better in FO + (FO - FC) map

## Slide 135: Resolution
http://bl831.als.lbl.gov/~jamesh/movies/resolution.mpeg

## Slide 136: Resolution: low-angle cutoff
http://bl831.als.lbl.gov/~jamesh/movies/lo_cut.mpeg

## Slide 137: R-factor
http://bl831.als.lbl.gov/~jamesh/movies/rfactor.mpeg

## Slide 138: Figure of Merit
http://bl831.als.lbl.gov/~jamesh/movies/dephase.mpeg

## Slide 139: Overloads
http://bl831.als.lbl.gov/~jamesh/movies/overloads.mpeg

## Slide 140: Completeness: missing wedge
http://bl831.als.lbl.gov/~jamesh/movies/osc.mpeg

## Slide 141: Completeness: random deletion
http://bl831.als.lbl.gov/~jamesh/movies/completeness.mpeg

## Slide 142: Model Refinement

## Slide 143: Accelerator-based light source terminology
Code name Translation
First Generation Bending Magnet
Second Generation Wiggler
Third Generation Undulator
Fourth Generation Free electron laser

## Slide 144: Electric charge
balloon
Wool Sweater
-
-
-
-
-
-
+ + + + + +

## Slide 145: Electric field of a moving charge

## Slide 146: Electric field of a moving charge
accelerating charges
make
electromagnetic waves
(light)

## Slide 147: Bending Magnet
N
electron beam
X-rays !
S

## Slide 148: ALS 8.3.1 optics
ADSC Quantum 315r
Superbend
Plane
Parabolic
mirror
Torroidal
mirror
Si(111)
monochromator
Protein Crystal
pinhole
Scatter
guard
divergence
slits
MacDowell et. al. (2004) J. Sync. Rad. 11 447-455
Notes: Next slide please

## Slide 149: ALS 8.3.1 optics
ADSC Quantum 315r
Superbend
Plane
Parabolic
mirror
Torroidal
mirror
Si(111)
monochromator
Protein Crystal
pinhole
Scatter
guard
divergence
slits
MacDowell et. al. (2004) J. Sync. Rad. 11 447-455
Notes: Most notable acheivements are high brightness with very little “wasted” light. The discovery made by Howard Padmore? At ALS showed that a 2:1 demagnification cancels nearly all of the spherical abberations in the parabolic/torroid system. 30-50% of the total x-ray intensity passes through the 100um pinhole. The flux through the 100um pinhole at the superbend beamlines is comparable to that of the 5.0.2 wiggler or an APS bending magnet, but only 150 Watts of power must be dissipated by the beamline optics. In contrast, wiggler beamlines must disipate kilowatts. The optics here are much simpler and not exotic in any way. The system is stable and robust. There is never a need to re-focus the beam at a new wavelength. Only a 2-minute automatic tuneup is recommended for very large energy changes (more than 1-2 keV).

## Slide 150: X-ray optics: mirrors
Notes: Next slide please

## Slide 151: Why are the mirrors so long?
3 mm
Must be < 0.2°
1000 mm
Notes: Next slide please

## Slide 152: Monochromator: silicon crystals
Notes: Next slide please

## Slide 153: Monochromators
λ = 1 Å
Si
Notes: Next slide please

## Slide 154: Monochromators
λ = 1 Å
Si
d = 3.135 Å
θ = 9.177°
λ = 1 Å
nλ = 2d sin(θ)
Notes: Next slide please

## Slide 155: Monochromators
nλ = 2d sin(θ)
λ = 1.1 Å
Si
d = 3.135 Å
θ = 9.177°
Notes: Next slide please

## Slide 156: Monochromators
nλ = 2d sin(θ)
λ = 1.1 Å
Si
d = 3.135 Å
θ = 10.104°
λ = 1.1 Å
Notes: Next slide please

## Slide 157: Monochromators
nλ = 2d sin(θ)
λ = 1 Å
Si
d = 3.135 Å
θ = 9.177°
λ = 1 Å
Notes: Next slide please

## Slide 158: Monochromators
Si
λ = 1 Å
λ = 1 Å
Si
d = 3.135 Å
Double Crystal Monochromator
Notes: Next slide please

## Slide 159: Monochromators
Si
λ = 1.1 Å
λ = 1.1 Å
Si
d = 3.135 Å
Double Crystal Monochromator
Notes: Next slide please

## Slide 160: ALS 4.2.2 optics
Westbrook Noir 1
Superbend
Plane
Parabolic
mirror
Plane
parabolic
mirror
sagital
monochromator
Protein Crystal
Sample
slits
divergence
slits
Notes: Most notable acheivements are high brightness with very little “wasted” light. The discovery made by Howard Padmore? At ALS showed that a 2:1 demagnification cancels nearly all of the spherical abberations in the parabolic/torroid system. 30-50% of the total x-ray intensity passes through the 100um pinhole. The flux through the 100um pinhole at the superbend beamlines is comparable to that of the 5.0.2 wiggler or an APS bending magnet, but only 150 Watts of power must be dissipated by the beamline optics. In contrast, wiggler beamlines must disipate kilowatts. The optics here are much simpler and not exotic in any way. The system is stable and robust. There is never a need to re-focus the beam at a new wavelength. Only a 2-minute automatic tuneup is recommended for very large energy changes (more than 1-2 keV).

## Slide 161: Monochromators
Si
λ = 1.1 Å
λ = 1.1 Å
Si
d = 3.135 Å
Double Crystal Monochromator
Notes: Next slide please

## Slide 162: Monochromators
nλ = 2d sin(θ)
λ = 1.54 Å
θ = 1.84°
λ = 1.54 Å
Mo/C
d = 24 Å
Multilayer optics
Notes: Next slide please

## Slide 163: ALS 12.3.1 optics
ADSC Quantum 315r
Superbend
Plane
Parabolic
mirror
Torroidal
mirror
Si(111)
monochromator
Protein Crystal
pinhole
Scatter
guard
divergence
slits
2 x 107 ph/s/μm2 ≈ 3000 s/xtal
Notes: Most notable acheivements are high brightness with very little “wasted” light. The discovery made by Howard Padmore? At ALS showed that a 2:1 demagnification cancels nearly all of the spherical abberations in the parabolic/torroid system. 30-50% of the total x-ray intensity passes through the 100um pinhole. The flux through the 100um pinhole at the superbend beamlines is comparable to that of the 5.0.2 wiggler or an APS bending magnet, but only 150 Watts of power must be dissipated by the beamline optics. In contrast, wiggler beamlines must disipate kilowatts. The optics here are much simpler and not exotic in any way. The system is stable and robust. There is never a need to re-focus the beam at a new wavelength. Only a 2-minute automatic tuneup is recommended for very large energy changes (more than 1-2 keV).

## Slide 164: ALS 12.3.1 optics
ADSC Quantum 315r
Superbend
Plane
Parabolic
mirror
Torroidal
mirror
1% ML
monochromator
Protein Crystal
pinhole
Scatter
guard
divergence
slits
4 x 108 ph/s/μm2 ≈ 160 s/xtal
Notes: Most notable acheivements are high brightness with very little “wasted” light. The discovery made by Howard Padmore? At ALS showed that a 2:1 demagnification cancels nearly all of the spherical abberations in the parabolic/torroid system. 30-50% of the total x-ray intensity passes through the 100um pinhole. The flux through the 100um pinhole at the superbend beamlines is comparable to that of the 5.0.2 wiggler or an APS bending magnet, but only 150 Watts of power must be dissipated by the beamline optics. In contrast, wiggler beamlines must disipate kilowatts. The optics here are much simpler and not exotic in any way. The system is stable and robust. There is never a need to re-focus the beam at a new wavelength. Only a 2-minute automatic tuneup is recommended for very large energy changes (more than 1-2 keV).

## Slide 165: ALS 12.3.1 double endstation

## Slide 166: Two Bending Magnets
N
electron beam
S
S
N
2x X-rays !

## Slide 167: Wiggler
N
S
N
S
N
S
N
S
S
N
S
N
S
N
S
N
8x X-rays !

## Slide 168: ALS 5.0.2 optics
ADSC Quantum 315
16-pole wiggler
Plane
Parabolic
mirror
Torroidal
mirror
Si(111)
monochromator
Protein Crystal
pinhole
Scatter
guard
divergence
slits
1.5 x 108 ph/s/μm2 ≈ 400 s/xtal
Notes: Most notable acheivements are high brightness with very little “wasted” light. The discovery made by Howard Padmore? At ALS showed that a 2:1 demagnification cancels nearly all of the spherical abberations in the parabolic/torroid system. 30-50% of the total x-ray intensity passes through the 100um pinhole. The flux through the 100um pinhole at the superbend beamlines is comparable to that of the 5.0.2 wiggler or an APS bending magnet, but only 150 Watts of power must be dissipated by the beamline optics. In contrast, wiggler beamlines must disipate kilowatts. The optics here are much simpler and not exotic in any way. The system is stable and robust. There is never a need to re-focus the beam at a new wavelength. Only a 2-minute automatic tuneup is recommended for very large energy changes (more than 1-2 keV).

## Slide 169: Undulator
N
S
N
S
N
S
N
S
S
N
S
N
S
N
S
N
> 8x X-rays !

## Slide 170: Undulator emission spectrum

## Slide 171: LCLS undulator hall

## Slide 172: LCLS at SLAC

## Slide 173: Electric field of a moving charge

## Slide 174: SASE effect

## Slide 175: Self-Amplified Stimulated Emission (SASE) effect

## Slide 176: 100 fs damage “threshold”
Lomb et al. (2011) Phys Rev B 84, 214111

## Slide 177: Snapshot from single virus particle
TEM
2 keV LCLS 200 fs Mimi virus single-shot.
200 nm
Reconstructed image
Resolution 20nm
Seibert, et al. (2011). Nature 470, 78-81.
Notes: what limits single-particle resolution ?
get 100 times more photons/shot if correctly focussed beam
get 100 times more with submicron beam focus on CXI chamber
Problem – scattering goes from N^2 to N (Wilson statistics).
So use a-priori information – modelling Axel Brummer
Make beam smaller than jet to elliminate streak
177

## Slide 178: Single-molecule imaging result
200 nm
Resolution ~2 mm
2.5 eV 0.1 s Michelin P255/70R18 LTX A/S 112T BSW
to what
resolution
is this
interesting?
Notes: what limits single-particle resolution ?
get 100 times more photons/shot if correctly focussed beam
get 100 times more with submicron beam focus on CXI chamber
Problem – scattering goes from N^2 to N (Wilson statistics).
So use a-priori information – modelling Axel Brummer
Make beam smaller than jet to elliminate streak
178

## Slide 179: Accelerator-based light source terminology
Code name Translation
First Generation Bending Magnet
Second Generation Wiggler
Third Generation Undulator
Fourth Generation Free electron laser

## Slide 180: lysozyme: real and reciprocal

## Slide 181: lysozyme: thermal motion

## Slide 182: ALS 8.3.1 optics
ADSC Quantum 315r
Protein Crystal
pinhole
Scatter
guard
MacDowell et. al. (2004) J. Sync. Rad. 11 447-455
X-ray
Source
Notes: Most notable acheivements are high brightness with very little “wasted” light. The discovery made by Howard Padmore? At ALS showed that a 2:1 demagnification cancels nearly all of the spherical abberations in the parabolic/torroid system. 30-50% of the total x-ray intensity passes through the 100um pinhole. The flux through the 100um pinhole at the superbend beamlines is comparable to that of the 5.0.2 wiggler or an APS bending magnet, but only 150 Watts of power must be dissipated by the beamline optics. In contrast, wiggler beamlines must disipate kilowatts. The optics here are much simpler and not exotic in any way. The system is stable and robust. There is never a need to re-focus the beam at a new wavelength. Only a 2-minute automatic tuneup is recommended for very large energy changes (more than 1-2 keV).

## Slide 184
0. Purify
protein
3. Calculate
e- density
2. Collect
data
1. Grow
crystals
4. Build
model
5. Refine
models
How do I get one of those maps?

## Slide 185: X-ray data are 3D!

## Slide 186: How do I get one of those maps?
Elves examine images and
set-up data processing
Elves run…
mosflm
scala
solve
mlphare
dm
arp/warp
Holton, J. & Alber, T. (2004) PNAS USA 101 1537-42.

## Slide 187: Auger emission
+

## Slide 188: Auger emission
++
e-

## Slide 189: Sample atoms

## Slide 190: Sample atoms

## Slide 191: Primary ionization
e-
+

## Slide 192: Secondary ionizations
e-
e-
e-
e-
e-
e-
+
+
+
+
+
+

## Slide 193: Ionization track
e-
e-
e-
e-
e-
e-
e-
e-
e-
+
+
+
+
+
+
+
+
+
e-
+

## Slide 196: The Future of Structural Biology
10 Å

## Slide 197: What is the structure of “disorder” ?
a grand challenge for structural biology in the 21st century

## Slide 198: Big Question:
When a horse gallops, is there ever a moment where all four hooves leave the ground?
AD 1872

## Slide 199: Muybridge’s galloping horse (1878)
Real space
reciprocal
Sum(intensity) – Sum(density) = diffuse scatter

## Slide 200: Muybridge’s galloping horse (1878)
Real space
reciprocal
Fincoh – Fcoherent with density phases

## Slide 201: Muybridge’s galloping horse (1878)
Real space
reciprocal
RMS variation in density

## Slide 202: average structure: 1-sigma contour

## Slide 203: 21st Century Big Question: how do biomolecules work?
Beernink, Endrizzi, Alber & Schachman (1999). PNAS USA 96, 5388-5393.

## Slide 204: Averaging destroys resolution
10 Å

## Slide 205
Great! We have lots of crystals
but we have a problem with
non-isomorphism

## Slide 206
Dear James
The story of the two forms of lysozyme crystals goes back to about 1964 when it was found that the diffraction patterns from different crystals could be placed in one of two classes depending on their intensities. This discovery was a big set back at the time and I can remember a lecture title being changed from the 'The structure of lysozyme' to 'The structure of lysozyme two steps forward and one step back'. Thereafter the crystals were screened based on intensities of the (11,11,l) rows to distinguish them (e.g. 11,11,4 > 11,11,5 in one form and vice versa in another). Data were collected only for those that fulfilled the Type II criteria. (These reflections were easy to measure on the linear diffractometer because crystals were mounted to rotate about the diagonal axis). As I recall both Type I and Type II could be found in the same crystallisation batch . Although sometimes the external morphology allowed recognition this was not infallible.
The structure was based on Type II crystals. Later a graduate student Helen Handoll examined Type I. The work, which was in the early days and before refinement programmes, seemed to suggest that the differences lay in the arrangement of water or chloride molecules (Lysozyme was crystallised from NaCl). But the work was never written up. Keith Wilson at one stage was following this up as lysozyme was being used to test data collection strategies but I do not know the outcome.
An account of this is given in International Table Volume F (Rossmann and Arnold edited 2001) p760.
Tony North was much involved in sorting this out and if you wanted more info he would be the person to contact. I hope this is helpful. Do let me know if you need more.
Best wishes
Louise
Non-isomorphism in lysozyme

## Slide 207
Johnson’s ratio
Structure factor (e-)
Non-isomorphism in lysozyme

## Slide 208
Johnson’s ratio
Structure factor (e-)
Non-isomorphism in lysozyme

## Slide 210
Dear James
The story of the two forms of lysozyme crystals goes back to about 1964 when it was found that the diffraction patterns from different crystals could be placed in one of two classes depending on their intensities. This discovery was a big set back at the time and I can remember a lecture title being changed from the 'The structure of lysozyme' to 'The structure of lysozyme two steps forward and one step back'. Thereafter the crystals were screened based on intensities of the (11,11,l) rows to distinguish them (e.g. 11,11,4 > 11,11,5 in one form and vice versa in another). Data were collected only for those that fulfilled the Type II criteria. (These reflections were easy to measure on the linear diffractometer because crystals were mounted to rotate about the diagonal axis). As I recall both Type I and Type II could be found in the same crystallisation batch . Although sometimes the external morphology allowed recognition this was not infallible.
The structure was based on Type II crystals. Later a graduate student Helen Handoll examined Type I. The work, which was in the early days and before refinement programmes, seemed to suggest that the differences lay in the arrangement of water or chloride molecules (Lysozyme was crystallised from NaCl). But the work was never written up. Keith Wilson at one stage was following this up as lysozyme was being used to test data collection strategies but I do not know the outcome.
An account of this is given in International Table Volume F (Rossmann and Arnold edited 2001) p760.
Tony North was much involved in sorting this out and if you wanted more info he would be the person to contact. I hope this is helpful. Do let me know if you need more.
Best wishes
Louise
Non-isomorphism in lysozyme

## Slide 211: Non-isomorphism in lysozyme
RH 84.2% vs 71.9%
Riso = 44.5%
RMSD = 0.18 Å

## Slide 212: Non-isomorphism = dehydration?
= 1 nL
100 μm

## Slide 213: Elastic deformation = non-isomorphism

## Slide 214: Elastic deformation = non-isomorphism

## Slide 215: Plastic deformation = poor diffraction

## Slide 216: Plastic deformation = poor diffraction

## Slide 217: The Future of Structural Biology
Problem:
XFELs cannot solve the “resolution problem”
overfitting
non-isomorphism
radiation damage
What can we do?
4D modelling
probe weak density
more accurate data
humidity control during growth

## Slide 218: The Future of Structural Biology
Problem:
XFELs cannot solve the “resolution problem”
overfitting
non-isomorphism
radiation damage
small & fast = noisy
What can we do?
4D modelling
probe weak density
more accurate data
humidity control during growth

## Slide 219: average structure: 1-sigma contour

## Slide 220: Disorder is not “noise”

## Slide 221: Ringer: systematic map sampling
Lang et al. (2010) Protein Sci. 19, 1420-31.
Fraser, et al. (2009) Nature 462, 669-73.

## Slide 222: Ringer: systematic map sampling
electron density (e-/Å3)
χ1 angle (degrees)

## Slide 223: Ringer: systematic map sampling
electron density (e-/Å3)
χ1 angle (degrees)

## Slide 224: Ringer: systematic map sampling
electron density (e-/Å3)
χ1 angle (degrees)

## Slide 225: Ringer: systematic map sampling
electron density (e-/Å3)
χ1 angle (degrees)

## Slide 226: Independent estimates of F000
--
PDB | Method | F000 | diff
3DKE | Matthews et. al. (2008) | 121700 | +6.7%
 | END map estimate | 130500 | 
2NVH | Matthews et. al. (2006) | 92800 | +3.0%
 | END map estimate | 95656 | 
1AHO | Cerutti et. al. (2010) | 23177 | -5.0%
 | END map estimate | 22020 | 
Lang et al. (2014) PNAS USA 111, 237-242.

## Slide 227: Ringer: systematic map sampling
electron density (e-/Å3)
χ1 angle (degrees)

## Slide 228: Ringer: systematic map sampling
electron density (e-/Å3)
χ1 angle (degrees)

## Slide 229: "error bar" for electron density
structure factor (e-/cell)
spot index
fractional coordinate
electron density (e-/Å3)
Lang et al. (2014) PNAS USA 111, 237-242.

## Slide 230: "error bar" for electron density
structure factor (e-/cell)
spot index
fractional coordinate
electron density (e-/Å3)
Lang et al. (2014) PNAS USA 111, 237-242.

## Slide 231: "error bar" for electron density
structure factor (e-/cell)
spot index
fractional coordinate
electron density (e-/Å3)
Lang et al. (2014) PNAS USA 111, 237-242.

## Slide 232: "error bar" for electron density
structure factor (e-/cell)
spot index
fractional coordinate
electron density (e-/Å3)
Lang et al. (2014) PNAS USA 111, 237-242.

## Slide 233: "error bar" for electron density
structure factor (e-/cell)
spot index
fractional coordinate
electron density (e-/Å3)
Lang et al. (2014) PNAS USA 111, 237-242.

## Slide 234: "error bar" for electron density
structure factor (e-/cell)
spot index
fractional coordinate
electron density (e-/Å3)
Lang et al. (2014) PNAS USA 111, 237-242.

## Slide 235: "error bar" for electron density
structure factor (e-/cell)
spot index
fractional coordinate
electron density (e-/Å3)
Lang et al. (2014) PNAS USA 111, 237-242.

## Slide 236: "error bar" for electron density
structure factor (e-/cell)
spot index
fractional coordinate
electron density (e-/Å3)
Lang et al. (2014) PNAS USA 111, 237-242.

## Slide 237: "error bar" for electron density
fractional coordinate
electron density (e-/Å3)
Refinement
Against
Perturbed
Input
Data
Lang et al. (2014) PNAS USA 111, 237-242.

## Slide 238: How does the “error” compare to “sigma”?
Electron Number Density (e-/Å3)
normalized counts
Lang et al. (2014) PNAS USA 111, 237-242.

## Slide 239: How does the “error” compare to “sigma”?
Electron Number Density (e-/Å3)
normalized counts
Lang et al. (2014) PNAS USA 111, 237-242.

## Slide 240: 1 “sigma” is not a good criterion
Electron Number Density (e-/Å3)
normalized counts
Lang et al. (2014) PNAS USA 111, 237-242.

## Slide 241: two kinds of wands
Hampton screw top
Hampton magnetic

## Slide 242: CryoTongs

## Slide 243: proper seating in tongs

## Slide 244: tongs held by magnet

## Slide 245: Foam dewars

## Slide 246: you must choose, but choose wisely!
Uni-puck
SPring-8 SPACE
ALS/BCSB
ESRF/SPINE
Rigaku
SSRL/SAM

## Slide 247: working with canes

## Slide 248: working with canes

## Slide 249: dewar ice
foam insulation
warm, moist air
cold N2
liquid N2
ice

## Slide 250: dewar ice is charged

## Slide 251: drizzling/annealing tool

## Slide 252: Ice-B-Gone

## Slide 253: What’s wrong with a little ice?
1) makes the crystal hard to see
2) adds annoying rings to the diffraction pattern
3) indicative of crystal warming up at some point

## Slide 254: plunge cooling
foam insulation
warm, moist air
cold N2
liquid N2
ice

## Slide 255: plunge cooling
cold N2
foam insulation
warm, moist air
liquid N2
ice

## Slide 256: Warkentin method
foam insulation
liquid N2
Warkentin (2006) J. Appl. Crystallogr. 39, 805-811.

## Slide 257: dewar ice
foam insulation
warm, moist air
cold N2
liquid N2
ice
HEAT
HEAT

## Slide 258: dewar ice
vacuum insulation
warm, moist air
cold N2
liquid N2
ice
HEAT
HEAT