## Slide 1: Recent Developments
ALS Superbend 8.3.1

## Slide 2: ALS 8.3.1 optics
ADSC Quantum 315r
Superbend
Pt & Rh
on invar
Plane
Parabolic
M1 mirror
Pt & Rh
on Si
Torroidal
M2 mirror
two flat
Si(111)
monochromator
Protein Crystal
100 μm
aperture
Scatter
guard
convergence
angle slits
MacDowell et. al. (2004) J. Sync. Rad. 11 447-455
Notes: 2
Most notable acheivements are high brightness with very little “wasted” light. The discovery made by Howard Padmore? At ALS showed that a 2:1 demagnification cancels nearly all of the spherical abberations in the parabolic/torroid system. 30-50% of the total x-ray intensity passes through the 100um pinhole. The flux through the 100um pinhole at the superbend beamlines is comparable to that of the 5.0.2 wiggler or an APS bending magnet, but only 150 Watts of power must be dissipated by the beamline optics. In contrast, wiggler beamlines must disipate kilowatts. The optics here are much simpler and not exotic in any way. The system is stable and robust. There is never a need to re-focus the beam at a new wavelength. Only a 2-minute automatic tuneup is recommended for very large energy changes (more than 1-2 keV).

## Slide 3: ALS 8.3.1 optics
ADSC Quantum 315r
Superbend
Pt & Rh
on invar
Plane
Parabolic
M1 mirror
Pt & Rh
on Si
Torroidal
M2 mirror
two flat
Si(111)
monochromator
Protein Crystal
100 μm
aperture
Scatter
guard
convergence
angle slits
MacDowell et. al. (2004) J. Sync. Rad. 11 447-455
Notes: 3
Most notable acheivements are high brightness with very little “wasted” light. The discovery made by Howard Padmore? At ALS showed that a 2:1 demagnification cancels nearly all of the spherical abberations in the parabolic/torroid system. 30-50% of the total x-ray intensity passes through the 100um pinhole. The flux through the 100um pinhole at the superbend beamlines is comparable to that of the 5.0.2 wiggler or an APS bending magnet, but only 150 Watts of power must be dissipated by the beamline optics. In contrast, wiggler beamlines must disipate kilowatts. The optics here are much simpler and not exotic in any way. The system is stable and robust. There is never a need to re-focus the beam at a new wavelength. Only a 2-minute automatic tuneup is recommended for very large energy changes (more than 1-2 keV).

## Slide 4
Lattice | New | Old
Unattenuated |  | 
120 μm Cu filter |  | 
Vert x Horiz FWHM (μm) | 60 x 176 | 137 x 183
Flux (1011 Ph/s) | 3.0 | 2.1
Flicker (%/√Hz) | 0.37 | 0.66
11.1 keV, convergence slits “wide” (3 x 0.3 mRad), pixel = 0.83 μm, FWHM from knife edge

## Slide 5: Beam Flicker
time (seconds)
normalized flux

## Slide 6: Beam Flicker
time (seconds)
normalized flux
pinhole
removed

## Slide 7: Beam Flicker
1/f noise or “flicker noise”
comes from everything

## Slide 8: Beam Flicker
1/f noise or “flicker noise”
comes from everything

## Slide 9: Beam Flicker
1/f noise or “flicker noise”
comes from everything

## Slide 10: Beam Flicker
1/f noise or “flicker noise”
comes from everything

## Slide 11: Beam Flicker
1/f noise or “flicker noise”
comes from everything

## Slide 12: Beam Flicker
1/f noise or “flicker noise”
comes from everything

## Slide 13: Beam Flicker
1/f noise or “flicker noise”
comes from everything

## Slide 14: Beam Flicker
1/f noise or “flicker noise”
comes from everything

## Slide 15: Beam Flicker
1/f noise or “flicker noise”
comes from everything

## Slide 16: Beam Flicker
1/f noise or “flicker noise”
comes from everything

## Slide 17: Minimum exposure time
Exposure time (ms)
Overall I/σ(I)
Same dose
for each
data set

## Slide 18: Optimal exposure time (faint spots)
thr Optimal exposure time for data set (s)
tref exposure time of reference image (s)
bgref background level near weak spots on
reference image (ADU)
bg0 ADC offset of detector (ADU)
bghr optimal background level (via thr)
σ0 rms read-out noise (ADU)
gain ADU/photon
m multiplicity of data set (including partials)
adjust exposure
so this is ~100

## Slide 19
Lattice | New | Old
Unattenuated |  | 
120 μm Cu filter |  | 
Vert x Horiz FWHM (μm) | 60 x 176 | 137 x 183
Flux (1011 Ph/s) | 3.0 | 2.1
Flicker (%/√Hz) | 0.37 | 0.66
11.1 keV, convergence slits “wide” (3 x 0.3 mRad), pixel = 0.83 μm, FWHM from knife edge

## Slide 20
Damage is done
by photons/area
proportional to dose (MGy)
not time
not heat

## Slide 21
The number of photons scattered by a given crystal before it is dead
is independent
of data collection time

## Slide 22: Types of Rad Damage
Global
Specific

## Slide 23
Global damage
Howells, et al. (2009) J. Electron Spectrosc. Relat. Phenom. 170, 4-12.

## Slide 24
maximum tolerable dose (MGy)
1 2 3 5 7 10 20 40 70 100
1
10
100
103
resolution (Å)
Howells et al. (2009) J. Electron. Spectrosc. Relat. Phenom. 170 4-12
Global damage

## Slide 25
resolution (Å)
maximum tolerable dose (MGy)
1 2 3 5 7 10 20 40 70 100
1
10
100
103
Howells et al. (2009) J. Electron. Spectrosc. Relat. Phenom. 170 4-12
Global damage
10 MGy/Å

## Slide 26: 10 MGy/Å
what the is a MGy?
http://bl831.als.lbl.gov/
damage_rates.pdf
Holton J. M. (2009) J. Synchrotron Rad. 16 133-42

## Slide 27
How long will my crystal last?

## Slide 28: Dose-doubling concentration
Na | 44
Mg | 27
P | 9
S | 7
Cl | 5
K | 3.3
Ca | 2.7
Fe | 1.0
Cu | 0.67
Zn | 0.58
molar, at the Se edge based on RADDOSE
As | 0.44
Se | 0.42
Br | 2.9
I | 0.47
Gd | 0.21
Ta | 0.13
Pt | 0.17
Au | 0.16
Hg | 0.16
U | 0.23

## Slide 29: Dose-doubling concentration
Na | 44
Mg | 27
P | 9
S | 7
Cl | 5
K | 3.3
Ca | 2.7
Fe | 1.0
Cu | 0.67
Zn | 0.58
molar, at the Se edge based on RADDOSE
As | 0.44
Se | 0.42
Br | 2.9
I | 0.47
Gd | 0.21
Ta | 0.13
Pt | 0.17
Au | 0.16
Hg | 0.16
U | 0.23

## Slide 30: Riso vs dose
Riso (%)
change in dose (MGy)
data taken from Banumathi, et al. (2004) Acta Cryst. D 60, 1085-1093.

## Slide 31: Riso vs dose
Riso (%)
change in dose (MGy)
data taken from Banumathi, et al. (2004) Acta Cryst. D 60, 1085-1093.
Riso ≈ 0.7 %/MGy

## Slide 32: Specific damage
final, refined phases

## Slide 33: Specific damage
experimentally-obtained phases

## Slide 34: Specific damage
Holton, J. M. (2007) J. Synch. Rad. 14, 51-72.

## Slide 35: Specific damage
1.0
0.0
fraction
unconverted
Holton, J. M. (2007) J. Synch. Rad. 14, 51-72.

## Slide 36: No Damage in the Dark!
Experiment time (hours)
Peak absorbance (electrons)
4.6 5.0 5.4 5.8 6.2 6.6
0 5 10 15 20 25
Holton (2007) JSR 14, 51
-CH2-Se-CH3
-CH2 Se-CH3

## Slide 37: No Damage in the Dark!
Absorbed Dose (MGy)
Fraction unconverted
0.0 0.2 0.4 0.6 0.8 1.0
0 20 40 60 80 100
radiation damage is
FAST!
Holton (2007) JSR 14, 51
-CH2-Se-CH3
-CH2 Se-CH3

## Slide 38: Damage is fast!
fluence (1015 photons/mm2)
Fraction unconverted
0.0 0.2 0.4 0.6 0.8 1.0
0 50 100 150 200 250 300 350 400
Holton (2007) JSR 14, 51
100% Se-C bonds broken
-CH2-Se-CH3
-CH2 Se-CH3

## Slide 39: Damage is fast!
fluence (1015 photons/mm2)
Fraction unconverted
0.0 0.2 0.4 0.6 0.8 1.0
0 50 100 150 200 250 300 350 400
fraction of Se atoms hit
50% Se atoms hit
100% Se-C bonds broken
-CH2-Se-CH3
-CH2 Se-CH3
Holton (2007) JSR 14, 51

## Slide 40: Specific damage “rate”
Holton (2009) J. Synchrotron Rad. 16 133-42
world records:
MGy reaction reference
~45 global damage Owen et al. (2006)
5 Se-Met Holton (2007)
4 Hg-S Ramagopal et al. (2004)
3 S-S Murray et al. (2002)
1 Br-RNA Olieric et al. (2007)
? Cl-C ???
0.5 Mn in PS II Yano et al. (2005)
0.02 Fe in myoglobin Denisov et al. (2007)

## Slide 41
How long will my crystal last?

## Slide 42
Where:
IDL - average damage-limited intensity (photons/hkl) at a given resolution
105 - converting R from μm to m, re from m to Å, ρ from g/cm3 to kg/m3 and MGy to Gy
re - classical electron radius (2.818 x 10-15 m/electron)
h - Planck’s constant (6.626 x 10-34 J∙s)
c - speed of light (299792458 m/s)
fdecayed - fractional progress toward completely faded spots at end of data set
ρ - density of crystal (~1.2 g/cm3)
R - radius of the spherical crystal (μm)
λ - X-ray wavelength (Å)
fNH - the Nave & Hill (2005) dose capture fraction (1 for large crystals)
nASU - number of proteins in the asymmetric unit
Mr - molecular weight of the protein (Daltons or g/mol)
VM - Matthews’s coefficient (~2.4 Å3/Dalton)
H - Howells’s criterion (10 MGy/Å)
θ - Bragg angle
a2 - number-averaged squared structure factor per protein atom (electron2)
Ma - number-averaged atomic weight of a protein atom (~7.1 Daltons)
B - average (Wilson) temperature factor (Å2)
μ - attenuation coefficient of sphere material (m-1)
μen - mass energy-absorption coefficient of sphere material (m-1)
Self-calibrated damage limit
Holton & Frankel (2010) Acta D 66 393-408.
4

## Slide 43
Holton & Frankel (2010) Acta D 66 393-408.

## Slide 44: Multi-crystal future?

## Slide 45: 140-fold multiplicity
SUBSET OF INTENSITY DATA WITH SIGNAL/NOISE >= -3.0 AS FUNCTION OF RESOLUTION
RESOLUTION NUMBER OF REFLECTIONS COMPLETENESS R-FACTOR R-FACTOR COMPARED I/SIGMA R-meas CC(1/2) Anomal SigAno Nano
LIMIT OBSERVED UNIQUE POSSIBLE OF DATA observed expected Corr
9.17 188919 1061 1071 99.1% 3.9% 4.8% 188919 257.47 3.9% 100.0* 91* 5.024 450
6.49 356827 1933 1933 100.0% 5.2% 5.5% 356827 214.33 5.2% 100.0* 86* 3.836 882
5.30 466503 2515 2515 100.0% 7.2% 7.0% 466503 165.13 7.2% 100.0* 76* 3.257 1175
4.59 548405 2974 2974 100.0% 7.2% 6.6% 548405 175.42 7.3% 100.0* 67* 2.589 1403
4.10 615117 3353 3355 99.9% 7.7% 6.7% 615117 174.13 7.7% 100.0* 59* 2.264 1594
3.74 684947 3734 3737 99.9% 9.4% 8.3% 684947 143.09 9.4% 100.0* 49* 1.953 1783
3.47 743557 4051 4051 100.0% 11.2% 10.1% 743557 120.17 11.2% 100.0* 39* 1.696 1942
3.24 799722 4366 4366 100.0% 14.1% 13.9% 799722 91.14 14.1% 100.0* 30* 1.333 2103
3.06 839252 4598 4603 99.9% 19.5% 20.2% 839252 65.79 19.5% 100.0* 23* 1.117 2214
2.90 891398 4904 4908 99.9% 29.0% 31.7% 891398 44.85 29.1% 99.9* 17* 1.008 2369
2.77 931622 5148 5151 99.9% 40.5% 44.8% 931622 32.58 40.6% 99.8* 11* 0.901 2493
2.65 969629 5384 5387 99.9% 52.8% 58.8% 969629 25.16 52.9% 99.8* 10* 0.866 2605
2.54 997213 5574 5574 100.0% 67.4% 76.0% 997213 19.47 67.6% 99.6* 2 0.804 2705
2.45 1031222 5889 5889 100.0% 88.9% 101.2% 1031222 14.58 89.2% 99.2* 4 0.831 2859
2.37 788617 6008 6008 100.0% 109.3% 125.5% 788617 9.97 109.7% 98.1* 5 0.829 2925
2.29 614145 6252 6252 100.0% 138.2% 161.4% 614145 6.87 138.9% 96.1* 1 0.760 3037
2.22 470283 6481 6481 100.0% 197.1% 231.7% 470283 4.03 198.6% 83.5* -1 0.721 3159
2.16 259836 6629 6631 100.0% 227.3% 268.7% 259831 2.41 230.8% 46.9* -1 0.677 3224
2.10 36259 4169 6807 61.2% 154.4% 169.4% 36229 1.28 163.6% 47.0* -2 0.660 1999
2.05 13567 3372 7037 47.9% 170.1% 187.0% 13503 0.68 196.5% 25.7* 3 0.629 1578
total 12247040 88395 94730 93.3% 15.7% 16.4% 12246941 54.30 15.8% 100.0* 12* 1.217 42499
16 crystals, 360° each, inverse beam, 7235 eV AS MX1

## Slide 46
RESOLUTION COMPLETENESS R-FACTOR I/SIGMA R-meas CC(1/2) Anomal SigAno Nano
LIMIT OF DATA observed Corr
9.17 99.1% 3.9% 257.47 3.9% 100.0* 91* 5.024 450
6.49 100.0% 5.2% 214.33 5.2% 100.0* 86* 3.836 882
5.30 100.0% 7.2% 165.13 7.2% 100.0* 76* 3.257 1175
4.59 100.0% 7.2% 175.42 7.3% 100.0* 67* 2.589 1403
4.10 99.9% 7.7% 174.13 7.7% 100.0* 59* 2.264 1594
3.74 99.9% 9.4% 143.09 9.4% 100.0* 49* 1.953 1783
3.47 100.0% 11.2% 120.17 11.2% 100.0* 39* 1.696 1942
3.24 100.0% 14.1% 91.14 14.1% 100.0* 30* 1.333 2103
3.06 99.9% 19.5% 65.79 19.5% 100.0* 23* 1.117 2214
2.90 99.9% 29.0% 44.85 29.1% 99.9* 17* 1.008 2369
2.77 99.9% 40.5% 32.58 40.6% 99.8* 11* 0.901 2493
2.65 99.9% 52.8% 25.16 52.9% 99.8* 10* 0.866 2605
2.54 100.0% 67.4% 19.47 67.6% 99.6* 2 0.804 2705
2.45 100.0% 88.9% 14.58 89.2% 99.2* 4 0.831 2859
2.37 100.0% 109.3% 9.97 109.7% 98.1* 5 0.829 2925
2.29 100.0% 138.2% 6.87 138.9% 96.1* 1 0.760 3037
2.22 100.0% 197.1% 4.03 198.6% 83.5* -1 0.721 3159
2.16 100.0% 227.3% 2.41 230.8% 46.9* -1 0.677 3224
2.10 61.2% 154.4% 1.28 163.6% 47.0* -2 0.660 1999
2.05 47.9% 170.1% 0.68 196.5% 25.7* 3 0.629 1578
total 93.3% 15.7% 54.30 15.8% 100.0* 12* 1.217 42499
140-fold multiplicity: 16 crystals, 360° each, inverse beam, 7235 eV

## Slide 47: 140-fold multiplicity
18 σ
Phased anomalous difference Fourier
data Courtesy of Tom Peat & Janet Newman
16 σ

## Slide 48: 140-fold multiplicity
15σ = PO4
Phased anomalous difference Fourier
data Courtesy of Tom Peat & Janet Newman

## Slide 49: 140-fold multiplicity
~2σ = Mg?
Phased anomalous difference Fourier
data Courtesy of Tom Peat & Janet Newman

## Slide 50: 140-fold multiplicity
8.2σ = Mg?
DELFAN residual anomalous difference
data Courtesy of Tom Peat & Janet Newman

## Slide 51: Discerning Na+ from Mg++
f’’ (electrons)
DELFAN peak height (σ)
Mg
Ne
Na
F
O
N

## Slide 52: 140-fold multiplicity
7.4σ = Na
DELFAN residual anomalous difference
data Courtesy of Tom Peat & Janet Newman

## Slide 53: 140-fold multiplicity
7.3σ = S
DELFAN residual anomalous difference
data Courtesy of Tom Peat & Janet Newman

## Slide 54: 140-fold multiplicity
12σ = S
Phased anomalous difference Fourier
data Courtesy of Tom Peat & Janet Newman

## Slide 55: Suggested anomalous protocol:
2 wavelengths are better than 1
- (peak + inf)/2, and remote
If you must do one, do 7235 eV
360° in < 5 MGy
move detector
4X exposure
goto 2

## Slide 56
Lattice | New | Old
Unattenuated |  | 
120 μm Cu filter |  | 
Vert x Horiz FWHM (μm) | 60 x 176 | 137 x 183
Flux (1011 Ph/s) | 3.0 | 2.1
Flicker (%/√Hz) | 0.37 | 0.66
11.1 keV, convergence slits “wide” (3 x 0.3 mRad), pixel = 0.83 μm, FWHM from knife edge

## Slide 57
Fluidigm
Topaz® DC
http://www.fluidigm.com/home/fluidigm/docs/Datasheet_1.96DC.pdf
200 μm
Soliman, A. et al. (2011). Acta Cryst. D 67, 646-656.
250 μm
NatX-ray
CrystalQuick™ X
~50 μm tape
MiTeGen
MicroRT™
25 μm wall
25 μm wall
http://www.mitegen.com/products/micrort/micrort.shtml
PDMS
PDMS
http://www.mitegen.com/mic_catalog.php?c=insituplates
100 μm
MiTeGen
in-situ-1
~50 μm tape
Ng, J. et al. (2008). Acta Cryst. D 64, 189-197.
1000 μm
Microlytic
Crystal Former™
in-situ diffraction systems
Notes: Next slide please

## Slide 58: X-ray scattering “rules”:
1 μm crystal ≈ 1 μm water
≈ 1 μm plastic
≈ 0.1 μm glass
≈ 1000 μm air

## Slide 59: Si(111) vs multilayers
0.014%
Si(111)
1%
multilayer
4 Å
1.9 Å
resolution
spectral
dispersion

## Slide 60: Wide-bandpass MX?
Mono | Si(111) | multilayer | Difference (%) | good
Energy (eV) | 11111 | 12027.5 | +8.2 | Same
Scale | 1.0000 | 1.0964 | +9.6 | Yes
Rmerge (%) | 7.5 | 8.4 | +12 | No
Ranom (%) | 5.6 | 5.5 | -1.8 | Same
I/sd | 13.2 | 12.3 | -6.8 | Same
I/sd (1.5Å) | 2.3 | 1.8 | -21.7 | No
Ano Patt | 5.96 | 4.75 | -20.3 | Same
Ano Phased | 41 | 39.33 | -4.1 | Same
FOM | 0.238 | 0.192 | -19.3 | Same
CC | 0.555 | 0.647 | -6.1 | Yes
Rcryst (%) | 26.70 | 26.48 | -0.2 | Same
Ligand peak | 8.213 | 8.250 | +0.4 | Same

## Slide 61: Adjustable bandpass
Si(111)
Si(111)
Existing mono
Ge(11 1 1)
Ge(11 1 1)
3 mm
16.26°
d = 0.51 Å
2θ = 147.478°
For 12660 eV
6 mm
0.0007 % (8 hours)

## Slide 62: Adjustable bandpass
Si(111)
Si(111)
Existing mono
3 mm
Ge(11 1 1)
Ge(11 1 1)
0.014 % (30 min)

## Slide 63: Adjustable bandpass
ML
ML
Existing mono
3 mm
Ge(11 1 1)
Ge(11 1 1)
1 % (~1 second)

## Slide 64: Adjustable bandpass
Existing mono
Ge(11 1 1)
Ge(11 1 1)
3 mm
16.26°
d = 0.51 Å
2θ = 147.478°
For 12660 eV
6 mm
ML
ML
0.0007 % (8 hours)

## Slide 65
We’d love to use multiple crystals
but we have a problem with
non-isomorphism

## Slide 66: Multi-crystal strategies
Kendrew et al. (1960) "Structure of Myoglobin” Nature 185, 422-427.

## Slide 67
Dear James
The story of the two forms of lysozyme crystals goes back to about 1964 when it was found that the diffraction patterns from different crystals could be placed in one of two classes depending on their intensities. This discovery was a big set back at the time and I can remember a lecture title being changed from the 'The structure of lysozyme' to 'The structure of lysozyme two steps forward and one step back'. Thereafter the crystals were screened based on intensities of the (11,11,l) rows to distinguish them (e.g. 11,11,4 > 11,11,5 in one form and vice versa in another). Data were collected only for those that fulfilled the Type II criteria. (These reflections were easy to measure on the linear diffractometer because crystals were mounted to rotate about the diagonal axis). As I recall both Type I and Type II could be found in the same crystallisation batch . Although sometimes the external morphology allowed recognition this was not infallible.
The structure was based on Type II crystals. Later a graduate student Helen Handoll examined Type I. The work, which was in the early days and before refinement programmes, seemed to suggest that the differences lay in the arrangement of water or chloride molecules (Lysozyme was crystallised from NaCl). But the work was never written up. Keith Wilson at one stage was following this up as lysozyme was being used to test data collection strategies but I do not know the outcome.
An account of this is given in International Table Volume F (Rossmann and Arnold edited 2001) p760.
Tony North was much involved in sorting this out and if you wanted more info he would be the person to contact. I hope this is helpful. Do let me know if you need more.
Best wishes
Louise
Non-isomorphism in lysozyme

## Slide 68
Johnson’s ratio
Structure factor (e-)
Non-isomorphism in lysozyme

## Slide 69
Johnson’s ratio
Structure factor (e-)
Non-isomorphism in lysozyme

## Slide 71: Non-isomorphism in lysozyme
RH 84.2% vs 71.9%
Riso = 44.5%
RMSD = 0.18 Å

## Slide 72
Dear James
The story of the two forms of lysozyme crystals goes back to about 1964 when it was found that the diffraction patterns from different crystals could be placed in one of two classes depending on their intensities. This discovery was a big set back at the time and I can remember a lecture title being changed from the 'The structure of lysozyme' to 'The structure of lysozyme two steps forward and one step back'. Thereafter the crystals were screened based on intensities of the (11,11,l) rows to distinguish them (e.g. 11,11,4 > 11,11,5 in one form and vice versa in another). Data were collected only for those that fulfilled the Type II criteria. (These reflections were easy to measure on the linear diffractometer because crystals were mounted to rotate about the diagonal axis). As I recall both Type I and Type II could be found in the same crystallisation batch . Although sometimes the external morphology allowed recognition this was not infallible.
The structure was based on Type II crystals. Later a graduate student Helen Handoll examined Type I. The work, which was in the early days and before refinement programmes, seemed to suggest that the differences lay in the arrangement of water or chloride molecules (Lysozyme was crystallised from NaCl). But the work was never written up. Keith Wilson at one stage was following this up as lysozyme was being used to test data collection strategies but I do not know the outcome.
An account of this is given in International Table Volume F (Rossmann and Arnold edited 2001) p760.
Tony North was much involved in sorting this out and if you wanted more info he would be the person to contact. I hope this is helpful. Do let me know if you need more.
Best wishes
Louise
Non-isomorphism in lysozyme

## Slide 73: Non-isomorphism = dehydration?
= 1 nL
100 μm

## Slide 74: Response to dehydration
www.proteros.de

## Slide 75: Response to dehydration
www.proteros.de

## Slide 76: Response to dehydration
www.proteros.de

## Slide 77: Response to dehydration
www.proteros.de

## Slide 78: Response to dehydration
www.proteros.de

## Slide 79: Response to dehydration
www.proteros.de

## Slide 80: Response to dehydration
www.proteros.de

## Slide 81: Response to dehydration
www.proteros.de

## Slide 82: Ostwald Ripening
http://xray.bmc.uu.se/terese/tutorial6.html

## Slide 83: … but does it work for membrane proteins?
Efremov, R. G. & Sazanov, L. A. (2011)."Structure of the membrane domain of respiratory complex I", Nature 476, 414-420.
“To improve diffraction properties, crystals were slowly cooled to 4 C, placed into micro-dialysis buttons, and mother liquor was exchanged by dialysis for 0.1 M sodium acetate (pH 4.8), 0.15 M sodium tartrate (pH 5.0), 30% PEG4000, 0.1% sodium cholate and 12.5% glycerol over a period of 5 days.”
CCP4BB:
“We used it for a large membrane protein complex and diffraction improved from ~7 to 2.7 A. The crystal is fished out and put into mother liquor solution in the button, sealed with dialysis membrane and the button is then placed into about 5 mls of mother liquor with slightly higher PEG concentration. Then you just exchange outside buffer every day or so for solutions containing higher concentrations of PEG. We went from ~9 to 30 % PEG4000 in about a week. You can easily observe crystal under microscope and if it cracks - you went too far/too quickly with PEG and need to use a bit less next time. Also, this method allows you to control all other components of the dehydrating solution - we needed to decrease salt concentration at the same time as increasing PEG. You can also introduce/increase cryo-protectant concentration at the same time. With these crystals, otherwise excellent dehydration machines already mentioned did not work, possibly because the process had to be really slow.”

## Slide 84: opened drop: you have ~30 seconds!
equilibrated drop

## Slide 85: opened drop: you have ~30 seconds!
oil

## Slide 86: oiled drop: you have ~3 hours
oil

## Slide 87: Dabbing to remove excess
clean surface

## Slide 88: Dabbing to remove excess
clean surface

## Slide 89: Dabbing to remove excess
clean surface

## Slide 90: Dabbing to remove excess
clean surface

## Slide 91: Dabbing to remove excess
clean surface

## Slide 92: Dabbing to remove excess
clean surface

## Slide 93: Dabbing to remove excess
clean surface

## Slide 94: My crystal “dissolved” in the oil!
no, it didn’t…

## Slide 95: plunge cooling
foam insulation
warm, moist air
cold N2
liquid N2
ice

## Slide 96: plunge cooling
cold N2
foam insulation
warm, moist air
liquid N2
ice

## Slide 97: Warkentin method
foam insulation
liquid N2
Warkentin (2006) J. Appl. Crystallogr. 39, 805-811.

## Slide 98: Elastic deformation = non-isomorphism

## Slide 99: Elastic deformation = non-isomorphism

## Slide 100: Plastic deformation = poor diffraction

## Slide 101: Plastic deformation = poor diffraction

## Slide 102: Now What?
10 Å

## Slide 103: Ways to improve diffraction
New crystal form
Purify
  - add a column
  - heat shock
Bind something
  - Known ligand
  - Silver bullets
Change the clone
  - chop off floppy bits
  - homolog
  - Ala -> Lys mutants
Make Protein Sit Still!!!
Newman J. (2006) Acta Cryst. D62 27-31.

## Slide 104: Proteins move
Beernink, Endrizzi, Alber & Schachman (1999). PNAS USA 96, 5388-5393.

## Slide 105: Muybridge’s galloping horse (1878)

## Slide 106: average structure: galloping horse

## Slide 107: average structure: galloping horse
1 “sigma” contour

## Slide 108: Horse: real and reciprocal

## Slide 109: Supporting a model with data

## Slide 110: Molecular Dynamics Simulation
1aho
Scorpion toxin
0.96 Å resolution
64 residues
Solvent: H20 + acetate
Cerutti et al. (2010).J. Phys. Chem. B 114, 12811-12824.
using real
crystal’s lattice

## Slide 111: 30 conformers from 24,000

## Slide 112: Electron density from 24,000 conformers

## Slide 113: Regular model with real data!

## Slide 114: Molecular Dynamics vs Observation
Fobs
1aho.cif
1aho.pdb
Fsim
Fcalc
Fcalc
Rcryst= 0.137
Rcryst= 0.116
Rvault = 0.69
refined_vs_Fsim.pdb
LSQ rmsd = 0.43Å
rmsd = 1.05 Å
1aho 64-residue scorpion toxin in water to 1.0 Å resolution
Rvault = 0.48 to 4 Å
Riso =

## Slide 115: Molecular Dynamics vs Observation
Sodium acetate trihydrate
5 atoms
3 waters
0.9 Å resolution
C2/c
Rvault = 0.01

## Slide 116: Molecular Dynamics vs Observation
fav8
8 residues
4 waters
1.0 Å resolution
P1

## Slide 117: Molecular Dynamics vs Observation
Fobs
fav8.fcf
fav8.cif
Fsim
Fcalc
Fcalc
Rcryst= 0.13
Rcryst= 0.16
Rvault= 0.23
refined_vs_Fsim.pdb
LSQ rmsd = 0.091Å & 0.15 Å
rmsd = 0.20 Å
“fav8” 8-residue aromatic peptide with 4 waters to 1.0 Å resolution
R1 = 0.041
With 4σ cutoff!
Rvault = 0.20 to 2Å

## Slide 118: MD-predicted water structure
Simulated
density
Janowski et al (2013) JACS 135, 7938-7948

## Slide 119: MD-predicted water structure
Simulated
density
2mFo-DFc
Janowski et al (2013) JACS 135, 7938-7948

## Slide 120: MD-predicted Val rotamer
Janowski et al (2013) JACS 135, 7938-7948

## Slide 121: MD-predicted Val rotamer
confs | Rwork | Rfree
one | 4.11% | 5.84%
A & B | 3.89% | 5.53%
Janowski et al (2013) JACS 135, 7938-7948

## Slide 122: Molecular Dynamics Simulation
1aho
Scorpion toxin
0.96 Å resolution
64 residues
Solvent: H20 + acetate
Cerutti et al. (2010).J. Phys. Chem. B 114, 12811-12824.
using real
crystal’s lattice

## Slide 123: Electron density from 24,000 conformers

## Slide 124: MLFSOM simulation

## Slide 125: Map from simulated data
1 “sigma” 2Fsim-Fc
2.5 “sigma” Fsim-Fc
1.0 Å data
Rcryst= 0.137
Rfree = 0.159

## Slide 126: The “right” difference map
1 “sigma” 2Fsim-Fc
Fright – (2Fsim-Fc)
1.0 Å data
Rcryst= 0.137
Rfree = 0.159

## Slide 127: Real data!
1 “sigma” 2Fsim-Fc
2.5 “sigma” Fsim-Fc
1.0 Å data
Rcryst= 0.116
Rfree = 0.135

## Slide 128: Ringer: systematic map sampling
Lang et al. (2010) Protein Sci. 19, 1420-31.
Fraser, et al. (2009) Nature 462, 669-73.

## Slide 129: Ringer: systematic map sampling
electron density (e-/Å3)
χ1 angle (degrees)

## Slide 130: Ringer: systematic map sampling
electron density (e-/Å3)
χ1 angle (degrees)

## Slide 131: Ringer: systematic map sampling
electron density (e-/Å3)
χ1 angle (degrees)

## Slide 132: Ringer: systematic map sampling
electron density (e-/Å3)
χ1 angle (degrees)

## Slide 133: Disorder is not noise

## Slide 134: "error bar" for electron density
structure factor (e-/cell)
spot index
fractional coordinate
electron density (e-/Å3)

## Slide 135: "error bar" for electron density
structure factor (e-/cell)
spot index
fractional coordinate
electron density (e-/Å3)

## Slide 136: "error bar" for electron density
structure factor (e-/cell)
spot index
fractional coordinate
electron density (e-/Å3)

## Slide 137: "error bar" for electron density
structure factor (e-/cell)
spot index
fractional coordinate
electron density (e-/Å3)

## Slide 138: "error bar" for electron density
structure factor (e-/cell)
spot index
fractional coordinate
electron density (e-/Å3)

## Slide 139: "error bar" for electron density
structure factor (e-/cell)
spot index
fractional coordinate
electron density (e-/Å3)

## Slide 140: "error bar" for electron density
structure factor (e-/cell)
spot index
fractional coordinate
electron density (e-/Å3)

## Slide 141: "error bar" for electron density
structure factor (e-/cell)
spot index
fractional coordinate
electron density (e-/Å3)

## Slide 142: "error bar" for electron density
fractional coordinate
electron density (e-/Å3)
Refinement
Against
Perturbed
Input
Data

## Slide 143: How does the “error” compare to “sigma”?
Electron Number Density (e-/Å3)
normalized counts

## Slide 144: How does the “error” compare to “sigma”?
Electron Number Density (e-/Å3)
normalized counts

## Slide 145: 1 “sigma” is not a good criterion
Electron Number Density (e-/Å3)
normalized counts

## Slide 146: 1.0 “sigma” multi-conformer (4nx4)

## Slide 147: 7x noise multi-conformer (4nx4)

## Slide 148: 1.0 “sigma” no ligand built (2pxr)

## Slide 149: 1.0 “sigma” multi-conformer (4nx4)

## Slide 150: 7x noise multi-conformer (4nx4)
Rcryst/Rfree =
0.1467/0.1944

## Slide 151: 7x noise no ligand built (2pxr)
Rcryst/Rfree =
0.1474/0.2032

## Slide 152: Summary
There is “life” below 1 “sigma”
100 ADU/pixel
10 μm for lysozyme
Anomalous = “Attenu-wait” at 7235 eV
The future is 4D models?
http://bl831.als.lbl.gov/END/RAPID/
http://bl831.als.lbl.gov/xtalsize.html
http://bl831.als.lbl.gov/~jamesh/powerpoint/Plexxikon_2014.ppt

## Slide 153: Classes of error in MX
Dependence on signal
Time
none sqrt proportional
none
1/sqrt
1/prop.
CCD
Read-out
Photon
counting
Beam flicker
Shutter jitter
Sample vibration
Detector calibration
attenuation
partiality
Non-isomorphism
Radiation damage

## Slide 154: attenuation factor
Bouguer, P. (1729). Essai d'optique sur la gradation de la lumière.
Lambert, J. H. (1760). Photometria: sive De mensura et gradibus luminis, colorum et umbrae. E. Klett.
Beer, A. (1852)."Bestimmung der Absorption des rothen Lichts in farbigen Flüssigkeiten", Ann. Phys. Chem 86, 78-90.
A = = exp(-μt)
IT
Ibeam
t
μ

## Slide 155: attenuation factor
Bouguer, P. (1729). Essai d'optique sur la gradation de la lumière.
Lambert, J. H. (1760). Photometria: sive De mensura et gradibus luminis, colorum et umbrae. E. Klett.
Beer, A. (1852)."Bestimmung der Absorption des rothen Lichts in farbigen Flüssigkeiten", Ann. Phys. Chem 86, 78-90.
A = = exp(-μ(tin+ tout))
IT
Ibeam
t
tin
tout
A = = exp(-μt)
IT
Ibeam
μ

## Slide 156: attenuation factor
Bouguer, P. (1729). Essai d'optique sur la gradation de la lumière.
Lambert, J. H. (1760). Photometria: sive De mensura et gradibus luminis, colorum et umbrae. E. Klett.
Beer, A. (1852)."Bestimmung der Absorption des rothen Lichts in farbigen Flüssigkeiten", Ann. Phys. Chem 86, 78-90.
A = = exp(-μ(tin+ tout))
IT
Ibeam
tin
tout
tin
tout
tin
tout
μ

## Slide 157: attenuation factor
Bouguer, P. (1729). Essai d'optique sur la gradation de la lumière.
Lambert, J. H. (1760). Photometria: sive De mensura et gradibus luminis, colorum et umbrae. E. Klett.
Beer, A. (1852)."Bestimmung der Absorption des rothen Lichts in farbigen Flüssigkeiten", Ann. Phys. Chem 86, 78-90.
A = = exp[-μxtal(txi+ txo)
-μsolvent(tsi + tso)]
IT
Ibeam
μxtal
txi
txo
tsi
tso
txi
txo
tsi
tso
txi
txo
tsi
tso
μsolvent

## Slide 158: Where do photons go?
beamstop
Transmitted (98%)
Protein
1A x-rays
attenuation correction cannot be > ~2%
for 100 μm xtal at 1 Å

## Slide 159: Classes of error in MX
Dependence on signal
Time
none sqrt proportional
none
1/sqrt
1/prop.
CCD
Read-out
Photon
counting
Beam flicker
Shutter jitter
Sample vibration
Detector calibration
attenuation
partiality
Non-isomorphism
Radiation damage

## Slide 160: Beam Flicker
1/f noise or “flicker noise”
comes from everything

## Slide 161: Beam Flicker
1/f noise or “flicker noise”
comes from everything

## Slide 162: Beam Flicker
1/f noise or “flicker noise”
comes from everything

## Slide 163: Beam Flicker
1/f noise or “flicker noise”
comes from everything

## Slide 164: Beam Flicker
1/f noise or “flicker noise”
comes from everything

## Slide 165: Beam Flicker
1/f noise or “flicker noise”
comes from everything

## Slide 166: Beam Flicker
1/f noise or “flicker noise”
comes from everything

## Slide 167: Beam Flicker
1/f noise or “flicker noise”
comes from everything

## Slide 168: Beam Flicker
1/f noise or “flicker noise”
comes from everything

## Slide 169: Beam Flicker
1/f noise or “flicker noise”
comes from everything

## Slide 170: Beam Flicker
time (seconds)
normalized flux

## Slide 171: Beam Flicker
time (seconds)
normalized flux
pinhole
removed

## Slide 172: Beam Flicker
flicker noise (%/√Hz)
CC to correct model

## Slide 173: Classes of error in MX
Dependence on signal
Time
none sqrt proportional
none
1/sqrt
1/prop.
CCD
Read-out
Photon
counting
Beam flicker
Shutter jitter
Sample vibration
Detector calibration
attenuation
partiality
Non-isomorphism
Radiation damage

## Slide 174: Shutter Jitter
open
closed
shutter jitter

## Slide 175: Shutter Jitter
frame number
PHIZ from MOSFLM (deg)

## Slide 176: Shutter Jitter
frame number
PHIZ from MOSFLM (deg)

## Slide 177: Shutter Jitter
frame number
shutter timing error (ms)
rms 0.47 ms

## Slide 178: Shutter Jitter
rms timing error (% exposure)
CC to correct model

## Slide 179: Classes of error in MX
Dependence on signal
Time
none sqrt proportional
none
1/sqrt
1/prop.
CCD
Read-out
Photon
counting
Beam flicker
Shutter jitter
Sample vibration
Detector calibration
attenuation
partiality
Non-isomorphism
Radiation damage

## Slide 180: xtal vibration noise
diffracted beam
incident beam
see: Alkire et al. (2008)."Is your cold-stream working for you or against you? An in-depth look at temperature and sample motion", J. Appl. Cryst. 41, 1122-1133.

## Slide 181: xtal vibration noise
incident beam
diffracted beam
see: Alkire et al. (2008)."Is your cold-stream working for you or against you? An in-depth look at temperature and sample motion", J. Appl. Cryst. 41, 1122-1133.

## Slide 182: xtal vibration noise
incident beam
diffracted beam
see: Alkire et al. (2008)."Is your cold-stream working for you or against you? An in-depth look at temperature and sample motion", J. Appl. Cryst. 41, 1122-1133.

## Slide 183: xtal vibration noise
incident beam
diffracted beam
see: Alkire et al. (2008)."Is your cold-stream working for you or against you? An in-depth look at temperature and sample motion", J. Appl. Cryst. 41, 1122-1133.

## Slide 184: xtal vibration noise
incident beam
diffracted beam
see: Alkire et al. (2008)."Is your cold-stream working for you or against you? An in-depth look at temperature and sample motion", J. Appl. Cryst. 41, 1122-1133.

## Slide 185: xtal vibration noise
incident beam
diffracted beam
see: Alkire et al. (2008)."Is your cold-stream working for you or against you? An in-depth look at temperature and sample motion", J. Appl. Cryst. 41, 1122-1133.

## Slide 186: xtal vibration noise
incident beam
diffracted beam
see: Alkire et al. (2008)."Is your cold-stream working for you or against you? An in-depth look at temperature and sample motion", J. Appl. Cryst. 41, 1122-1133.

## Slide 187: xtal vibration noise
incident beam
see: Alkire et al. (2008)."Is your cold-stream working for you or against you? An in-depth look at temperature and sample motion", J. Appl. Cryst. 41, 1122-1133.

## Slide 188: xtal vibration noise
incident beam
diffracted beam
see: Alkire et al. (2008)."Is your cold-stream working for you or against you? An in-depth look at temperature and sample motion", J. Appl. Cryst. 41, 1122-1133.

## Slide 189: xtal vibration noise
incident beam
see: Alkire et al. (2008)."Is your cold-stream working for you or against you? An in-depth look at temperature and sample motion", J. Appl. Cryst. 41, 1122-1133.

## Slide 190: Classes of error in MX
Dependence on signal
Time
none sqrt proportional
none
1/sqrt
1/prop.
CCD
Read-out
Photon
counting
Beam flicker
Shutter jitter
Sample vibration
Detector calibration
attenuation
partiality
Non-isomorphism
Radiation damage

## Slide 191: Detector calibration
plastic
air
fibers
Gd2O2S:Tb

## Slide 192: Detector calibration
plastic
air
fibers
Gd2O2S:Tb
x-rays

## Slide 193: Detector calibration
plastic
air
fibers
Gd2O2S:Tb
x-rays

## Slide 194: Detector calibration
plastic
air
fibers
Gd2O2S:Tb
x-rays

## Slide 195: Detector calibration
plastic
air
fibers
Gd2O2S:Tb
x-rays

## Slide 196: Detector calibration
plastic
air
fibers
Gd2O2S:Tb
x-rays

## Slide 197: Detector calibration
plastic
air
fibers
Gd2O2S:Tb
x-rays

## Slide 198: Detector calibration: 7247 eV
target:
oil
distance:
900 mm
2θ: 12°

## Slide 199: Detector calibration: 7235 eV
target:
oil
distance:
900 mm
2θ: 12°

## Slide 200: Detector calibration: ALS 8.3.1

## Slide 201: Detector calibration errors: detector 2

## Slide 202: Detector calibration errors: detector 3

## Slide 203: Detector calibration
calibration error (%)
megapixels

## Slide 204: Detector calibration
photon energy (keV)
calibration error (%)
good!
good!
bad!

## Slide 205
7247 eV

## Slide 206
7235 eV

## Slide 207: Detector calibration
7223 eV

## Slide 208: Calibration Error

## Slide 209: Calibration Error

## Slide 210: Calibration Error

## Slide 211: Calibration Error

## Slide 212: Calibration Error
down

## Slide 213: Calibration Error
down
up

## Slide 214: Calibration Error
down
up
Rseparate

## Slide 215: Calibration Error
odd
even
Rmixed

## Slide 216: Calibration Error
separate:
2.5%

## Slide 217: Calibration Error
separate:
mixed:
2.5%
0.9%

## Slide 218: Calibration Error
separate:
mixed:
2.5%
0.9%
2.5%2-0.9%2 = 2.3%2

## Slide 219: Detector calibration errors
Dan Schuette PhD Thesis (2008) Fig 6.22 page 198, Gruner Lab, Cornell University.
Pixel arrays are not immune!

## Slide 220: Calibration Error: Q315r vs Pilatus
Holton, Frankel, Gonzalez, Waterman and Wang (2010) in preparation
3.5
3.0
2.5
2.0
1.5
1.0
0.5
average change in spot intensity (%)
distance between spots (mm)
0.1 1 10 100
Pilatus
Q315r
anomalous mates
> 100 mm apart

## Slide 221: What is holding us back?
Weak spots (high-res)
  background
MAD/SAD (small differences)
  fractional errors
( if not rad dam! )

## Slide 222: What is holding us back?
Weak spots (high-res)
  background
  solution: use as few pixels as possible
MAD/SAD (small differences)
  fractional errors
  solution: use as many pixels as possible
( if not rad dam! )

## Slide 223
Holton & Frankel (2010) Acta D 66 393-408.

## Slide 224
Holton & Frankel (2010) Acta D 66 393-408.

## Slide 225
Holton & Frankel (2010) Acta D 66 393-408.

## Slide 226
Holton & Frankel (2010) Acta D 66 393-408.

## Slide 227
anomalous signal
Crick, F. H. C. & Magdoff, B. S. (1956) Acta Crystallogr. 9, 901-908.
Hendrickson, W. A. & Teeter, M. M. (1981) Nature 290, 107-113.
# sites
MW (Da)
ΔF
F
≈ 1.2 f”
√
World record!
ΔF/F = 0.5%
Wang, Dauter & Dauter (2006) Acta Cryst. D 62, 1475-1483.

## Slide 228: Fractional error
mult > (—)2
~3%
<ΔF/F>

## Slide 229
Holton & Frankel (2010) Acta D 66 393-408.

## Slide 230: Dose slicing
crystal’s useful life
N
photons
N
photons
N
photons
unacceptable
damage
unacceptable
read noise