## Slide 1: Acknowledgements
UCSF LBNL SLAC
ALS 8.3.1: TomAlberTron
NIH NIGMS R01 GM124149
NIH NIGMS P50 GM082250
NIH NIGMS P30 GM124169
NIH NIGMS R01 GM117126
DOE-BER Integrated Diffraction Analysis Technologies (IDAT)
Plexxikon, Inc.
NIH NIGMS P41 GM103393
NSF DBI-1625906
The Advanced Light Source is supported by the Director, Office of Science, Office of Basic Energy Sciences, Materials Sciences Division, of the US Department of Energy under contract No. DE-AC02-05CH11231 at Lawrence Berkeley National Laboratory.
Robert Stroud James Fraser Nevan Krogan
Paul Adams John Tainer Greg Hura Nick Sauter
Aina Cohen Art Lyubimov Jeney Wierman

## Slide 2
ber
structural
bio
portal.org

## Slide 3: Powerpoint file available:
https://bl831.als.lbl.gov/~jamesh/powerpoint/RapiData_XDS_2021.pptx
JMHolton@lbl.gov

## Slide 4: XDS X-ray diffraction data reduction engine
Diffraction image data
List of spot intensities (text)

## Slide 5
XDS native user interface: a text file!

## Slide 7
XDS native user interface

## Slide 8
The “prompt”
Cursor
does not move
with mouse
Mouse
can be used
To select and paste
But all pasted text
Goes to the cursor

## Slide 9: How to start XDS
~]$ cd
~]$ mkdir xds_tutorial
~]$ cd xds_tutorial
~]$ ln –sf /data/holton/demo .
~]$ pxproc_dialog.sh
~]$ ssh pxproc12
@pxproc ~]$ cd xds_tutorial
]$ xdsgui
Space!
underscore
prompt
Do not type!
Enter
Space!

## Slide 10
xdsgui – a graphical interface for XDS

## Slide 11
xdsgui – load an image

## Slide 12
xdsgui – a GUI for text!

## Slide 13
xdsconv – make an “mtz” file
output_file_name.mtz

## Slide 14: Supply & Demand:

## Slide 15
Nov 2020- Dec 2021
20-Nov
20-Dec
21-Jan
21-Feb
21-Mar
21-Apr
21-May
21-Jun
21-Jul
21-Aug
21-Sep
21-Oct
21-Nov
21-Dec

## Slide 16
Nov 2022- Dec 2023
22-Nov
22-Dec
23-Jan
23-Feb
23-Mar
23-Apr
23-May
23-Jun
23-Jul
23-Aug
23-Sep
23-Oct
23-Nov
23-Dec
assuming here that basis schedules from 2021 will repeat

## Slide 17
Nov 2023- Dec 2024
23-Nov
23-Dec
24-Jan
24-Feb
24-Mar
24-Apr
24-May
24-Jun
24-Jul
24-Aug
24-Sep
24-Oct
24-Nov
24-Dec
assuming here that basis schedules from 2021 will repeat

## Slide 18
xdsgui – nuts, bolts, bells & whistles

## Slide 19
XYCORR– setting things up

## Slide 20: Beam center vs “origin”
direct beam

## Slide 21: Beam center vs “origin”
origin
direct beam

## Slide 22: Correct beam center
155.063 155.647
159.353 155.063
159.301 155.011
1512.73 1554.57
155.063 159.356 -250
ADXV
HKL2000
MOSFLM
XDS
DIALS

## Slide 23
XYCORR– setting things up

## Slide 24
XYCORR– setting things up

## Slide 25: Pilatus: 1-module shift

## Slide 26
Pilatus: 1-module shift

## Slide 27
Pilatus: 1-module shift

## Slide 28: Detector pixels are 3D
51 μm
31 μm thick
17%
loss
Typical CCD
Mean Penetration depth 12 keV:
17.5 μm Gd2O2S 237 μm Si
8%
side
loss
172 μm
26% loss
320 μm thick
Pilatus
53%
side
loss
photon
75 μm
15%
loss
450 μm thick
Eiger
76%
side
loss
photon
Arrival angle of 2 Å spot: 29°

## Slide 29
XYCORR– setting things up

## Slide 30
INIT – look at the background

## Slide 31
INIT – look at the background

## Slide 32
COLSPOT – the spot picker

## Slide 33
COLSPOT – the spot picker

## Slide 34
IDXREF – auto-indexing

## Slide 35: XPARM.XDS
XPARM.XDS VERSION Feb 5, 2021 BUILT=20210323
1 270.0000 0.2000 0.999971 0.007404 -0.001831
0.979460 0.000711 0.000291 1.020970
96 78.8215 78.8215 37.1784 90.000 90.000 90.000
36.338257 9.327419 -69.320709
26.357738 70.532730 23.307367
30.559622 -16.002089 13.866352
1 2463 2527 0.172000 0.172000
1231.500000 1263.500000 189.160004
1.000000 0.000000 0.000000
0.000000 1.000000 0.000000
0.000000 0.000000 1.000000
1 1 2463 1 2527
0.00 0.00 0.00 1.00000 0.00000 0.00000 0.00000 1.00000 0.00000
First image number
Starting phi
Delta-phi
spindle vector
beam vector
unit cell
a
b
c
space
group
number
wavelen
cell vectors
real space
det panels
origin
distance
pixels
pixel size
x
y
z
Detector fast/slow
Panel #

## Slide 36
IDXREF – auto-indexing
BRAVAIS- POSSIBLE SPACE-GROUPS FOR PROTEIN CRYSTALS
TYPE [SPACE GROUP NUMBER,SYMBOL]
aP [1,P1]
mP [3,P2] [4,P2(1)]
mC,mI [5,C2]
oP [16,P222] [17,P222(1)] [18,P2(1)2(1)2] [19,P2(1)2(1)2(1)]
oC [21,C222] [20,C222(1)]
oF [22,F222]
oI [23,I222] [24,I2(1)2(1)2(1)]
tP [75,P4] [76,P4(1)] [77,P4(2)] [78,P4(3)] [89,P422] [90,P42(1)2]
[91,P4(1)22] [92,P4(1)2(1)2] [93,P4(2)22] [94,P4(2)2(1)2]
[95,P4(3)22] [96,P4(3)2(1)2]
tI [79,I4] [80,I4(1)] [97,I422] [98,I4(1)22]
hP [143,P3] [144,P3(1)] [145,P3(2)] [149,P312] [150,P321] [151,P3(1)12]
[152,P3(1)21] [153,P3(2)12] [154,P3(2)21] [168,P6] [169,P6(1)]
[170,P6(5)] [171,P6(2)] [172,P6(4)] [173,P6(3)] [177,P622]
[178,P6(1)22] [179,P6(5)22] [180,P6(2)22] [181,P6(4)22] [182,P6(3)22]
hR [146,R3] [155,R32]
cP [195,P23] [198,P2(1)3] [207,P432] [208,P4(2)32] [212,P4(3)32]
[213,P4(1)32]
cF [196,F23] [209,F432] [210,F4(1)32]
cI [197,I23] [199,I2(1)3] [211,I432] [214,I4(1)32]

## Slide 37: Trick Question: What is the space group?
a = 63 b = 63 c = 63
α = 90 β = 90 γ = 90

## Slide 38: What you know when:
Autoindexing: Lattice (14)
Scaling: Point Group (43)
Phasing: Space Group (84)

## Slide 39: Pitfalls: sub-space groups
P622
What else?
P6122 P6522 P6222 P6422 P6322
P6 P61 P65 P62 P64 P63
P312 P321 P3112 P3121 P3212 P3221
P3 P31 P32
P2 P21 C2 C222 C2221
P1
zanuda

## Slide 40: Asymmetric Unit (ASU)
Smallest part of the crystal
that you can use to build up the rest of the crystal
with symmetry operations
Unit Cell
Smallest part of the crystal
that you can use to build up the rest of the crystal
with translation symmetry only
“NCS ASU”
Builds ASU with Non-crystallographic symmetry operations

## Slide 41: Frequency of Space Groups in PDB
13007 | P212121 | 566 | R32 | 191 | P6322 | 74 | P4322 | 36 | P3212
9629 | P21 | 491 | P6522 | 187 | P213 | 71 | P4122 | 30 | P622
5300 | C2 | 474 | P61 | 161 | P6 | 70 | P64 | 24 | F4132
3021 | P21212 | 421 | P65 | 144 | P4212 | 70 | P6222 | 22 | P4332
2543 | C2221 | 353 | P41 | 142 | I23 | 67 | P6422 | 22 | P422
2109 | P1 | 351 | P63 | 132 | F432 | 66 | F222 | 22 | P23
2030 | P43212 | 286 | P31 | 124 | P321 | 63 | P62 | 19 | P3112
1565 | P3221 | 257 | I422 | 115 | I41 | 59 | P4132 | 18 | P4222
1359 | P41212 | 253 | P42212 | 103 | P3 | 46 | P2221 | 18 | I4132
1198 | P3121 | 249 | I4 | 99 | I213 | 43 | P4 | 16 | P432
1147 | I222 | 246 | P43 | 94 | C222 | 43 | I432 | 8 | P4232
825 | R3 | 239 | I4122 | 93 | P2 | 40 | P42 | 3 | P222
620 | P6122 | 193 | P32 | 82 | I212121 | 37 | F23 | 2 | P312

## Slide 42
DEFPIX – pixels good and bad

## Slide 43
DEFPIX – pixels good and bad

## Slide 44
INTEGRATE – measure the spots!

## Slide 45
INTEGRATE – measure the spots!

## Slide 46
Ewald’s “mosaic” picture

## Slide 47
Ewald’s “mosaic” picture

## Slide 48
Ewald’s “mosaic” picture
What is
this stuff?

## Slide 49
Darwin’s original picture

## Slide 50
“mosaicity” with visible light

## Slide 51: mosaic spread = 0 º

## Slide 52: mosaic spread = 0.1º

## Slide 53: mosaic spread = 0.2º

## Slide 54: mosaic spread = 0.4º

## Slide 55: mosaic spread = 0.6º

## Slide 56: mosaic spread = 0.8º

## Slide 57: mosaic spread = 1.0º

## Slide 58: mosaic spread = 1.5º

## Slide 59: mosaic spread = 2.0º

## Slide 60: mosaic spread = 2.5º

## Slide 61: mosaic spread = 3.2º

## Slide 62: mosaic spread = 6.4º

## Slide 63: mosaic spread = 12.8º

## Slide 64: Spot Shape
Ewald sphere
(h,k,l)

## Slide 65: Spot Shape
Ewald sphere
(h,k,l)
d*

## Slide 66: Spot Shape
Ewald sphere
spindle axis
(h,k,l)
d*

## Slide 67: Spot Shape
Ewald sphere
spindle axis
Φ circle
(h,k,l)

## Slide 68: Spot Shape
Ewald sphere
spindle axis
Φ circle
(h,k,l)

## Slide 69: Spot Shape
Ewald sphere
spindle axis
Φ circle
diffracted ray
(h,k,l)

## Slide 70: Spot Shape
Ewald sphere
spindle axis
Φ circle
diffracted ray
(h,k,l)
d*
λ*
λ*

## Slide 71: Spot Shape
Ewald sphere
spindle axis
Φ circle
diffracted ray
(h,k,l)
d*
λ*
λ*
θ
λ = 2 d sin(θ)

## Slide 72: mosaic spread
Ewald sphere
spindle axis
Φ circle
diffracted ray
(h,k,l)
λ*
λ*
d*
d*

## Slide 73: mosaic spread
Ewald sphere
spindle axis
Φ circle
diffracted rays
(h,k,l)
d*
d*

## Slide 74
INTEGRATE – measure the spots!

## Slide 75: beam divergence
Ewald sphere
spindle axis
Φ circle
diffracted ray
(h,k,l)
d*
λ*
λ*

## Slide 76: beam divergence
spindle axis
Φ circle
diffracted ray
(h,k,l)
d*
Ewald sphere
λ*
λ*

## Slide 77: divergence = 0 º

## Slide 78: divergence = 0.3 º

## Slide 79: spectral dispersion
Ewald sphere
spindle axis
Φ circle
diffracted ray
(h,k,l)
d*
λ*
λ*

## Slide 80: spectral dispersion
Ewald sphere
spindle axis
Φ circle
diffracted ray
(h,k,l)
d*
λ’*
λ’*

## Slide 81: dispersion = 0

## Slide 82: dispersion = 0.014% (Si 111)

## Slide 83: dispersion = 0.25% (CuKα)

## Slide 84: dispersion = 0.6%

## Slide 85: dispersion = 1.3%

## Slide 86: dispersion = 2.6%

## Slide 87: dispersion = 5.1%

## Slide 88: spot shape
Ewald sphere
spindle axis
Φ circle
diffracted ray
(h,k,l)
d*
λ’*
λ’*
Ewald sphere
spindle axis
Φ circle
diffracted ray
(h,k,l)
d*
λ*
λ*
spindle axis
Φ circle
diffracted ray
(h,k,l)
d*
Ewald sphere
λ*
λ*
Ewald sphere
spindle axis
Φ circle
diffracted rays
(h,k,l)
d*
d*

## Slide 89
CORRECT – scale all the “factors”

## Slide 90: Polarization factor
1.0
0.45
+

## Slide 91: Lorentz factor
+
50x
1.2x
550x
3x
10x

## Slide 92: The “Lorentz zone”

## Slide 93: A few other corrections…
Meisburger, et. al. Nat Commun 11, 1271 (2020). https://doi.org/10.1038

## Slide 94: attenuation factor
Bouguer, P. (1729). Essai d'optique sur la gradation de la lumière.
Lambert, J. H. (1760). Photometria: sive De mensura et gradibus luminis, colorum et umbrae. E. Klett.
Beer, A. (1852)."Bestimmung der Absorption des rothen Lichts in farbigen Flüssigkeiten", Ann. Phys. Chem 86, 78-90.
A = = exp[-μxtal(txi+ txo)
-μsolvent(tsi + tso)]
IT
Ibeam
μxtal
txi
txo
tsi
tso
txi
txo
tsi
tso
txi
txo
tsi
tso
μsolvent

## Slide 95: B factor from image gazing
B = 500

## Slide 96: B factor from image gazing
B = 20

## Slide 97
CORRECT – after all the “factors”

## Slide 98
xdscc12 – any outlier images?

## Slide 99
xdsstat – radiation damage “limit”

## Slide 100
xdsconv – make an “mtz” file

## Slide 101
xscale – merge multiple runs

## Slide 102
CORRECT – scale all the “factors”

## Slide 103: Re-indexing cheat sheet
Dauter, Z. (1999). "Data-collection strategies." Acta Cryst. D 55(10): 1703-1717.

## Slide 104: Zero-dose extrapolation (ZDE)
Diederichs, K. (2006). "Some aspects of quantitative analysis and correction of radiation damage." Acta Cryst. D 62(1): 96-101.
OUTPUT_FILE= multi-merge.hkl
INPUT_FILE= ../xds_1/XDS_ASCII.HKL
CRYSTAL_NAME= ip
INPUT_FILE= ../xds_2/XDS_ASCII.HKL
! same crystal, but translated
CRYSTAL_NAME= ip

## Slide 105
xscale – final stats

## Slide 106: XDS: CORRECT.LP or XSCALE.LP
SUBSET OF INTENSITY DATA WITH SIGNAL/NOISE >= -3.0 AS FUNCTION OF RESOLUTION
RESOLUTION NUMBER OF REFLECTIONS COMPLETENESS R-FACTOR R-FACTOR COMPARED I/SIGMA R-meas CC(1/2) Anomal SigAno Nano
LIMIT OBSERVED UNIQUE POSSIBLE OF DATA observed expected Corr
6.39 4103 2517 3135 80.3% 1.7% 1.8% 3018 33.77 2.3% 99.9* 21* 1.012 533
4.54 7339 4970 5544 89.6% 2.6% 2.6% 4585 22.56 3.6% 99.8* 9 0.914 796
3.72 9442 6724 7186 93.6% 2.8% 2.9% 5327 19.99 4.0% 99.7* 9 0.859 877
3.22 11077 7997 8486 94.2% 4.8% 4.9% 6094 12.27 6.8% 99.3* -1 0.784 913
2.88 12408 8851 9595 92.2% 10.0% 10.1% 7068 6.01 14.2% 97.8* -2 0.799 1012
2.63 13673 9561 10662 89.7% 20.8% 20.8% 8185 3.10 29.4% 88.8* -4 0.776 1274
2.44 14360 9852 11510 85.6% 34.5% 34.7% 8981 1.90 48.8% 75.9* 2 0.765 1525
2.28 9501 6495 12416 52.3% 61.8% 59.6% 5991 1.14 87.3% 53.5* -2 0.722 1031
2.15 4567 3307 13210 25.0% 120.5% 120.3% 2520 0.59 170.4% 21.9* 4 0.693 408
total 86470 60274 81744 73.7% 6.7% 6.7% 51769 8.97 9.5% 99.5* 2 0.804 8369

## Slide 107: XDS: CORRECT.LP or XSCALE.LP
******************************************************************************
CORRECTION PARAMETERS FOR THE STANDARD ERROR OF REFLECTION INTENSITIES
******************************************************************************
The variance v0(I) of the intensity I obtained from counting statistics is
replaced by v(I)=a*(v0(I)+b*I^2). The model parameters a, b are chosen to
minimize the discrepancies between v(I) and the variance estimated from
sample statistics of symmetry related reflections. This model implicates
an asymptotic limit ISa=1/SQRT(a*b) for the highest I/Sigma(I) that the
experimental setup can produce (Diederichs (2010) Acta Cryst D66, 733-740).
a b ISa
3.806E+00 1.080E-04 49.32

## Slide 108: CCP4: aimless log
$TABLE: Analysis against resolution, XDSdataset:
$GRAPHS:I/sigma, Mean Mn(I)/sd(Mn(I)):0|0.216023x0|137.14:2,13,14:
:Rmerge, Rfull, Rmeas, Rpim v Resolution:0|0.216023x0|1.70834:2,4,5,6,7:
:Average I, RMSdeviation and Sd:0|0.216023x0|1650.8:2,10,11,12:
:Fractional bias:0|0.216023x0|0:2,15:
$$
N 1/d^2 Dmid Rmrg Rfull Rcum Rmeas Rpim Nmeas AvI RMSdev sd I/RMS Mn(I/sd) FrcBias $$ $$
1 0.0064 12.55 0.020 0.020 0.020 0.021 0.006 13115 1651 57 42 29.2 137.1 -
2 0.0191 7.24 0.027 0.027 0.024 0.028 0.008 24753 1171 47 42 25.0 105.2 -
3 0.0318 5.61 0.038 0.038 0.029 0.040 0.012 32197 857 46 43 18.4 79.6 -
4 0.0445 4.74 0.034 0.034 0.031 0.035 0.010 37743 1212 57 53 21.4 91.2 -
5 0.0572 4.18 0.036 0.036 0.032 0.038 0.011 42642 1181 59 57 19.9 83.8 -
6 0.0699 3.78 0.049 0.049 0.036 0.052 0.015 47224 883 59 57 15.1 65.1 -
7 0.0826 3.48 0.065 0.065 0.040 0.068 0.020 51052 685 59 58 11.7 50.9 -
8 0.0953 3.24 0.096 0.096 0.045 0.100 0.029 54636 448 56 56 8.0 35.0 -
9 0.1080 3.04 0.151 0.151 0.050 0.158 0.046 58072 268 53 53 5.1 22.7 -
10 0.1207 2.88 0.229 0.229 0.056 0.240 0.070 60731 171 51 51 3.3 15.4 -
11 0.1334 2.74 0.314 0.314 0.063 0.329 0.097 63807 125 51 51 2.4 11.3 -
12 0.1461 2.62 0.406 0.406 0.070 0.425 0.125 66241 98 51 52 1.9 8.7 -
13 0.1588 2.51 0.537 0.537 0.078 0.562 0.166 68272 76 53 53 1.4 6.5 -
14 0.1715 2.41 0.685 0.685 0.084 0.727 0.237 54170 61 53 54 1.1 4.4 -
15 0.1843 2.33 0.886 0.886 0.088 0.950 0.333 39781 48 55 56 0.9 3.1 -
16 0.1970 2.25 1.249 1.249 0.092 1.355 0.508 29754 35 57 59 0.6 2.1 -

## Slide 109: CCP4: aimless log
$TABLE: Correlations CC(1/2) within dataset, XDSdataset:
$GRAPHS: Anom & Imean CCs v resolution:0|0.216023x0|1:2,4,7:
: RMS correlation ratio :0|0.216023x0|2.20344:2,6:
$$
N 1/d^2 Dmid CCanom Nanom RCRanom CC1/2 NImean $$ $$
1 0.0064 12.55 0.659 499 2.203 1.000 669
2 0.0191 7.24 0.550 975 1.853 1.000 1155
3 0.0318 5.61 0.527 1295 1.798 1.000 1479
16 0.1970 2.25 0.037 2123 1.038 0.711 2275
17 0.2097 2.18 0.043 1682 1.044 0.460 1877

## Slide 110: Summary
XYCORR
INIT
COLSPOT
IDXREF
DEFPIX
INTEGRATE
CORRECT

## Slide 111: 100% SeMet incorporation
Trivial to solve

## Slide 112: 100% S -Met incorporation
Impossible to solve

## Slide 113: Systematic error does not “average out”
Unless you can randomize their source

## Slide 114
Pick-up tool mark
Bragg
glitch
oxygen
inclusions
“tree rings”
1% high
average
1% low
~3x105 photon/pixel → 0.3% error
Pilatus: subtract smooth baseline
Gollwitzer & Krumrey (2014)
J. Appl. Cryst. 47, 378

## Slide 115: Fractional error
mult > (—)2
~3%
<ΔF/F>

## Slide 116
anomalous signal
Crick, F. H. C. & Magdoff, B. S. (1956) Acta Crystallogr. 9, 901-908.
Hendrickson, W. A. & Teeter, M. M. (1981) Nature 290, 107-113.
# sites
MW (Da)
ΔF
F
≈ 1.2 f”
√
World record!
ΔF/F = 0.5%
Wang, Dauter & Dauter (2006) Acta Cryst. D 62, 1475-1483.

## Slide 117
Holton & Frankel (2010) Acta D 66 393-408.

## Slide 118: what the is a MGy?
http://bl831.als.lbl.gov/
damage_rates.pdf
Holton J. M. (2009) J. Synchrotron Rad. 16 133-42
1 MGy = 106 J/kg

## Slide 119
How long will my crystal last?
synch line type flux beamsize flux density dose max xtal min site
ph/s μm ph/μm2/s rate lifetime lifetime
ALS 4.2.2 MAD 2.2e11 75x80 3.7e+07 27.3 kGy/s 18 m 73 s
ALS 5.0.1 mono 1.6e11 100 2.0e+07 10.4 kGy/s 48 m 3.2 m
ALS 5.0.2 MAD 8e11 100 1.0e+08 51.8 kGy/s 9.6 m 39 s
ALS 5.0.3 mono 1.7e11 100 2.2e+07 11 kGy/s 45 m 3 m
ALS 8.2.1 MAD 1.8e11 100 2.3e+07 11.7 kGy/s 43 m 2.9 m
ALS 8.2.2 MAD 2.3e11 100 2.9e+07 14.9 kGy/s 34 m 2.2 m
ALS 8.3.1 MAD 9e11 70 2.3e+08 119 kGy/s 4.2 m 17 s
ALS 8.3.1 typical 6e11 70 1.6e+08 115 kGy/s 4.3 m 17 s
ALS 12.3.1 MAD 1.8e11 100 2.3e+07 11.7 kGy/s 43 m 2.9 m
ALS 12.3.1 ML 4.0e12 100 5.1e+08 513 kGy/s 58 s 3.9 s
APS 8-BM MAD 1e11 200 2.5e+06 1.27 kGy/s 6.6 h 26 m
APS 14-BM-C mono 5.8e10 200 1.4e+06 738 Gy/s 11 h 45 m
APS 14-BM-D MAD 3.3e9 200 8.2e+04 42 Gy/s 8.3 d 13 h
APS 14-ID-B MAD 6.0e10 200 1.5e+06 763 Gy/s 11 h 44 m
APS 17-BM MAD 1.1e11 200 2.8e+06 1.4 kGy/s 6 h 24 m
APS 17-ID MAD 2.3e11 200 5.8e+06 2.93 kGy/s 2.8 h 11 m
APS 19-BM MAD 2.0e11 70x60 4.8e+07 24.2 kGy/s 21 m 83 s
APS 19-ID MAD 1.3e13 80x40 4.1e+09 2.07 MGy/s 15 s 0.97 s
APS 19-ID typical 5.5e11 100x100 5.5e+07 28 kGy/s 18 m 71 s
APS 22-BM MAD 7e12 80x40 2.2e+09 1.23 MGy/s 24 s 1.6 s
APS 22-ID MAD 7e12 80x40 2.2e+09 1.23 MGy/s 24 s 1.6 s
APS 22-ID typical 1.5e12 80 2.3e+08 119 kGy/s 4.2 m 17 s
APS 23-ID-B MAD 1e13 75x25 5.3e+09 3.01 MGy/s 10 s 0.66 s
APS 23-ID typical 1.5e12 80 2.3e+08 119 kGy/s 4.2 m 17 s
APS 24-ID-C MAD 1.3e13 20x60 1.1e+10 5.23 MGy/s 5.7 s 0.38 s
APS 24-ID-E MAD 0.5e13 20x100 2.5e+09 1.19 MGy/s 25 s 1.7 s
APS 31-ID MAD 2e12 70 4.1e+08 194 kGy/s 2.6 m 10 s

## Slide 121
Holton & Frankel (2010) Acta D 66 393-408.

## Slide 122
Holton (2009) J. Synchrotron Rad. 16 133-42
Radiation Damage World Records
MGy reaction reference
~45 global damage Owen et al. (2006)
10/Å global damage Howells et al. (2009)
5 Se-Met Holton (2007)
4 Hg-S Ramagopal et al. (2004)
4 R-C-COOH Garman et al. (2015)
3 S-S Murray et al. (2002)
1 Br-RNA Olieric et al. (2007)
~1? Cl-C ???
0.5 Mn in PS II Yano et al. (2005)
0.06 putidaredoxin Corbett et al. (2007)
0.02 Fe in myoglobin Denisov et al. (2007)

## Slide 123: Life-halving concentration
Na | 44
Mg | 27
P | 9
S | 7
Cl | 5
K | 3.3
Ca | 2.7
Fe | 1.0
Cu | 0.67
Zn | 0.58
molar, at the Se edge based on RADDOSE
As | 0.44
Se | 0.42
Br | 2.9
I | 0.47
Gd | 0.21
Ta | 0.13
Pt | 0.17
Au | 0.16
Hg | 0.16
U | 0.23

## Slide 124: Life-halving concentration
Na | 44
Mg | 27
P | 9
S | 7
Cl | 5
K | 3.3
Ca | 2.7
Fe | 1.0
Cu | 0.67
Zn | 0.58
molar, at the Se edge based on RADDOSE
As | 0.44
Se | 0.42
Br | 2.9
I | 0.47
Gd | 0.21
Ta | 0.13
Pt | 0.17
Au | 0.16
Hg | 0.16
U | 0.23

## Slide 125
The number of photons scattered
before crystal is dead
is independent
of flux & time
1 um3 = 105 photons (roomT)
1 um3 = 106 photons (synch)
1 um3 = 108 photons (XFEL)
Henderson, 1990; Gonzalez & Nave, 1994; Glaeser et al., 2000; Sliz et al., 2003; Leiros et al., 2006; Owen et al., 2006; Garman & McSweeney, 2006; Garman & Nave, 2009; Holton, 2009
(synch)

## Slide 126: Dose slicing
crystal’s useful life
N
photons
N
photons
N
photons
unacceptable
completeness
unacceptable
read noise

## Slide 127: Dose slicing: is that a spot?

## Slide 128: Dose slicing: is that a spot?

## Slide 129: Dose slicing: is that a spot?

## Slide 130: Dose slicing: is that a spot?

## Slide 131: Dose slicing: is that a spot?

## Slide 132: Dose slicing: is that a spot?

## Slide 133: What if?
You only have a few small crystals…
Should you:
  Collect 360° from each?
  Collect 10° from each at 36x exposure?
  Glue 36 xtals together, then collect 360° ?
  Glue, and do 12960° faint exposures?

## Slide 134
Chart: 
Resolution (Å)
CC vs right answer
“true” resolution limit

## Slide 135
Dose slice collection scenario |  |  | 
fine | coarse | glue | Glue & fine
3.17700 | 3.14057 | 3.17604 | 3.12086
“true” resolution vs strategy
301,640,334 photons/xtal

## Slide 136
processing procedure | resolution limit |  |  | 
 | fine | coarse | glue | glue&fine
XDS/XSCALE | 3.17700 | 3.14057 | 3.17604 | 3.12086
XDS/aimless | 3.16251 | 3.15922 | 3.16675 | 3.12162
XDS/noscale | 3.27289 | 3.15617 | 3.38929 | 3.12162
MOSFLM-defaults | x | x | x | x
MOSFLM-cheat | 3.19755 | 3.18166 | 3.20354 | 3.19755
HKL2000-defaults | x | x | x | x
HKL2000-expert | 3.21803 | 3.12646 | 3.11551 | 3.21803
HKL2000-cheat | 3.16119 | 3.10338 | 3.10338 | 3.16300
DIALS-defaults | x | x | 3.42125 | x
DIALS-expert | 3.10296 | 3.09547 | 3.07653 | 3.08465
d*Trek | ? | ? | 4.34076 | ?
EVAL15 | 3.14352 | 3.10383 | 3.09936 | 3.14105
EVAL15-sadabs | 3.13920 | 3.10797 | 3.10171 | 3.13700

## Slide 137: Minimum exposure time:
~1 photon/pixel
  PAD w XDS, DIALS
~10 photon/pixel
PAD w HKL2k, MOSFLM
~30 photon/pixel
CCD detector
adjust exposure
so this is ~15
(Rayonix)

## Slide 138
sample
detector
x-ray beam
scattering

## Slide 139
sample
detector
x-ray beam
scattering

## Slide 140
sample
detector
x-ray beam
scattering

## Slide 141
to source
to detector
d
d∙sin(θ)
θ
atom #1
atom #2
Bragg’s Law
nλ = 2d sin(θ)

## Slide 142
to source
to detector
d
d∙sin(θ)
θ
atom #1
atom #2
Bragg’s Law
nλ = 2d sin(θ)

## Slide 143
sample
x-ray beam
anomalous scattering
detector

## Slide 144
sample
detector
x-ray beam
anomalous scattering

## Slide 145
sample
x-ray beam
anomalous scattering
detector

## Slide 146
sample
detector
x-ray beam
anomalous scattering

## Slide 147
sample
detector
x-ray beam
anomalous scattering

## Slide 148
detector
sample
detector
x-ray beam
scattering

## Slide 151
scattering from a structure
sample
detector
beam center
+

## Slide 152
forward Fourier Transform
inverse Fourier Transform
no phase
+

## Slide 153
scattering from a structure
colored by phase
sample
detector
inverse Fourier Transform
+

## Slide 154
scattering from a lattice
colored by phase
sample
detector
beam center
+

## Slide 155
scattering from a lattice
colored by phase
sample
detector
beam center
+

## Slide 156
scattering from a lattice
colored by phase
sample
detector
beam center
+

## Slide 157
scattering from a crystal structure
colored by phase
sample
detector
beam center
+

## Slide 158: lysozyme: real and reciprocal
+
+

## Slide 159: Diffraction from a 3D lattice
real space
reciprocal space
+

## Slide 160: Diffraction from a 3D lattice
real space
reciprocal space
+

## Slide 161: Expected Rmerge as Iobs → 0

## Slide 162
Average value
Averaging Gaussian error
Number of samples averaged

## Slide 163
Average value
Averaging Gauss/Gauss error
Number of samples averaged

## Slide 164
Average outer-shell Rmerge
Rmerge at the resolution limit in PDB
Year

## Slide 165: Optimum resolution cutoff is:
Too optimistic: add nothing but noise
Too pessimistic: series-termination error
Happy medium?
Simulate:
  Random atoms, compute F2
  Add Gaussian noise, RMS = 1
  Truncate
  Subtract “right” map, RMS difference

## Slide 166
Chart: 
| 0.8 | 0.81 | 0.82 | 0.83 | 0.84 | 0.85 | 0.86 | 0.87 | 0.88 | 0.89 | 0.9 | 0.91 | 0.92 | 0.93 | 0.94 | 0.95 | 0.96 | 0.97 | 0.98 | 0.99 | 1 | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.0900000000000001 | 1.1000000000000001 | 1.1100000000000001 | 1.1200000000000001 | 1.1299999999999999 | 1.1399999999999999 | 1.1499999999999999 | 1.1599999999999999 | 1.17 | 1.18 | 1.19 | 1.2 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.29 | 1.3
total error | 0.0063854 | 0.006377 | 0.0063704 | 0.0063613 | 0.0063546 | 0.0063466 | 0.006339 | 0.0063331 | 0.0063281 | 0.0063218 | 0.0063169 | 0.0063122 | 0.0063076 | 0.0063035 | 0.0062986 | 0.0062952 | 0.006292 | 0.0062882 | 0.0062851 | 0.0062825 | 0.0062808 | 0.0062806 | 0.0062829 | 0.0062871 | 0.0063049 | 0.0063317 | 0.0063768 | 0.006453 | 0.00655 | 0.0067899 | 0.006935 | 0.0072772 | 0.0075335 | 0.0079267 | 0.0082267 | 0.0085917 | 0.0090308 | 0.0096403 | 0.010142 | 0.010578 | 0.011234 | 0.011798 | 0.012663 | 0.013292 | 0.013934 | 0.014784 | 0.015263 | 0.016206 | 0.01711 | 0.01817 | 0.018793
series termination | 0.0001167 | 0.00014118 | 0.00016693 | 0.00020738 | 0.00024052 | 0.00028835 | 0.00034393 | 0.00040287 | 0.00046367 | 0.00054497 | 0.00062312 | 0.00072092 | 0.00082139 | 0.00091787 | 0.0010631 | 0.0011963 | 0.0013389 | 0.0015127 | 0.0017075 | 0.0019087 | 0.0021401 | 0.0023446 | 0.0026241 | 0.0028107 | 0.0031805 | 0.0034765 | 0.0037849 | 0.0041158 | 0.0044212 | 0.0049665 | 0.0052371 | 0.0057965 | 0.0061838 | 0.0067416 | 0.0071474 | 0.0076234 | 0.0081753 | 0.0089132 | 0.0095026 | 0.010003 | 0.010739 | 0.011361 | 0.012297 | 0.012969 | 0.013647 | 0.014539 | 0.015036 | 0.016012 | 0.016942 | 0.018027 | 0.018662
noise in data | 0.0063844 | 0.0063755 | 0.0063682 | 0.0063579 | 0.0063501 | 0.0063401 | 0.0063297 | 0.0063203 | 0.0063111 | 0.0062983 | 0.0062861 | 0.0062709 | 0.0062539 | 0.0062363 | 0.0062082 | 0.0061804 | 0.0061479 | 0.0061034 | 0.0060487 | 0.0059854 | 0.0059049 | 0.0058264 | 0.0057085 | 0.0056237 | 0.0054437 | 0.0052917 | 0.0051318 | 0.0049697 | 0.0048324 | 0.0046297 | 0.0045459 | 0.0043996 | 0.0043026 | 0.0041691 | 0.0040733 | 0.0039623 | 0.0038365 | 0.0036725 | 0.0035436 | 0.0034402 | 0.0032968 | 0.0031835 | 0.0030235 | 0.002914 | 0.0028109 | 0.0026816 | 0.0026168 | 0.0024966 | 0.0023891 | 0.0022725 | 0.0022113
Resolution cutoff (Å)
Error in map (rms e-)
Optimal resolution cutoff

## Slide 167
Chart: 
| 0.8 | 0.81 | 0.82 | 0.83 | 0.84 | 0.85 | 0.86 | 0.87 | 0.88 | 0.89 | 0.9 | 0.91 | 0.92 | 0.93 | 0.94 | 0.95 | 0.96 | 0.97 | 0.98 | 0.99 | 1 | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.0900000000000001 | 1.1000000000000001 | 1.1100000000000001 | 1.1200000000000001 | 1.1299999999999999 | 1.1399999999999999 | 1.1499999999999999 | 1.1599999999999999 | 1.17 | 1.18 | 1.19 | 1.2 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.29 | 1.3
total error | 0.011234 | 0.011798 | 0.012663 | 0.013292 | 0.013934 | 0.014784 | 0.015263 | 0.016206 | 0.01711 | 0.01817 | 0.018793
CC1/2 | 0.00027543 | 0.0009316 | 0.0011407 | 0.00093939 | -0.00085547 | 8.519e-05 | 0.0017707 | 0.0018655 | 0.00044742 | -0.00083644 | -8.9753e-05 | -0.0012856 | -0.0022441 | -0.0011232 | -0.0006217 | -0.00063863 | 0.00039935 | 0.0020736 | 0.0023749 | 0.001991 | 0.0015254 | 0.0011874 | 0.0012966 | 0.0070442 | 0.010659 | 0.014306 | 0.021539 | 0.031659 | 0.044804 | 0.054608 | 0.085054 | 0.10512 | 0.15123 | 0.19219 | 0.24658 | 0.30221 | 0.374 | 0.43569 | 0.50225 | 0.58077 | 0.63948 | 0.71345 | 0.75573 | 0.79925 | 0.8353 | 0.86091 | 0.89448 | 0.91138 | 0.92903 | 0.94132 | 0.95425
CCtrue | -0.0018739 | -0.0009566 | 0.00078927 | -0.00047161 | 0.00086676 | 0.0013379 | 0.00021407 | 0.001888 | 0.0027656 | 0.0036656 | 0.0050632 | 0.0063693 | 0.0067999 | 0.010765 | 0.013871 | 0.01712 | 0.023754 | 0.030523 | 0.038896 | 0.047732 | 0.06188 | 0.078774 | 0.090571 | 0.11775 | 0.13935 | 0.1692 | 0.20286 | 0.24067 | 0.29172 | 0.32687 | 0.40014 | 0.43863 | 0.51218 | 0.56571 | 0.62646 | 0.68068 | 0.73761 | 0.77832 | 0.81873 | 0.8591 | 0.88409 | 0.9127 | 0.92782 | 0.94254 | 0.95429 | 0.96199 | 0.97176 | 0.97663 | 0.98146 | 0.98479 | 0.98825
Resolution cutoff (Å)
Correlation Coefficient
Optimal resolution cutoff

## Slide 168
Chart: 
| 0.8 | 0.81 | 0.82 | 0.83 | 0.84 | 0.85 | 0.86 | 0.87 | 0.88 | 0.89 | 0.9 | 0.91 | 0.92 | 0.93 | 0.94 | 0.95 | 0.96 | 0.97 | 0.98 | 0.99 | 1 | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.0900000000000001 | 1.1000000000000001 | 1.1100000000000001 | 1.1200000000000001 | 1.1299999999999999 | 1.1399999999999999 | 1.1499999999999999 | 1.1599999999999999 | 1.17 | 1.18 | 1.19 | 1.2 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.29 | 1.3
total error | 0.011234 | 0.011798 | 0.012663 | 0.013292 | 0.013934 | 0.014784 | 0.015263 | 0.016206 | 0.01711 | 0.01817 | 0.018793
CC1/2 | 0.00027543 | 0.0009316 | 0.0011407 | 0.00093939 | -0.00085547 | 8.519e-05 | 0.0017707 | 0.0018655 | 0.00044742 | -0.00083644 | -8.9753e-05 | -0.0012856 | -0.0022441 | -0.0011232 | -0.0006217 | -0.00063863 | 0.00039935 | 0.0020736 | 0.0023749 | 0.001991 | 0.0015254 | 0.0011874 | 0.0012966 | 0.0070442 | 0.010659 | 0.014306 | 0.021539 | 0.031659 | 0.044804 | 0.054608 | 0.085054 | 0.10512 | 0.15123 | 0.19219 | 0.24658 | 0.30221 | 0.374 | 0.43569 | 0.50225 | 0.58077 | 0.63948 | 0.71345 | 0.75573 | 0.79925 | 0.8353 | 0.86091 | 0.89448 | 0.91138 | 0.92903 | 0.94132 | 0.95425
CCtrue | -0.0018739 | -0.0009566 | 0.00078927 | -0.00047161 | 0.00086676 | 0.0013379 | 0.00021407 | 0.001888 | 0.0027656 | 0.0036656 | 0.0050632 | 0.0063693 | 0.0067999 | 0.010765 | 0.013871 | 0.01712 | 0.023754 | 0.030523 | 0.038896 | 0.047732 | 0.06188 | 0.078774 | 0.090571 | 0.11775 | 0.13935 | 0.1692 | 0.20286 | 0.24067 | 0.29172 | 0.32687 | 0.40014 | 0.43863 | 0.51218 | 0.56571 | 0.62646 | 0.68068 | 0.73761 | 0.77832 | 0.81873 | 0.8591 | 0.88409 | 0.9127 | 0.92782 | 0.94254 | 0.95429 | 0.96199 | 0.97176 | 0.97663 | 0.98146 | 0.98479 | 0.98825
CC* | 0.023467 | 0.043145 | 0.047737 | 0.043325 | 0 | 0.013052 | 0.059457 | 0.061026 | 0.029907 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.028256 | 0.064333 | 0.068837 | 0.063041 | 0.055192 | 0.048703 | 0.05089 | 0.11828 | 0.14523 | 0.16795 | 0.20535 | 0.24774 | 0.29286 | 0.32181 | 0.39595 | 0.43617 | 0.51257 | 0.56782 | 0.62898 | 0.68128 | 0.73783 | 0.77906 | 0.81772 | 0.8572 | 0.88323 | 0.91256 | 0.92783 | 0.94256 | 0.95408 | 0.9619 | 0.97175 | 0.97654 | 0.98143 | 0.98477 | 0.98823
Resolution cutoff (Å)
Correlation Coefficient
Optimal resolution cutoff

## Slide 169
Chart: 
log( intensity )
Wilson Plot
2.2 1.6 1.3 1.1 1.0 0.9 0.85 0.8 0.75 0.7
resolution (Å)
8
2

## Slide 170: Optimum resolution cutoff is:
0.0 Å