## Slide 1: Acknowledgements
UCSF LBNL SLAC
ALS 8.3.1: TomAlberTron
NIH NIGMS R01 GM124149 (Holton)
NIH NIAID P50 AI150476 (Krogan)
NIH NIGMS P30 GM124169 (Adams)
DOE-BER IDAT (Hura)
NIH NIGMS P30 GM133894 (Hodgson)
The Advanced Light Source is supported by the Director, Office of Science, Office of Basic Energy Sciences, Materials Sciences Division, of the US Department of Energy under contract No. DE-AC02-05CH11231 at Lawrence Berkeley National Laboratory.
Robert Stroud James Fraser Nevan Krogan
Paul Adams John Tainer Greg Hura Nick Sauter
Aina Cohen Art Lyubimov Jeney Wierman
https://bl831.als.lbl.gov/~jamesh/powerpoint/RapiDataXDStutorial_2024.pptx

## Slide 2: Useful links
Fear no Space Groups
phasertng

## Slide 3: XDS X-ray diffraction data reduction engine
Diffraction image data
List of spot intensities (text)

## Slide 4
XDS native user interface: a text file!

## Slide 6: A better image viewer

## Slide 7
XDS native user interface

## Slide 8
The “prompt”
Cursor
does not move
with mouse
Mouse
can be used
To select and paste
But all pasted text
Goes to the cursor

## Slide 9: too… much… typing!
1) Make sure X11 is running
2) Select some text with mouse
3) Move mouse over to your terminal
4) Hit middle mouse button
5) No keyboard required!

## Slide 10: How to start XDS
smbnxs:~> ssh pxproc27
pxproc:~> cd /data/$USER
pxproc:~> mkdir xds_tutorial
pxproc:~> cd xds_tutorial
pxproc:~> ln –sf /data/holton/demo .
pxproc:~> xdsgui
Space!
underscore
prompt
Do not type!
Enter
Space!
Space!
period!

## Slide 11
xdsgui – a graphical interface for XDS

## Slide 12
xdsgui – load an image
/data/holton/demo/1rb5_1_00001.cbf

## Slide 13
xdsgui – a GUI for text!
/data/holton/demo/1rb5_1_?????.cbf

## Slide 14
XYCORR– setting things up

## Slide 15
INIT – look at the background

## Slide 16
INIT – look at the background

## Slide 17
COLSPOT – the spot picker

## Slide 18
COLSPOT – the spot picker

## Slide 19
IDXREF – auto-indexing

## Slide 20: XPARM.XDS
XPARM.XDS VERSION Feb 5, 2021 BUILT=20210323
1 270.0000 0.2000 0.999971 0.007404 -0.001831
0.979460 0.000711 0.000291 1.020970
96 78.8215 78.8215 37.1784 90.000 90.000 90.000
36.338257 9.327419 -69.320709
26.357738 70.532730 23.307367
30.559622 -16.002089 13.866352
1 2463 2527 0.172000 0.172000
1231.500000 1263.500000 189.160004
1.000000 0.000000 0.000000
0.000000 1.000000 0.000000
0.000000 0.000000 1.000000
1 1 2463 1 2527
0.00 0.00 0.00 1.00000 0.00000 0.00000 0.00000 1.00000 0.00000
First image number
Starting phi
Delta-phi
spindle vector
beam vector
unit cell
a
b
c
space
group
number
wavelen
cell vectors
real space
det panels
origin
distance
pixels
pixel size
x
y
z
Detector fast/slow
Panel #

## Slide 21
IDXREF – auto-indexing
BRAVAIS- POSSIBLE SPACE-GROUPS FOR PROTEIN CRYSTALS
TYPE [SPACE GROUP NUMBER,SYMBOL]
aP [1,P1]
mP [3,P2] [4,P2(1)]
mC,mI [5,C2]
oP [16,P222] [17,P222(1)] [18,P2(1)2(1)2] [19,P2(1)2(1)2(1)]
oC [21,C222] [20,C222(1)]
oF [22,F222]
oI [23,I222] [24,I2(1)2(1)2(1)]
tP [75,P4] [76,P4(1)] [77,P4(2)] [78,P4(3)] [89,P422] [90,P42(1)2]
[91,P4(1)22] [92,P4(1)2(1)2] [93,P4(2)22] [94,P4(2)2(1)2]
[95,P4(3)22] [96,P4(3)2(1)2]
tI [79,I4] [80,I4(1)] [97,I422] [98,I4(1)22]
hP [143,P3] [144,P3(1)] [145,P3(2)] [149,P312] [150,P321] [151,P3(1)12]
[152,P3(1)21] [153,P3(2)12] [154,P3(2)21] [168,P6] [169,P6(1)]
[170,P6(5)] [171,P6(2)] [172,P6(4)] [173,P6(3)] [177,P622]
[178,P6(1)22] [179,P6(5)22] [180,P6(2)22] [181,P6(4)22] [182,P6(3)22]
hR [146,R3] [155,R32]
cP [195,P23] [198,P2(1)3] [207,P432] [208,P4(2)32] [212,P4(3)32]
[213,P4(1)32]
cF [196,F23] [209,F432] [210,F4(1)32]
cI [197,I23] [199,I2(1)3] [211,I432] [214,I4(1)32]

## Slide 22
INTEGRATE – measure the spots!

## Slide 23
CORRECT – scale all the “factors”

## Slide 24: Beam center vs “origin”
origin
direct beam

## Slide 25
xdsconv – make an “mtz” file
output_file_name.mtz

## Slide 26: Adding noise
12 + 12 = 1.42
32 + 12 = 3.22
σtotal2 = σ12 + σ22

## Slide 27: Adding noise
12 + 12 = 1.42
32 + 12 = 3.22
σtotal2 = σ12 + σ22

## Slide 28: Adding noise
12 + 12 = 1.42
32 + 12 = 3.22
102 + 12 = 10.052

## Slide 29: signal vs noise
“If you don’t have good data, then you must learn statistics.”
-James Holton

## Slide 30: Sources of error for anomalous differences
Shutter jitter (rms 0.5 ms)
Beam flicker (0.15%/sqrt(Hz))
Attenuation correction (< 2%)
Radiation damage (1%/MGy)
Detector calibration (3%)

## Slide 31
anomalous signal
Crick, F. H. C. & Magdoff, B. S. (1956) Acta Crystallogr. 9, 901-908.
Hendrickson, W. A. & Teeter, M. M. (1981) Nature 290, 107-113.
# sites
MW (Da)
ΔF
F
≈ 1.2 f”
√

## Slide 32
Holton & Frankel (2010) Acta D 66 393-408.

## Slide 33
xdsconv – make an “mtz” file

## Slide 34
xscale – merge multiple runs

## Slide 35: Re-indexing cheat sheet
Dauter, Z. (1999). "Data-collection strategies." Acta Cryst. D 55(10): 1703-1717.

## Slide 36: Zero-dose extrapolation (ZDE)
Diederichs, K. (2006). "Some aspects of quantitative analysis and correction of radiation damage." Acta Cryst. D 62(1): 96-101.
OUTPUT_FILE= multi-merge.hkl
INPUT_FILE= ../xds_1/XDS_ASCII.HKL
CRYSTAL_NAME= ip
INPUT_FILE= ../xds_2/XDS_ASCII.HKL
! same crystal, but translated
CRYSTAL_NAME= ip

## Slide 37
xscale – final stats

## Slide 38: XDS: CORRECT.LP or XSCALE.LP
SUBSET OF INTENSITY DATA WITH SIGNAL/NOISE >= -3.0 AS FUNCTION OF RESOLUTION
RESOLUTION NUMBER OF REFLECTIONS COMPLETENESS R-FACTOR R-FACTOR COMPARED I/SIGMA R-meas CC(1/2) Anomal SigAno Nano
LIMIT OBSERVED UNIQUE POSSIBLE OF DATA observed expected Corr
6.39 4103 2517 3135 80.3% 1.7% 1.8% 3018 33.77 2.3% 99.9* 21* 1.012 533
4.54 7339 4970 5544 89.6% 2.6% 2.6% 4585 22.56 3.6% 99.8* 9 0.914 796
3.72 9442 6724 7186 93.6% 2.8% 2.9% 5327 19.99 4.0% 99.7* 9 0.859 877
3.22 11077 7997 8486 94.2% 4.8% 4.9% 6094 12.27 6.8% 99.3* -1 0.784 913
2.88 12408 8851 9595 92.2% 10.0% 10.1% 7068 6.01 14.2% 97.8* -2 0.799 1012
2.63 13673 9561 10662 89.7% 20.8% 20.8% 8185 3.10 29.4% 88.8* -4 0.776 1274
2.44 14360 9852 11510 85.6% 34.5% 34.7% 8981 1.90 48.8% 75.9* 2 0.765 1525
2.28 9501 6495 12416 52.3% 61.8% 59.6% 5991 1.14 87.3% 53.5* -2 0.722 1031
2.15 4567 3307 13210 25.0% 120.5% 120.3% 2520 0.59 170.4% 21.9* 4 0.693 408
total 86470 60274 81744 73.7% 6.7% 6.7% 51769 8.97 9.5% 99.5* 2 0.804 8369

## Slide 39: XDS: CORRECT.LP or XSCALE.LP
******************************************************************************
CORRECTION PARAMETERS FOR THE STANDARD ERROR OF REFLECTION INTENSITIES
******************************************************************************
The variance v0(I) of the intensity I obtained from counting statistics is
replaced by v(I)=a*(v0(I)+b*I^2). The model parameters a, b are chosen to
minimize the discrepancies between v(I) and the variance estimated from
sample statistics of symmetry related reflections. This model implicates
an asymptotic limit ISa=1/SQRT(a*b) for the highest I/Sigma(I) that the
experimental setup can produce (Diederichs (2010) Acta Cryst D66, 733-740).
a b ISa
3.806E+00 1.080E-04 49.32

## Slide 40: CCP4: aimless log
$TABLE: Analysis against resolution, XDSdataset:
$GRAPHS:I/sigma, Mean Mn(I)/sd(Mn(I)):0|0.216023x0|137.14:2,13,14:
:Rmerge, Rfull, Rmeas, Rpim v Resolution:0|0.216023x0|1.70834:2,4,5,6,7:
:Average I, RMSdeviation and Sd:0|0.216023x0|1650.8:2,10,11,12:
:Fractional bias:0|0.216023x0|0:2,15:
$$
N 1/d^2 Dmid Rmrg Rfull Rcum Rmeas Rpim Nmeas AvI RMSdev sd I/RMS Mn(I/sd) FrcBias $$ $$
1 0.0064 12.55 0.020 0.020 0.020 0.021 0.006 13115 1651 57 42 29.2 137.1 -
2 0.0191 7.24 0.027 0.027 0.024 0.028 0.008 24753 1171 47 42 25.0 105.2 -
3 0.0318 5.61 0.038 0.038 0.029 0.040 0.012 32197 857 46 43 18.4 79.6 -
4 0.0445 4.74 0.034 0.034 0.031 0.035 0.010 37743 1212 57 53 21.4 91.2 -
5 0.0572 4.18 0.036 0.036 0.032 0.038 0.011 42642 1181 59 57 19.9 83.8 -
6 0.0699 3.78 0.049 0.049 0.036 0.052 0.015 47224 883 59 57 15.1 65.1 -
7 0.0826 3.48 0.065 0.065 0.040 0.068 0.020 51052 685 59 58 11.7 50.9 -
8 0.0953 3.24 0.096 0.096 0.045 0.100 0.029 54636 448 56 56 8.0 35.0 -
9 0.1080 3.04 0.151 0.151 0.050 0.158 0.046 58072 268 53 53 5.1 22.7 -
10 0.1207 2.88 0.229 0.229 0.056 0.240 0.070 60731 171 51 51 3.3 15.4 -
11 0.1334 2.74 0.314 0.314 0.063 0.329 0.097 63807 125 51 51 2.4 11.3 -
12 0.1461 2.62 0.406 0.406 0.070 0.425 0.125 66241 98 51 52 1.9 8.7 -
13 0.1588 2.51 0.537 0.537 0.078 0.562 0.166 68272 76 53 53 1.4 6.5 -
14 0.1715 2.41 0.685 0.685 0.084 0.727 0.237 54170 61 53 54 1.1 4.4 -
15 0.1843 2.33 0.886 0.886 0.088 0.950 0.333 39781 48 55 56 0.9 3.1 -
16 0.1970 2.25 1.249 1.249 0.092 1.355 0.508 29754 35 57 59 0.6 2.1 -

## Slide 41: CCP4: aimless log
$TABLE: Correlations CC(1/2) within dataset, XDSdataset:
$GRAPHS: Anom & Imean CCs v resolution:0|0.216023x0|1:2,4,7:
: RMS correlation ratio :0|0.216023x0|2.20344:2,6:
$$
N 1/d^2 Dmid CCanom Nanom RCRanom CC1/2 NImean $$ $$
1 0.0064 12.55 0.659 499 2.203 1.000 669
2 0.0191 7.24 0.550 975 1.853 1.000 1155
3 0.0318 5.61 0.527 1295 1.798 1.000 1479
16 0.1970 2.25 0.037 2123 1.038 0.711 2275
17 0.2097 2.18 0.043 1682 1.044 0.460 1877

## Slide 42: Resolution cutoff?

## Slide 43
0.8
0.7
0.6
0.5
0.4
0.3
0.2
0.1
0
Average outer-shell Rmerge
Year
Rmerge at the resolution limit in PDB

## Slide 44: Expected Rmerge as Iobs → 0

## Slide 45: Optimum resolution cutoff is:
Too optimistic: add nothing but noise
Too pessimistic: series-termination error
Happy medium?
Simulate:
  Random atoms, compute F2
  Add Gaussian noise, RMS = 1
  Truncate
  Subtract “right” map, RMS difference

## Slide 46
Chart: 
| 0.8 | 0.81 | 0.82 | 0.83 | 0.84 | 0.85 | 0.86 | 0.87 | 0.88 | 0.89 | 0.9 | 0.91 | 0.92 | 0.93 | 0.94 | 0.95 | 0.96 | 0.97 | 0.98 | 0.99 | 1 | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.0900000000000001 | 1.1000000000000001 | 1.1100000000000001 | 1.1200000000000001 | 1.1299999999999999 | 1.1399999999999999 | 1.1499999999999999 | 1.1599999999999999 | 1.17 | 1.18 | 1.19 | 1.2 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.29 | 1.3
total error | 0.0063854 | 0.006377 | 0.0063704 | 0.0063613 | 0.0063546 | 0.0063466 | 0.006339 | 0.0063331 | 0.0063281 | 0.0063218 | 0.0063169 | 0.0063122 | 0.0063076 | 0.0063035 | 0.0062986 | 0.0062952 | 0.006292 | 0.0062882 | 0.0062851 | 0.0062825 | 0.0062808 | 0.0062806 | 0.0062829 | 0.0062871 | 0.0063049 | 0.0063317 | 0.0063768 | 0.006453 | 0.00655 | 0.0067899 | 0.006935 | 0.0072772 | 0.0075335 | 0.0079267 | 0.0082267 | 0.0085917 | 0.0090308 | 0.0096403 | 0.010142 | 0.010578 | 0.011234 | 0.011798 | 0.012663 | 0.013292 | 0.013934 | 0.014784 | 0.015263 | 0.016206 | 0.01711 | 0.01817 | 0.018793
series termination | 0.0001167 | 0.00014118 | 0.00016693 | 0.00020738 | 0.00024052 | 0.00028835 | 0.00034393 | 0.00040287 | 0.00046367 | 0.00054497 | 0.00062312 | 0.00072092 | 0.00082139 | 0.00091787 | 0.0010631 | 0.0011963 | 0.0013389 | 0.0015127 | 0.0017075 | 0.0019087 | 0.0021401 | 0.0023446 | 0.0026241 | 0.0028107 | 0.0031805 | 0.0034765 | 0.0037849 | 0.0041158 | 0.0044212 | 0.0049665 | 0.0052371 | 0.0057965 | 0.0061838 | 0.0067416 | 0.0071474 | 0.0076234 | 0.0081753 | 0.0089132 | 0.0095026 | 0.010003 | 0.010739 | 0.011361 | 0.012297 | 0.012969 | 0.013647 | 0.014539 | 0.015036 | 0.016012 | 0.016942 | 0.018027 | 0.018662
noise in data | 0.0063844 | 0.0063755 | 0.0063682 | 0.0063579 | 0.0063501 | 0.0063401 | 0.0063297 | 0.0063203 | 0.0063111 | 0.0062983 | 0.0062861 | 0.0062709 | 0.0062539 | 0.0062363 | 0.0062082 | 0.0061804 | 0.0061479 | 0.0061034 | 0.0060487 | 0.0059854 | 0.0059049 | 0.0058264 | 0.0057085 | 0.0056237 | 0.0054437 | 0.0052917 | 0.0051318 | 0.0049697 | 0.0048324 | 0.0046297 | 0.0045459 | 0.0043996 | 0.0043026 | 0.0041691 | 0.0040733 | 0.0039623 | 0.0038365 | 0.0036725 | 0.0035436 | 0.0034402 | 0.0032968 | 0.0031835 | 0.0030235 | 0.002914 | 0.0028109 | 0.0026816 | 0.0026168 | 0.0024966 | 0.0023891 | 0.0022725 | 0.0022113
Resolution cutoff (Å)
Error in map (rms e-)
Optimal resolution cutoff

## Slide 47
Chart: 
| 0.8 | 0.81 | 0.82 | 0.83 | 0.84 | 0.85 | 0.86 | 0.87 | 0.88 | 0.89 | 0.9 | 0.91 | 0.92 | 0.93 | 0.94 | 0.95 | 0.96 | 0.97 | 0.98 | 0.99 | 1 | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.0900000000000001 | 1.1000000000000001 | 1.1100000000000001 | 1.1200000000000001 | 1.1299999999999999 | 1.1399999999999999 | 1.1499999999999999 | 1.1599999999999999 | 1.17 | 1.18 | 1.19 | 1.2 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.29 | 1.3
total error | 0.011234 | 0.011798 | 0.012663 | 0.013292 | 0.013934 | 0.014784 | 0.015263 | 0.016206 | 0.01711 | 0.01817 | 0.018793
CC1/2 | 0.00027543 | 0.0009316 | 0.0011407 | 0.00093939 | -0.00085547 | 8.519e-05 | 0.0017707 | 0.0018655 | 0.00044742 | -0.00083644 | -8.9753e-05 | -0.0012856 | -0.0022441 | -0.0011232 | -0.0006217 | -0.00063863 | 0.00039935 | 0.0020736 | 0.0023749 | 0.001991 | 0.0015254 | 0.0011874 | 0.0012966 | 0.0070442 | 0.010659 | 0.014306 | 0.021539 | 0.031659 | 0.044804 | 0.054608 | 0.085054 | 0.10512 | 0.15123 | 0.19219 | 0.24658 | 0.30221 | 0.374 | 0.43569 | 0.50225 | 0.58077 | 0.63948 | 0.71345 | 0.75573 | 0.79925 | 0.8353 | 0.86091 | 0.89448 | 0.91138 | 0.92903 | 0.94132 | 0.95425
CCtrue | -0.0018739 | -0.0009566 | 0.00078927 | -0.00047161 | 0.00086676 | 0.0013379 | 0.00021407 | 0.001888 | 0.0027656 | 0.0036656 | 0.0050632 | 0.0063693 | 0.0067999 | 0.010765 | 0.013871 | 0.01712 | 0.023754 | 0.030523 | 0.038896 | 0.047732 | 0.06188 | 0.078774 | 0.090571 | 0.11775 | 0.13935 | 0.1692 | 0.20286 | 0.24067 | 0.29172 | 0.32687 | 0.40014 | 0.43863 | 0.51218 | 0.56571 | 0.62646 | 0.68068 | 0.73761 | 0.77832 | 0.81873 | 0.8591 | 0.88409 | 0.9127 | 0.92782 | 0.94254 | 0.95429 | 0.96199 | 0.97176 | 0.97663 | 0.98146 | 0.98479 | 0.98825
Resolution cutoff (Å)
Correlation Coefficient
Optimal resolution cutoff

## Slide 48
Chart: 
| 0.8 | 0.81 | 0.82 | 0.83 | 0.84 | 0.85 | 0.86 | 0.87 | 0.88 | 0.89 | 0.9 | 0.91 | 0.92 | 0.93 | 0.94 | 0.95 | 0.96 | 0.97 | 0.98 | 0.99 | 1 | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.0900000000000001 | 1.1000000000000001 | 1.1100000000000001 | 1.1200000000000001 | 1.1299999999999999 | 1.1399999999999999 | 1.1499999999999999 | 1.1599999999999999 | 1.17 | 1.18 | 1.19 | 1.2 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.29 | 1.3
total error | 0.011234 | 0.011798 | 0.012663 | 0.013292 | 0.013934 | 0.014784 | 0.015263 | 0.016206 | 0.01711 | 0.01817 | 0.018793
CC1/2 | 0.00027543 | 0.0009316 | 0.0011407 | 0.00093939 | -0.00085547 | 8.519e-05 | 0.0017707 | 0.0018655 | 0.00044742 | -0.00083644 | -8.9753e-05 | -0.0012856 | -0.0022441 | -0.0011232 | -0.0006217 | -0.00063863 | 0.00039935 | 0.0020736 | 0.0023749 | 0.001991 | 0.0015254 | 0.0011874 | 0.0012966 | 0.0070442 | 0.010659 | 0.014306 | 0.021539 | 0.031659 | 0.044804 | 0.054608 | 0.085054 | 0.10512 | 0.15123 | 0.19219 | 0.24658 | 0.30221 | 0.374 | 0.43569 | 0.50225 | 0.58077 | 0.63948 | 0.71345 | 0.75573 | 0.79925 | 0.8353 | 0.86091 | 0.89448 | 0.91138 | 0.92903 | 0.94132 | 0.95425
CCtrue | -0.0018739 | -0.0009566 | 0.00078927 | -0.00047161 | 0.00086676 | 0.0013379 | 0.00021407 | 0.001888 | 0.0027656 | 0.0036656 | 0.0050632 | 0.0063693 | 0.0067999 | 0.010765 | 0.013871 | 0.01712 | 0.023754 | 0.030523 | 0.038896 | 0.047732 | 0.06188 | 0.078774 | 0.090571 | 0.11775 | 0.13935 | 0.1692 | 0.20286 | 0.24067 | 0.29172 | 0.32687 | 0.40014 | 0.43863 | 0.51218 | 0.56571 | 0.62646 | 0.68068 | 0.73761 | 0.77832 | 0.81873 | 0.8591 | 0.88409 | 0.9127 | 0.92782 | 0.94254 | 0.95429 | 0.96199 | 0.97176 | 0.97663 | 0.98146 | 0.98479 | 0.98825
CC* | 0.023467 | 0.043145 | 0.047737 | 0.043325 | 0 | 0.013052 | 0.059457 | 0.061026 | 0.029907 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.028256 | 0.064333 | 0.068837 | 0.063041 | 0.055192 | 0.048703 | 0.05089 | 0.11828 | 0.14523 | 0.16795 | 0.20535 | 0.24774 | 0.29286 | 0.32181 | 0.39595 | 0.43617 | 0.51257 | 0.56782 | 0.62898 | 0.68128 | 0.73783 | 0.77906 | 0.81772 | 0.8572 | 0.88323 | 0.91256 | 0.92783 | 0.94256 | 0.95408 | 0.9619 | 0.97175 | 0.97654 | 0.98143 | 0.98477 | 0.98823
Resolution cutoff (Å)
Correlation Coefficient
Optimal resolution cutoff

## Slide 49
log( intensity )
Wilson Plot
2.2 1.6 1.3 1.1 1.0 0.9 0.85 0.8 0.75 0.7
resolution (Å)
8
2

## Slide 50: Optimum resolution cutoff is:
0.0 Å

## Slide 51
detector
sample
detector
x-ray beam
scattering

## Slide 53
sample
detector
x-ray beam
scattering

## Slide 54
sample
detector
x-ray beam
scattering

## Slide 55
sample
detector
x-ray beam
scattering

## Slide 56
to source
to detector
d
d∙sin(θ)
θ
atom #1
atom #2
Bragg’s Law
nλ = 2d sin(θ)

## Slide 57
to source
to detector
d
d∙sin(θ)
θ
atom #1
atom #2
Bragg’s Law
nλ = 2d sin(θ)

## Slide 59
scattering from a structure
sample
detector
direct beam
+

## Slide 60
forward Fourier Transform
inverse Fourier Transform
no phase

## Slide 61
scattering from a structure
colored by phase
sample
detector
inverse Fourier Transform

## Slide 62
scattering from a lattice
colored by phase
sample
detector
beam center
+

## Slide 63
scattering from a lattice
colored by phase
sample
detector

## Slide 64
scattering from a lattice
colored by phase
sample
detector

## Slide 65
scattering from a crystal structure
colored by phase
sample
detector
beam center
+

## Slide 66: Major Phasing techniques
Molecular Replacement
Multiple Isomorphous Replacement
Multiwavelength Anomalous Diffraction
Single-wavelength Anomalous Diffraction

## Slide 67
inverse Fourier Transform
no phase

## Slide 68
inverse Fourier Transform
no phase

## Slide 69: Major Phasing techniques
Molecular Replacement
Multiple Isomorphous Replacement
Multiwavelength Anomalous Diffraction
Single-wavelength Anomalous Diffraction

## Slide 70
sample
x-ray beam
anomalous scattering
detector

## Slide 71
sample
detector
x-ray beam
anomalous scattering

## Slide 72
sample
x-ray beam
anomalous scattering
detector

## Slide 73
sample
detector
x-ray beam
anomalous scattering

## Slide 74
sample
detector
x-ray beam
anomalous scattering