## Slide 1: Where do the X-rays come from?

## Slide 2: X-rays (aka “Röntgenstrahlung”)
Röntgen, W. C. (1898)."Ueber eine neue Art von Strahlen", Annalen der Physik 300, 1-11.
The discovery soon after this of X-rays by Röntgen (22), the first investigator to use the type of tube described above, is generally considered to be a good example of a lucky discovery.
P.E.A. von Lenard
1906 Nobel Lecture

## Slide 3: Electricity + vacuum = X-rays
1654 ~10 Torr
Otto von Guericke’s
“Magdeburg hemispheres”
Benvenuti, C. & Chiggiato, P. (1993). "Obtention of pressures in the 10−14 torr range by means of a Zr-V-Fe non evaporable getter", Vacuum 44, 511-513.
1993 ~10-14 Torr
1865 < 10-2 Torr
Hg “fall tube” pump
1870s < 10-3 Torr
1894 < 10-5 Torr
Sprengel pump
Sprengel, H. (1865). "Researches on the vacuum", J. Chem. Soc. 18, 9b-21.
Redhead, Paul A., ed. Vacuum Science and Technology: Pioneers of the 20th Century. Vol. 2. Springer, 1994.

## Slide 4: Fluorescent X-ray sources: “sealed tube” and “rotating anode”

## Slide 5: EXPERIMENTAL SETUP – Home Source

## Slide 6: Electric charge
balloon
Wool Sweater
-
-
-
-
-
-
+ + + + + +

## Slide 7: Electric field of a moving charge

## Slide 8: Electric field of a moving charge
accelerating charges
make
electromagnetic waves
(light)

## Slide 9: Bending Magnet
N
electron beam
X-rays !
S

## Slide 10: Two Bending Magnets
N
electron beam
S
S
N
2x X-rays !

## Slide 11: Wiggler
N
S
N
S
N
S
N
S
S
N
S
N
S
N
S
N
8x X-rays !

## Slide 12: Undulator
N
S
N
S
N
S
N
S
S
N
S
N
S
N
S
N
> 8x X-rays !

## Slide 13: Undulator emission spectrum

## Slide 14: LCLS undulator hall

## Slide 15: LCLS at SLAC

## Slide 16: Electric field of a moving charge

## Slide 17: SASE effect

## Slide 18: Self-Amplified Stimulated Emission (SASE) effect

## Slide 19: Accelerator-based light source terminology
Code name Translation
First Generation Bending Magnet
Second Generation Wiggler
Third Generation Undulator
Fourth Generation Free electron laser

## Slide 20
Holton (2009) J. Synchrotron Rad. 16 133-42
Radiation Damage World Records
MGy reaction reference
~45 global damage Owen et al. (2006)
10/Å global damage Howells et al. (2009)
5 Se-Met Holton (2007)
4 Hg-S Ramagopal et al. (2004)
4 R-C-COOH Garman et al. (2015)
3 S-S Murray et al. (2002)
1 Br-RNA Olieric et al. (2007)
~1? Cl-C ???
0.5 Mn in PS II Yano et al. (2005)
0.06 putidaredoxin Corbett et al. (2007)
0.02 Fe in myoglobin Denisov et al. (2007)

## Slide 21: what the is a MGy?
http://bl831.als.lbl.gov/
damage_rates.pdf
Holton J. M. (2009) J. Synchrotron Rad. 16 133-42

## Slide 22
How long will my crystal last?
synch line type flux beamsize flux density dose max xtal min site
ph/s μm ph/μm2/s rate lifetime lifetime
ALS 4.2.2 MAD 1e12 75x80 1.7e+08 124 kGy/s 4 m 16 s
ALS 5.0.1 mono 2e11 100 2.5e+07 13 kGy/s 39 m 2.6 m
ALS 5.0.2 MAD 1.5e12 100 1.5e+08 76.3 kGy/s 6.6 m 26 s
ALS 5.0.3 mono 3e11 100 3.8e+07 19.4 kGy/s 26 m 1.7 m
ALS 8.2.1 MAD 3.5e11 100 4.5e+07 22.7 kGy/s 22 m 88 s
ALS 8.2.2 MAD 3.5e11 100 4.5e+07 22.7 kGy/s 22 m 88 s
ALS 8.3.1 MAD 9e11 70 2.3e+08 119 kGy/s 4.2 m 17 s
ALS 12.3.1 MAD 2e11 65x90 3.4e+07 17.4 kGy/s 29 m 1.9 m
ALS 12.3.1 ML 4.0e13 65x90 6.8e+09 6.89 MGy/s 4.4 s 0.29 s
SSRL 1-5 MAD 1.7e10 200 4.2e+05 202 Gy/s 41 h 2.8 h
SSRL 7-1 mono 2.6e11 200 6.5e+06 3.09 kGy/s 2.7 h 11 m
SSRL 9-1 mono 3.9e10 200 9.8e+05 463 Gy/s 18 h 72 m
SSRL 9-2 MAD 4.8e11 200 1.2e+07 5.7 kGy/s 88 m 5.8 m
SSRL 11-1 MAD 3.9e11 200 9.8e+06 4.63 kGy/s 1.8 h 7.2 m
SSRL 11-3 mono 2.6e10 200 6.5e+05 302 Gy/s 28 h 1.8 h
SSRL 12-2 MAD 4e12 90x5 8.9e+09 5.01 MGy/s 6 s 0.4 s
SSRL 12-1 MAD 4e12 5x5 1.6e+11 90 MGy/s 0.35 s 22 ms
SSRL 12-1 ML 3e14 5x5 1.1e+13 6.4 GGy/s 5 ms 300 us
APS 24-ID-C MAD 1.3e13 20x60 1.1e+10 5.23 MGy/s 5.7 s 0.38 s
APS "typical" MAD 1.5e12 80 2.3e+08 119 kGy/s 4.2 m 17 s
synch line type flux beamsize flux density dose max xtal min site
ph/s μm ph/μm2/s rate lifetime lifetime

## Slide 23
Where:
IDL - average damage-limited intensity (photons/hkl) at a given resolution
105 - converting R from μm to m, re from m to Å, ρ from g/cm3 to kg/m3 and MGy to Gy
re - classical electron radius (2.818 x 10-15 m/electron)
h - Planck’s constant (6.626 x 10-34 J∙s)
c - speed of light (299792458 m/s)
fdecayed - fractional progress toward completely faded spots at end of data set
ρ - density of crystal (~1.2 g/cm3)
R - radius of the spherical crystal (μm)
λ - X-ray wavelength (Å)
fNH - the Nave & Hill (2005) dose capture fraction (1 for large crystals)
nASU - number of proteins in the asymmetric unit
Mr - molecular weight of the protein (Daltons or g/mol)
VM - Matthews’s coefficient (~2.4 Å3/Dalton)
H - Howells’s criterion (10 MGy/Å)
θ - Bragg angle
a2 - number-averaged squared structure factor per protein atom (electron2)
Ma - number-averaged atomic weight of a protein atom (~7.1 Daltons)
B - average (Wilson) temperature factor (Å2)
μ - attenuation coefficient of sphere material (m-1)
μen - mass energy-absorption coefficient of sphere material (m-1)
Self-calibrated damage limit
Holton & Frankel (2010) Acta D 66 393-408.
No flux
No symmetry
No exposure time

## Slide 24
The number of photons scattered
before crystal is dead
is independent
of data collection time
1 um3 = 105 photons (roomT)
1 um3 = 106 photons (synch)
1 um3 = 108 photons (XFEL)
Henderson, 1990; Gonzalez & Nave, 1994; Glaeser et al., 2000; Sliz et al., 2003; Leiros et al., 2006; Owen et al., 2006; Garman & McSweeney, 2006; Garman & Nave, 2009; Holton, 2009
(synch)

## Slide 25: Dose slicing
N
photons
N
photons
crystal’s useful life
unacceptable
completeness
N
photons
unacceptable
read noise

## Slide 26: What do I do?
You only have a few small crystals…
Should you:
  Collect 360° from each?
  Collect 10° from each at 36x exposure?
  Glue 36 xtals together, then collect 360° ?
  Glue together and do 12960° ?

## Slide 27: Dose slicing: is that a spot?

## Slide 28: Dose slicing: is that a spot?

## Slide 29: Dose slicing: is that a spot?

## Slide 30: Dose slicing: is that a spot?

## Slide 31: Dose slicing: is that a spot?

## Slide 32: Dose slicing: is that a spot?

## Slide 33
Resolution (Å)
CC vs right answer
“true” resolution limit

## Slide 34
Dose slice collection scenario |  |  | 
fine | coarse | glue | Glue & fine
3.17700 | 3.14057 | 3.17604 | 3.12086
“true” resolution vs strategy
301,640,334 photons/xtal

## Slide 35
processing procedure | resolution limit |  |  | 
 | fine | coarse | glue | glue&fine
XDS/XSCALE | 3.17700 | 3.14057 | 3.17604 | 3.12086
XDS/aimless | 3.16251 | 3.15922 | 3.16675 | 3.12162
XDS/noscale | 3.27289 | 3.15617 | 3.38929 | 3.12162
MOSFLM-defaults | x | x | x | x
MOSFLM-cheat | 3.19755 | 3.18166 | 3.20354 | 3.19755
HKL2000-defaults | x | x | x | x
HKL2000-expert | 3.21803 | 3.12646 | 3.11551 | 3.21803
HKL2000-cheat | 3.16119 | 3.10338 | 3.10338 | 3.16300
DIALS-defaults | x | x | 3.42125 | x
DIALS-expert | 3.10296 | 3.09547 | 3.07653 | 3.08465
d*Trek | ? | ? | 4.34076 | ?
EVAL15 | 3.14352 | 3.10383 | 3.09936 | 3.14105
EVAL15-sadabs | 3.13920 | 3.10797 | 3.10171 | 3.13700

## Slide 36: Optimum exposure time?
~1 photon/pixel
  XDS, DIALS
~10 photon/pixel
HKL2000, MOSFLM
~30 photon/pixel
CCD detector

## Slide 38
Holton & Frankel (2010) Acta D 66 393-408.

## Slide 39: What does all that stuff in the concrete tunnel do?

## Slide 40: X-ray optics: mirrors
Notes: Next slide please

## Slide 41: Why are the mirrors so long?
3 mm
Must be < 0.2°
1000 mm
Notes: Next slide please

## Slide 42: Monochromators
λ = 1 Å
Si
Notes: Next slide please

## Slide 43: Monochromators
λ = 1 Å
Si
d = 3.135 Å
θ = 9.177°
λ = 1 Å
nλ = 2d sin(θ)
Notes: Next slide please

## Slide 44: Monochromators
nλ = 2d sin(θ)
λ = 1.1 Å
Si
d = 3.135 Å
θ = 9.177°
Notes: Next slide please

## Slide 45: Monochromators
nλ = 2d sin(θ)
λ = 1.1 Å
Si
d = 3.135 Å
θ = 10.104°
λ = 1.1 Å
Notes: Next slide please

## Slide 46: Monochromators
nλ = 2d sin(θ)
λ = 1 Å
Si
d = 3.135 Å
θ = 9.177°
λ = 1 Å
Notes: Next slide please

## Slide 47: Monochromators
Si
λ = 1 Å
λ = 1 Å
Si
d = 3.135 Å
DCM = Double Crystal Monochromator
Notes: Next slide please

## Slide 48: Monochromators
Si
λ = 1.1 Å
λ = 1.1 Å
Si
d = 3.135 Å
DCM = Double Crystal Monochromator
Notes: Next slide please

## Slide 49: Monochromators
nλ = 2d sin(θ)
λ = 1.54 Å
θ = 1.84°
λ = 1.54 Å
Mo/C
d = 24 Å
Multilayer optics
Notes: Next slide please

## Slide 50: The “beam line”
Area Detector
Source
Primary
mirror
Secondary
mirror
Si(111)
monochromator
Sample
focused
beam
slits
Scatter
guard
Unfocused
beam
slits
N
S

## Slide 51: The “beam line”
Area Detector
Source
Primary
mirror
Secondary
mirror
Si(111)
monochromator
Sample
focused
beam
slits
Scatter
guard
Unfocused
beam
slits
N
S
Notes: Most notable acheivements are high brightness with very little “wasted” light. The discovery made by Howard Padmore? At ALS showed that a 2:1 demagnification cancels nearly all of the spherical abberations in the parabolic/torroid system. 30-50% of the total x-ray intensity passes through the 100um pinhole. The flux through the 100um pinhole at the superbend beamlines is comparable to that of the 5.0.2 wiggler or an APS bending magnet, but only 150 Watts of power must be dissipated by the beamline optics. In contrast, wiggler beamlines must disipate kilowatts. The optics here are much simpler and not exotic in any way. The system is stable and robust. There is never a need to re-focus the beam at a new wavelength. Only a 2-minute automatic tuneup is recommended for very large energy changes (more than 1-2 keV).

## Slide 52: X-ray optics
Area Detector
Protein Crystal
pinhole
Scatter
guard
Cite the “beamline paper!”
Know “final” beam properties
X-ray
Source
Notes: Most notable acheivements are high brightness with very little “wasted” light. The discovery made by Howard Padmore? At ALS showed that a 2:1 demagnification cancels nearly all of the spherical abberations in the parabolic/torroid system. 30-50% of the total x-ray intensity passes through the 100um pinhole. The flux through the 100um pinhole at the superbend beamlines is comparable to that of the 5.0.2 wiggler or an APS bending magnet, but only 150 Watts of power must be dissipated by the beamline optics. In contrast, wiggler beamlines must disipate kilowatts. The optics here are much simpler and not exotic in any way. The system is stable and robust. There is never a need to re-focus the beam at a new wavelength. Only a 2-minute automatic tuneup is recommended for very large energy changes (more than 1-2 keV).

## Slide 53: The truth about x-ray beams
Term units significance
Flux photons/s duration of experiment
Beam Size μm match to crystal
Divergence mrad spot size vs distance
Wavelength Å resolution and absorption
Dispersion Δλ/λ spot size
Flux density ph/s/area scattering/damage rate
Fluence ph/area scattering/damage

## Slide 54: The truth about x-ray beams
Term units significance
Flux photons/s duration of experiment
Beam Size μm match to crystal
Divergence mrad spot size vs distance
Wavelength Å resolution and absorption
Dispersion Δλ/λ spot size
Flux density ph/s/area scattering/damage rate
Fluence ph/area scattering/damage

## Slide 55: The truth about x-ray beams
quantity | units | home source | 12-2
flux | Photons/second | 2.5 x 109 | 2 x 1012
exposure | seconds | 400 | 1

## Slide 56: The truth about x-ray beams
quantity | units | home source | 12-2
flux | Photons/second | 2.5 x 109 | 2 x 1012
exposure | seconds | 400 | 1
Dispersion | wavelength range / wavelength | 0.2% (Kα1 - Kα2) | 0.014% (Si111)
Divergence | milliRadian | 4.8 | 1.8 (h) 0.18 (v)
Beam size | microns | 100 | 24 (h) 12 (v)
Spectral brightness | Photons/s/mm2/ mR2/0.1%BW | 5.4 x 109 | 1.5 x 1017

## Slide 57: The truth about x-ray beams
quantity | units | home source | 12-2
flux | Photons/second | 72 x 106 | 2 x 1012
exposure | time | 4 hours | 1 second
Dispersion | wavelength range / wavelength | 0.2% (Kα1 - Kα2) | 0.014% (Si111)
Divergence | milliRadian | 4.8 | 1.8 (h) 0.18 (v)
Beam size | microns | 17 | 24 (h) 12 (v)
Spectral brightness | Photons/s/mm2/ mR2/0.1%BW | 5.4 x 109 | 1.5 x 1017

## Slide 58: The truth about x-ray beams
quantity | units | home source | 12-2
flux | Photons/second | 1 x 106 | 2 x 1012
exposure | time | 10 days | 1 second
Dispersion | wavelength range / wavelength | 0.2% (Kα1 - Kα2) | 0.014% (Si111)
Divergence | milliRadian | 0.6 | 1.8 (h) 0.18 (v)
Beam size | microns | 17 | 24 (h) 12 (v)
Spectral brightness | Photons/s/mm2/ mR2/0.1%BW | 5.4 x 109 | 1.5 x 1017

## Slide 59: The truth about x-ray beams
quantity | units | home source | 12-2
flux | Photons/second | 8 x 104 | 2 x 1012
exposure | time | 5 months | 1 second
Dispersion | wavelength range / wavelength | 0.014% (Si111) | 0.014% (Si111)
Divergence | milliRadian | 0.6 | 1.8 (h) 0.18 (v)
Beam size | microns | 17 | 24 (h) 12 (v)
Spectral brightness | Photons/s/mm2/ mR2/0.1%BW | 5.4 x 109 | 1.5 x 1017

## Slide 60: The truth about x-ray beams
Term units significance
Flux photons/s duration of experiment
Beam Size μm match to crystal
Divergence mrad spot size vs distance
Wavelength Å resolution and absorption
Dispersion Δλ/λ spot size
Flux density ph/s/area scattering/damage rate
Fluence ph/area scattering/damage

## Slide 61: Si(111) vs multilayers
0.014%
Si(111)
1%
multilayer
4 Å
1.9 Å
resolution
spectral
dispersion
Beam Center

## Slide 62: spectral dispersion
Ewald sphere
spindle axis
Φ circle
diffracted ray
(h,k,l)
d*
λ*
λ*

## Slide 63: spectral dispersion
Ewald sphere
spindle axis
Φ circle
diffracted ray
(h,k,l)
d*
λ’*
λ’*

## Slide 64: The truth about x-ray beams
Term units significance
Flux photons/s duration of experiment
Beam Size μm match to crystal
Divergence mrad spot size vs distance
Wavelength Å resolution and absorption
Dispersion Δλ/λ spot size
Flux density ph/s/area scattering/damage rate
Fluence ph/area scattering/damage

## Slide 65: beam divergence
Ewald sphere
spindle axis
Φ circle
diffracted ray
(h,k,l)
d*
λ*
λ*

## Slide 66: beam divergence
spindle axis
Φ circle
diffracted ray
(h,k,l)
d*
Ewald sphere
λ*
λ*

## Slide 67: divergence = 0 º

## Slide 68: divergence = 0.3 º

## Slide 69: Divergence Arndt & Wonacott (1977)

## Slide 70: The truth about x-ray beams
Term units significance
Flux photons/s duration of experiment
Beam Size μm match to crystal
Divergence mrad spot size vs distance
Wavelength Å resolution and absorption
Dispersion Δλ/λ spot size
Flux density ph/s/area scattering/damage rate
Fluence ph/area scattering/damage

## Slide 71: Air absorption
Distance from sample (mm)
transmittance

## Slide 72: The truth about x-ray beams
Term units significance
Flux photons/s duration of experiment
Beam Size μm match to crystal
Divergence mrad spot size vs distance
Wavelength Å resolution and absorption
Dispersion Δλ/λ spot size
Flux density ph/s/area scattering/damage rate
Fluence ph/area scattering/damage

## Slide 73: Data quality - simple rules:
  Put your crystal into the beam

## Slide 74: “crystal” = thing you want to shoot
Membrane Protein Expression Center © 2012

## Slide 75: Data quality - simple rules:
  Put your crystal into the beam
  Shoot the whole crystal

## Slide 76: shoot the whole crystal

## Slide 77: shoot the whole crystal

## Slide 78: shoot the whole crystal

## Slide 79: shoot the whole crystal

## Slide 80: shoot the whole crystal

## Slide 81: How many crystals do you see?
Shoot the crystal (singular)

## Slide 82: mosaic spread
Ewald sphere
spindle axis
Φ circle
diffracted ray
(h,k,l)
λ*
λ*
d*
d*

## Slide 83: mosaic spread
Ewald sphere
spindle axis
Φ circle
diffracted rays
(h,k,l)
d*
d*

## Slide 84: mosaic spread = 0 º

## Slide 85: mosaic spread = 0.1º

## Slide 86: mosaic spread = 0.2º

## Slide 87: mosaic spread = 0.4º

## Slide 88: mosaic spread = 0.6º

## Slide 89: mosaic spread = 0.8º

## Slide 90: mosaic spread = 1.0º

## Slide 91: mosaic spread = 1.5º

## Slide 92: mosaic spread = 2.0º

## Slide 93: mosaic spread = 2.5º

## Slide 94: mosaic spread = 3.2º

## Slide 95: mosaic spread = 6.4º

## Slide 96: mosaic spread = 12.8º

## Slide 97: Data quality - simple rules:
  Put your crystal into the beam
  Shoot the whole crystal
  Shoot nothing but the crystal

## Slide 98: “crystal” = thing you want to shoot
Membrane Protein Expression Center © 2012

## Slide 99: shoot nothing but the crystal

## Slide 100: shoot nothing but the crystal

## Slide 101: shoot nothing but the crystal

## Slide 102: shoot nothing but the crystal

## Slide 103: shoot nothing but the crystal

## Slide 104: Gaussian beam
50% of flux
the other 50%

## Slide 105: Collimated beam
100% of flux
0%

## Slide 106: Tipping point for µfocus:
10 µm
~10 mm
~1 mrad =0.06°

## Slide 107: Data quality - simple rules:
  Put your crystal into the beam
  Shoot the whole crystal
  Shoot nothing but the crystal
  Back off!

## Slide 108: Background scattering

## Slide 109: Mosquito tips cost $0.15 each
So do Pilatus pixels!
We bought 6,224,001
Use them!
(they are reusable)

## Slide 110: Background scattering
$100,000.00
$100,000.00
$100,000.00
$100,000.00
$100,000.00
$100,000.00
$100,000.00
$100,000.00
real estate is
expensive
use it!

## Slide 111: Fine Slicing
Pflugrath, J. W. (1999)."The finer things in X-ray diffraction data collection", Acta Cryst. D 55, 1718-1725.
background
background

## Slide 112: Optimal exposure time (faint spots)
thr Optimal exposure time for data set (s)
tref exposure time of reference image (s)
bgref background level near weak spots on
reference image (ADU)
bg0 ADC offset of detector (ADU)
bghr optimal background level (via thr)
σ0 rms read-out noise (ADU)
gain ADU/photon
m multiplicity of data set (including partials)
adjust exposure
so this is ~100

## Slide 113: Data quality - simple rules:
  Put your crystal into the beam
  Shoot the whole crystal
  Shoot nothing but the crystal
  Back off!
  The crystal must rotate

## Slide 114: The crystal rotates!

## Slide 115: The crystal rotates!
Membrane Protein Expression Center © 2012

## Slide 116: Multi-crystal strategies
Kendrew et al. (1960) "Structure of Myoglobin” Nature 185, 422-427.

## Slide 117: Basic Principles
“Hell, there are NO RULES here - we're trying to accomplish something.”
Thomas A. Edison – inventor
“You’ve got to have an ASSAY.”
Arthur Kornberg – Nobel Laureate
“Control, control, you must learn CONTROL!”
Yoda – Jedi Master

## Slide 118: Summary
X-rays come from: vac. + elec. + $$$
photons/area
Divergence
Dispersion
Collimation
1 um3 = 106 photons
assay, control and open mind

## Slide 119: Optimum resolution cutoff is:
Too optimistic: add nothing but noise
Too pessimistic: series-termination error
Happy medium?
Simulate:
  Random atoms, compute F2
  Add Gaussian noise, RMS = 1
  Truncate
  Subtract “right” map, RMS difference

## Slide 120
Resolution cutoff (Å)
Error in map (rms e-)
Optimal resolution cutoff

## Slide 121
Resolution cutoff (Å)
Correlation Coefficient
Optimal resolution cutoff

## Slide 122
Resolution cutoff (Å)
Correlation Coefficient
Optimal resolution cutoff

## Slide 123
Resolution cutoff (Å)
Error in map (rms e-)
Optimal resolution cutoff

## Slide 124
log( intensity )
Wilson Plot
2.2 1.6 1.3 1.1 1.0 0.9 0.85 0.8 0.75 0.7
resolution (Å)
8
2

## Slide 125: Optimum resolution cutoff is:
0.0 Å

## Slide 126: Trick Question: What is the space group?
a = 63 b = 63 c = 63
α = 90 β = 90 γ = 90

## Slide 127: What you know when:
Autoindexing: Lattice (14)
Scaling: Point Group (43)
Refinement: Space Group (84)

## Slide 128: There are 16 different “beam centers”
Origin
Beam X
Beam Y

## Slide 129: There are 16 different “beam centers”
Origin
Beam Y
Beam X

## Slide 130: There are 16 different “beam centers”
Origin
Beam Y
Beam X

## Slide 131: There are 16 different “beam centers”
Origin
Beam X
Beam Y

## Slide 132: There are 16 different “beam centers”
Origin
Beam X
Beam Y
Cameraman
View
Beam
View

## Slide 133: There are 16 different “beam centers”
Origin?
Origin?
Origin?
Origin?
Origin?
Origin?
Origin?
Origin?

## Slide 134: There are 16 different “beam centers”
Xbeam, Ybeam
Ybeam, Xwidth – Xbeam
Xwidth-Xbeam, Ywidth-Ybeam
Ybeam, Xbeam
Xbeam, Ywidth-Ybeam
Xbeam, Ybeam-Ywidth
Xbeam, -Ybeam
Xwidth-Xbeam, Ybeam
Xwidth-Xbeam, Ywidth-Ybeam
Xwidth-Xbeam, Ybeam-Ywidth
Xwidth-Xbeam, -Ybeam
Xbeam, Ybeam
Xbeam, Ywidth-Ybeam
Xbeam, Ybeam-Ywidth
Xbeam, -Ybeam