## Slide 1: PowerPoint File available:
http://bl831.als.lbl.gov/
~jamesh/powerpoint/
SHSSS_Berlin_2015.pptx

## Slide 2: Acknowledgements
UCSF LBNL SLAC
ALS 8.3.1 creator: Tom Alber
Center for Structure of Membrane Proteins (CSMP)
Membrane Protein Expression Center II (MPEC)
Center for HIV Accessory and Regulatory Complexes (HARC)
UC Multicampus Research Programs and Initiatives (MRPI)
UCSF Program for Breakthrough Biomedical Research (PBBR) Integrated Diffraction Analysis Technologies (IDAT)
Plexxikon, Inc.
M D Anderson CRC
Synchrotron Radiation Structural Biology Resource (SLAC)
The Advanced Light Source is supported by the Director, Office of Science, Office of Basic Energy Sciences, Materials Sciences Division, of the US Department of Energy under contract No. DE-AC02-05CH11231 at Lawrence Berkeley National Laboratory.
Robert Stroud James Fraser Christine Gee Tom Peat Janet Newman
Chris Nielsen Clemens Schulze-Briese Meitian Wang Aina Cohen Ana Gonzalez

## Slide 3: Can you count to 1,000,000 ?
= 0.1%
sqrt(1,000,000)
1,000,000
= 3%
sqrt(1,000)
1,000
> 1000 is a waste!
photon
spot
Theoretically:
In reality:
ISa ~ 33
Rmeas ≈ 0.1% ?
ISa = 1000
Rmeas = ≈ 3%

## Slide 4
Required signal-to-noise (I/σ)
Solve-able proteins (%)
Current
technology
Goal
Threshold of a revolution in phasing
S-SAD
7 keV

## Slide 5: SHSSS! The dominant source of error for anomalous difference measurements
Spatial
Heterogeneity in
Sharp
Spot
Sensitivity
Flat
Field

## Slide 6: SHSSS! The dominant source of error for anomalous difference measurements
Spatial
Heterogeneity in
Sharp
Spot
Sensitivity
Flat
Field
90%
110%
x 1.11 =
x 0.91 =
100%
100%

## Slide 7: SHSSS! The dominant source of error for anomalous difference measurements
Spatial
Heterogeneity in
Sharp
Spot
Sensitivity
Sharp
Spot
20%
110%
x 1.11 =
x 0.91 =
100%
22%

## Slide 8: SHSSS! The dominant source of error for anomalous difference measurements
Spatial
Heterogeneity in
Sharp
Spot
Sensitivity
Sharp
Spot
90%
30%
x 1.11 =
x 0.91 =
27.3%
100%
Integral: 127.3

## Slide 9: SHSSS! The dominant source of error for anomalous difference measurements
Spatial
Heterogeneity in
Sharp
Spot
Sensitivity
Sharp
Spot
20%
110%
x 1.11 =
x 0.91 =
100%
22.2%
Integral: 122.2

## Slide 10
Spatial Heterogeneity in Sharp Spot Sensitivity

## Slide 11
Spatial Heterogeneity in Sharp Spot Sensitivity

## Slide 12
Spot centroid position (pixels)
Relative spot intensity
Spatial Heterogeneity in Sharp Spot Sensitivity

## Slide 13
Spot centroid position (pixels)
Relative spot intensity
Spatial Heterogeneity in Sharp Spot Sensitivity

## Slide 14: SHSSS for MX spots

## Slide 15: SHSSS for MX spots
down

## Slide 16: SHSSS for MX spots
down
up

## Slide 17: SHSSS for MX spots
down
up
Rseparate

## Slide 18: SHSSS for MX spots
odd
even
Rmixed

## Slide 19: SHSSS for MX spots
separate:
2.5%

## Slide 20: SHSSS for MX spots
separate:
mixed:
2.5%
0.9%

## Slide 21: SHSSS for MX spots
separate:
mixed:
2.5%
0.9%
2.5%2-0.9%2 = 2.3%2

## Slide 22
Rsepar = sqrt(5.92/mult + 2.12)
Rmixed = sqrt(5.92/mult + 0.22)
Rdiff (%) between half-sets
5
4
3
2
1
0
Image multiplicity
ADSC Q315r
3-pixel shift

## Slide 23
Rdiff (%) between half-sets
5
4
3
2
1
0
Pilatus
3-pixel shift
Image multiplicity
Rsepar = sqrt(3.72/mult + 0.52)
Rmixed = sqrt(3.72/mult + 0.32)

## Slide 24
Rdiff (%) between half-sets
5
4
3
2
1
0
Image multiplicity
Pilatus
1-module shift
Rsepar = sqrt(4.12/mult + 3.22)
Rmixed = sqrt(4.12/mult + 0.22)

## Slide 25
3.5
3.0
2.5
2.0
1.5
1.0
0.5
SHSSS
Systematic component of Rsseparate (%)
distance between spots (mm)
0.1 1 10 100
CCD detector
anomalous mates are always on
different modules
different
modules
Pilatus
SN001
Spatial Heterogeneity in Sharp Spot Sensitivity
Notes: Comparision of SHSSS measured by the detector-shift experiment on three different detectors:
The Q315R in ALS 8.3.1
The first Pilatus detector at the Swiss Light Soruce
A newer Pilatus detector at Stanford
Improvements in calibration technology have finally made
Pilatus competitive with calibration on CCDs

## Slide 26: Pilatus: 1-module shift

## Slide 27
Pilatus: 1-module shift

## Slide 28
Pilatus: 1-module shift

## Slide 29
Pilatus: 1-module shift - aligned

## Slide 30
Pilatus: 1-module shift - ratio
+3%
- 3%

## Slide 31
3.5
3.0
2.5
2.0
1.5
1.0
0.5
SHSSS
Systematic component of Rsseparate (%)
distance between spots (mm)
0.1 1 10 100
CCD detector
anomalous mates are always on
different modules
different
modules
Pilatus
SN001
Pilatus
SN113
Notes: Comparision of SHSSS measured by the detector-shift experiment on three different detectors:
The Q315R in ALS 8.3.1
The first Pilatus detector at the Swiss Light Soruce
A newer Pilatus detector at Stanford
Improvements in calibration technology have finally made
Pilatus competitive with calibration on CCDs

## Slide 32
Pick-up tool mark
Bragg
glitch
oxygen
inclusions
“tree rings”
1% high
average
1% low
~3x105 photon/pixel → 0.3% error
Pilatus: subtract smooth baseline
Gollwitzer & Krumrey (2014)
J. Appl. Cryst. 47, 378

## Slide 33: Beam Flicker
time (seconds)
normalized flux through pinhole
pinhole
removed
0.15%/√Hz

## Slide 34
Source of error | realistic simulation | No SHSSS | Perfect detector
Photon counting | + | + | +
Shutter jitter | + | + | -
Beam flicker | + | + | -
Sample absorption | + | + | -
Radiation damage | + | + | -
Imperfect spindle | + | + | -
vignette | + | + | -
Corner correction | + | + | -
SHSSS | + | - | -
Rmeas (∞-10 Å) | 2.8% | 0.7% | 0.7%
I/σ asymptotic | 26.8 | 74.2 | 81.0
Threshold of a revolution in phasing
Holton et al (2014) "R-factor gap", FEBS Journal 281, 4046-4060.

## Slide 35: Oh! SHSSS! how do we get rid of it?
Calibrate it out
Average over it
go MAD

## Slide 36: Oh! SHSSS! how do we get rid of it?
Calibrate it out
Average over it
go MAD

## Slide 37: Detector pixels are 3D
51 μm
31 μm thick
17%
loss
Typical CCD
Mean Penetration depth 12 keV:
17.5 μm Gd2O2S 237 μm Si
8%
side
loss
172 μm
26% loss
320 μm thick
Pilatus
53%
side
loss
photon
75 μm
15%
loss
450 μm thick
Eiger
76%
side
loss
photon
Arrival angle of 2 Å spot: 29°

## Slide 38: Detector pixels are 3D
N
N
N
θ
λ
detection
event
incoming
photon

## Slide 39: Detector pixels are 3D
thickness
width
N
N
N
Bragg
glitch
oxygen
inclusion
θ
λ
detection
event
incoming
photon

## Slide 40: Oh! SHSSS! how do we get rid of it?
Calibrate it out
Average over it
go MAD

## Slide 41: Averaging over SHSSS
mult > (—)2
~3%
<ΔF/F>
“true” multiplicity = different pixels

## Slide 42: Averaging with systematic error

## Slide 43: CCD calibration: 7235 eV

## Slide 44: CCD calibration: 7247 eV

## Slide 45: Gadox calibration vs energy
photon energy (keV)
Relative absorption depth
same = good!
bad!
Mar?

## Slide 46: Detector calibration: 7247 eV
target:
oil
distance:
900 mm
2θ: 12°

## Slide 47: Detector calibration: 7235 eV
target:
oil
distance:
900 mm
2θ: 12°

## Slide 48: ADSC Q315r SN 926 (ALS 8.3.1)
-10%
+10%

## Slide 49
ADSC Q315r SN )

## Slide 50: ADSC Q315r SN )

## Slide 51: Detector calibration
calibration error (%)
megapixels

## Slide 52
7247 eV

## Slide 53
7235 eV

## Slide 54: Detector calibration
7223 eV

## Slide 55
Never use same pixel twice
Detector calibration
- 7235 eV
Radiation Damage
- 1% error per MGy
Isomorphous crystals
- Humidity!
- Friend or foe?
Averaging over SHSSS

## Slide 56
Never use same pixel twice
Detector calibration
- 7235 eV
Radiation Damage
- 1% error per MGy
Isomorphous crystals
- Humidity!
- Friend or foe?
Averaging over SHSSS

## Slide 57
Never use same pixel twice
Detector calibration
- 7235 eV
Radiation Damage
- 1% error per MGy
Isomorphous crystals
- Humidity!
- Friend or foe?
Averaging over SHSSS

## Slide 58: Riso vs dose
Riso (%)
change in dose (MGy)
data taken from Banumathi, et al. (2004) Acta Cryst. D 60, 1085-1093.
Riso ≈ 0.7 %/MGy

## Slide 59: Damage Limit → Dose slicing
N
photons
N
photons
unacceptable
damage
N
photons
unacceptable
read noise
1 um3 xtal = 106 photons

## Slide 60: 140-fold multiplicity
SUBSET OF INTENSITY DATA WITH SIGNAL/NOISE >= -3.0 AS FUNCTION OF RESOLUTION
RESOLUTION NUMBER OF REFLECTIONS COMPLETENESS R-FACTOR R-FACTOR COMPARED I/SIGMA R-meas CC(1/2) Anomal SigAno Nano
LIMIT OBSERVED UNIQUE POSSIBLE OF DATA observed expected Corr
9.17 188919 1061 1071 99.1% 3.9% 4.8% 188919 257.47 3.9% 100.0* 91* 5.024 450
6.49 356827 1933 1933 100.0% 5.2% 5.5% 356827 214.33 5.2% 100.0* 86* 3.836 882
5.30 466503 2515 2515 100.0% 7.2% 7.0% 466503 165.13 7.2% 100.0* 76* 3.257 1175
4.59 548405 2974 2974 100.0% 7.2% 6.6% 548405 175.42 7.3% 100.0* 67* 2.589 1403
4.10 615117 3353 3355 99.9% 7.7% 6.7% 615117 174.13 7.7% 100.0* 59* 2.264 1594
3.74 684947 3734 3737 99.9% 9.4% 8.3% 684947 143.09 9.4% 100.0* 49* 1.953 1783
3.47 743557 4051 4051 100.0% 11.2% 10.1% 743557 120.17 11.2% 100.0* 39* 1.696 1942
3.24 799722 4366 4366 100.0% 14.1% 13.9% 799722 91.14 14.1% 100.0* 30* 1.333 2103
3.06 839252 4598 4603 99.9% 19.5% 20.2% 839252 65.79 19.5% 100.0* 23* 1.117 2214
2.90 891398 4904 4908 99.9% 29.0% 31.7% 891398 44.85 29.1% 99.9* 17* 1.008 2369
2.77 931622 5148 5151 99.9% 40.5% 44.8% 931622 32.58 40.6% 99.8* 11* 0.901 2493
2.65 969629 5384 5387 99.9% 52.8% 58.8% 969629 25.16 52.9% 99.8* 10* 0.866 2605
2.54 997213 5574 5574 100.0% 67.4% 76.0% 997213 19.47 67.6% 99.6* 2 0.804 2705
2.45 1031222 5889 5889 100.0% 88.9% 101.2% 1031222 14.58 89.2% 99.2* 4 0.831 2859
2.37 788617 6008 6008 100.0% 109.3% 125.5% 788617 9.97 109.7% 98.1* 5 0.829 2925
2.29 614145 6252 6252 100.0% 138.2% 161.4% 614145 6.87 138.9% 96.1* 1 0.760 3037
2.22 470283 6481 6481 100.0% 197.1% 231.7% 470283 4.03 198.6% 83.5* -1 0.721 3159
2.16 259836 6629 6631 100.0% 227.3% 268.7% 259831 2.41 230.8% 46.9* -1 0.677 3224
2.10 36259 4169 6807 61.2% 154.4% 169.4% 36229 1.28 163.6% 47.0* -2 0.660 1999
2.05 13567 3372 7037 47.9% 170.1% 187.0% 13503 0.68 196.5% 25.7* 3 0.629 1578
total 12247040 88395 94730 93.3% 15.7% 16.4% 12246941 54.30 15.8% 100.0* 12* 1.217 42499
16 crystals, 360° each, < 1 MGy, inverse beam
7235 eV ADSC Q210r Australian Synchrotron MX1
I/SIGMA
257.47

## Slide 61: 140-fold multiplicity
18 σ
Phased anomalous difference Fourier
data Courtesy of Tom Peat & Janet Newman
16 σ
Not
Lysozyme!

## Slide 62: 140-fold multiplicity
15σ = PO4
Phased anomalous difference Fourier
data Courtesy of Tom Peat & Janet Newman
Not
Lysozyme!

## Slide 63: 140-fold multiplicity
~2σ = Mg?
Phased anomalous difference Fourier
data Courtesy of Tom Peat & Janet Newman
Not
Lysozyme!

## Slide 64: 140-fold multiplicity
8.2σ = Na
DELFAN residual anomalous difference
data Courtesy of Tom Peat & Janet Newman
10 σ is enough for phasing!
Bunkoczi et al. (2015)."weak, single-wavelength anomalous", Nat. Meth. 12, 127-130.
Not
Lysozyme!
f” = 0.15

## Slide 65: Discerning Na+ from Mg++
f’’ (electrons)
DELFAN peak height (σ)
Mg
Ne
Na
F
O
N

## Slide 66
Never use same pixel twice
Detector calibration
- 7235 eV
Radiation Damage
- 1% error per MGy
Isomorphous crystals
- Humidity!
- Friend or foe?
Averaging over SHSSS

## Slide 67: Non-isomorphism in lysozyme
RH: 84.2%
vs 71.9%
Riso = 44.5%
3aw6
3aw7
Δcell = 0.7 %
RMSD = 0.18 Å

## Slide 68
Never use same pixel twice
Detector calibration
- 7235 eV
Radiation Damage
- 1% error per MGy
Isomorphous crystals
- Humidity!
- Friend or foe?
Averaging over SHSSS

## Slide 69
h,k,l | structure factors (F) |  |  | 
 | #1 | #2 | #3 | #4
5,3,4 | 523.7 | 559.8 | 579.9 | 603.2
5,4,4 | 168.2 | 166.6 | 177.2 | 196.1
5,5,4 | 34.9 | 26.4 | 19.2 | 17.3
6,1,4 | 305.7 | 301.1 | 298.1 | 296.3
6,2,4 | 353.0 | 353.9 | 356.2 | 366.9
6,3,4 | 300.9 | 285.3 | 273.5 | 259.4
6,3,5 | 223.8 | 226.3 | 234.6 | 251.4
… | … | … | … | …
X-ray Data Sets
h,k,l | 100% | 17 % | 13 % | 9 %
 | 1st | 2nd | 3rd | 4th
5,3,4 | -1.46 | -0.02 | 1.84 | -0.72
5,4,4 | -0.88 | -0.29 | -0.34 | 0.15
5,5,4 | -0.42 | -0.65 | 1.02 | 1.47
6,1,4 | -0.90 | -0.85 | 1.44 | -0.40
6,2,4 | -1.20 | -0.37 | 0.67 | 0.01
6,3,4 | -0.75 | 0.31 | 0.48 | 0.00
6,3,5 | -0.75 | -0.85 | 0.72 | -0.82
… | … | … | … | …
data
set
Singular value decomp.
SVD
vector
value
Singular values & vectors
Correlation
Coefficients
-0.68
CC2
CC3
CC1
CC1 CC2 CC3
data set #1
Data sets
Positioned
in
“Correlation
Space”
-0.24
data set #2
data set #3
0.32 …

## Slide 71
Correlation to 2nd & 3rd singular vectors
Structure factor (electrons)

## Slide 72
Correlation to 2nd & 3rd singular vectors
Structure factor (electrons)

## Slide 73
Correlation to 2nd & 3rd singular vectors
Structure factor (electrons)

## Slide 74
CC = 0.02
Phases from non-isomorphism?
DMMULTI – fake data - 4 deg rotation: 8 “xtals”

## Slide 75
CC = 0.8
DMMULTI – fake data - 4 deg rotation: 8 “xtals”
Phases from non-isomorphism!

## Slide 76: Oh! SHSSS! how do we get rid of it?
Calibrate it out
Average over it
go MAD

## Slide 77
Photon Energy (eV)
Anomalous f” (electrons)
2472.7 2473.5
Bohic et al. Anal. Chem. 2008; 80(24):9557. doi: 10.1021/ac801817k

## Slide 78: Two wavelengths are better than one

## Slide 79
Required signal-to-noise (I/σ)
Solve-able proteins (%)
Current
detectors

## Slide 80
Excellent signal at S edge
- 2472.7 eV
x Radiation Damage
x Self absorption
Pro & Con of S-MAD

## Slide 81: S-MAD Sample size Dilemma

## Slide 82: Darwin’s Formula
Ispot - photons/spot (fully-recorded)
Ibeam - incident (photons/s/m2 )
re - classical electron radius (2.818x10-15 m)
Vxtal - volume of crystal (in m3)
Vcell - volume of unit cell (in m3)
λ - x-ray wavelength (in meters!)
ω - rotation speed (radians/s)
L - Lorentz factor (speed/speed)
P - polarization factor
A - attenuation correction
F0 - structure factor at T = 0
B - Debye-Waller-Ott factor
s - 0.5/d-spacing
C. G. Darwin (1914)
P A | F0e-Bs2 |2
Ispot = Ibeam re2
Vxtal
Vcell
λ3 L
ωVcell

## Slide 83
Where:
IDL - average damage-limited intensity (photons/hkl) at a given resolution
105 - converting R from μm to m, re from m to Å, ρ from g/cm3 to kg/m3 and MGy to Gy
re - classical electron radius (2.818 x 10-15 m/electron)
h - Planck’s constant (6.626 x 10-34 J∙s)
c - speed of light (299792458 m/s)
fdecayed - fractional progress toward completely faded spots at end of data set
ρ - density of crystal (~1.2 g/cm3)
R - radius of the spherical crystal (μm)
λ - X-ray wavelength (Å)
fNH - the Nave & Hill (2005) dose capture fraction (1 for large crystals)
nASU - number of proteins in the asymmetric unit
Mr - molecular weight of the protein (Daltons or g/mol)
VM - Matthews’s coefficient (~2.4 Å3/Dalton)
H - Howells’s criterion (10 MGy/Å)
θ - Bragg angle
a2 - number-averaged squared structure factor per protein atom (electron2)
Ma - number-averaged atomic weight of a protein atom (~7.1 Daltons)
B - average (Wilson) temperature factor (Å2)
μ - attenuation coefficient of sphere material (m-1)
μen - mass energy-absorption coefficient of sphere material (m-1)
Self-calibrated damage limit
Holton & Frankel (2010) Acta D 66 393-408.

## Slide 84
wavelength dependence
crystal diameter (μm)
Damage-limited minimum crystal size
Holton & Frankel (2010) Acta D 66 393-408.

## Slide 85
Holton & Frankel (2010) Acta D 66 393-408.

## Slide 86: Where do photons go?
beamstop
97%
transmitted
1 Å x-rays
attenuation correction
error cannot be > ~3%
100 μm
thick
protein

## Slide 87: attenuation correction
Bouguer, P. (1729). Essai d'optique sur la gradation de la lumière.
Lambert, J. H. (1760). Photometria: sive De mensura et gradibus luminis, colorum et umbrae. E. Klett.
Beer, A. (1852)."Bestimmung der Absorption des rothen Lichts in farbigen Flüssigkeiten", Ann. Phys. Chem 86, 78-90.
A = = exp(-μt)
IT
Ibeam
t
μ

## Slide 88: attenuation correction
Bouguer, P. (1729). Essai d'optique sur la gradation de la lumière.
Lambert, J. H. (1760). Photometria: sive De mensura et gradibus luminis, colorum et umbrae. E. Klett.
Beer, A. (1852)."Bestimmung der Absorption des rothen Lichts in farbigen Flüssigkeiten", Ann. Phys. Chem 86, 78-90.
A = = exp(-μ(tin+ tout))
IT
Ibeam
t
tin
tout
A = = exp(-μt)
IT
Ibeam
μ

## Slide 89: attenuation correction
Bouguer, P. (1729). Essai d'optique sur la gradation de la lumière.
Lambert, J. H. (1760). Photometria: sive De mensura et gradibus luminis, colorum et umbrae. E. Klett.
Beer, A. (1852)."Bestimmung der Absorption des rothen Lichts in farbigen Flüssigkeiten", Ann. Phys. Chem 86, 78-90.
A = = exp(-μ(tin+ tout))
IT
Ibeam
tin
tout
tin
tout
tin
tout
μ

## Slide 90: attenuation correction
Bouguer, P. (1729). Essai d'optique sur la gradation de la lumière.
Lambert, J. H. (1760). Photometria: sive De mensura et gradibus luminis, colorum et umbrae. E. Klett.
Beer, A. (1852)."Bestimmung der Absorption des rothen Lichts in farbigen Flüssigkeiten", Ann. Phys. Chem 86, 78-90.
A = = exp[-μxtal(txi+ txo)
-μsolvent(tsi + tso)]
IT
Ibeam
μxtal
txi
txo
tsi
tso
txi
txo
tsi
tso
txi
txo
tsi
tso
μsolvent

## Slide 91: Scattering/absorption “rules”
1 μm crystal ≈ 1 μm water
≈ 1 μm plastic
≈ 0.1 μm glass
≈ 1000 μm air
~ 50 mm He
≈ 370 mm water vapor @ 4C

## Slide 92: Where do photons go?
beamstop
97%
transmitted
1 Å x-rays
attenuation correction
error cannot be > ~3%
100 μm
thick
protein

## Slide 93: Where do photons go?
beamstop
2%
transmitted
5 Å x-rays
attenuation correction
error can be ~98%
100 μm
thick
protein

## Slide 94: Where do photons go?
beamstop
96%
transmitted
5 Å x-rays
attenuation correction
error cannot be > ~4%
1 μm
thick
protein

## Slide 95: “sweet spot” for sample size ?
S
P
Mg
Na
C
F
N
O
Cl
Ar
K
Ca
Ne
Se

## Slide 96
Excellent signal at S edge
- 2472.7 eV
x Radiation Damage
- min size ~ 100 μm
x Self absorption
- 20 μm error = 50% error
Pro & Con of S-MAD

## Slide 97: Dose-rate dependence of damage
dose rate (kGy/s)
maximum useful dose (MGy)
1 um3 xtal =
106 photons
1 um3 xtal =
108 photons

## Slide 98
0.006°
1 μm xtal
1 Å x-rays
Lovelace et al. (2006)."topography", J. Appl. Cryst. 39, 425-432.
The “partiality problem” at XFEL

## Slide 99
mult > (—)2
~100%
<ΔF/F>
Gd lyso: ΔF/F = 8.7% → mult = 132
The “partiality problem” at XFEL
Barends et al. (2014) Nature 505, 244-247.

## Slide 100: Osculating Ewald Spheres: SINBAD
Ewald sphere 2
diffracted ray
λ*
λ*
θ
1
Ewald sphere
λ*
(h,k,l)
diffracted ray
λ*
θ
d*
(-h,-k,-l)
d*

## Slide 101
h,k,l
-h,-k,-l
Detector
Detector
Detector
Detector
λ = 5 Å
sample
injector
Si(111)
52.87deg
Si(111)
Si(111)
Si(111)
Si(111)
Si(111)
2 Multilayer mirrors
d=2nm, W/B4C
KB Horiz focus
KB vertical focus
~ 1m
Notes: Proposed 5-crystal monochromator for colliding-beam crystallography. The generation of two opposing Ewald spheres means that for every h,k,l index reflected from the protein crystal, the Friedel mate is reflected in the opposite direction, allowing for very accurate anomalous difference measurements. The wavelength need not be set precisely at the sulfur or phosphorous edge, but merely at a wavelength where the f” of those elements are strong. For fs-class pulses, the distances above will have to be precise to withing ~3 microns, and for accurate reproduction of the partiality of the Friedel mates, the angular precision of the colliding beams must be within ~1% of the sample crystal’s rocking width, or ~100 microrads.

## Slide 102: “sweet spot” for sample size ?
2 Å
2sin(90°)
3 Å
2sin(90°)
S
P
Mg
Na
C
F
N
O
Cl
Ar
K
Ca
Ne
Se
4 Å
2sin(90°)

## Slide 103
Excellent signal at S edge
- 2472.7 eV
x Radiation Damage
- min size ~ 100 μm
x Self absorption
- 20 μm error = 50% error
Solutions
- outrun damage at XFEL
- Bragg geometry?
Pro & Con of S-MAD

## Slide 105
Diamond
($100, reusable)

## Slide 106
Diamond
($100, reusable)
oil ?

## Slide 107
Diamond
($100, reusable)
oil ?

## Slide 108
oil ?
λ=2d sinθ
2.5 Å data
with 5 Å X-rays
Only analytic absorption correction:
International Tables for Crystallography, Vol. C, 2nd ed., chapter 6.3

## Slide 109
Excellent signal at S edge
- 2472.7 eV
x Radiation Damage
- min size ~ 100 μm
x Self absorption
- 20 μm error = 50% error
Solutions
- outrun damage at XFEL
- Bragg geometry?
Pro & Con of S-MAD

## Slide 110: The Way Forward:
Calibrate it out
  - Impossible for CCDs
  - 3D pixel model on PADs
Average over SHSSS
  - non-isomorphism: foe or friend?
go MAD
  XFEL/SINBAD cancels errors
  Return to Bragg geometry (analytic absorption)
http://bl831.als.lbl.gov/~jamesh/powerpoint/
SHSSS_Berlin_2015.pptx