## Slide 1: Acknowledgements
UCSF LBNL SLAC
ALS 8.3.1 creator: Tom Alber
Integrated Diffraction Analysis Technologies (IDAT)
Center for Structure of Membrane Proteins (CSMP)
Membrane Protein Expression Center II (MPEC)
Center for HIV Accessory and Regulatory Complexes (HARC)
UC Multicampus Research Programs and Initiatives (MRPI)
UCSF Program for Breakthrough Biomedical Research (PBBR) Plexxikon, Inc.
M D Anderson CRC
The Advanced Light Source is supported by the Director, Office of Science, Office of Basic Energy Sciences, Materials Sciences Division, of the US Department of Energy under contract No. DE-AC02-05CH11231 at Lawrence Berkeley National Laboratory.
Chris Neilson Aina Cohen Ana Gonzalez Mike Soltis
Meitian Wang Clemens Schulze-Briese Joe Ferarra

## Slide 2: SHSSS! The dominant source of error for anomalous difference measurements
Spatial
Heterogeneity in
Sharp
Spot
Sensitivity
Flat
Field

## Slide 3: SHSSS! The dominant source of error for anomalous difference measurements
Spatial
Heterogeneity in
Sharp
Spot
Sensitivity
Flat
Field
90%
110%
x 1.11 =
x 0.91 =
100%
100%

## Slide 4: SHSSS! The dominant source of error for anomalous difference measurements
Spatial
Heterogeneity in
Sharp
Spot
Sensitivity
Sharp
Spot
20%
110%
x 1.11 =
x 0.91 =
100%
22%

## Slide 5: SHSSS! The dominant source of error for anomalous difference measurements
Spatial
Heterogeneity in
Sharp
Spot
Sensitivity
Sharp
Spot
90%
30%
x 1.11 =
x 0.91 =
27.3%
100%
Integral: 127.3

## Slide 6: SHSSS! The dominant source of error for anomalous difference measurements
Spatial
Heterogeneity in
Sharp
Spot
Sensitivity
Sharp
Spot
20%
110%
x 1.11 =
x 0.91 =
100%
22.2%
Integral: 122.2

## Slide 7: Can you count to 1,000,000?
= 0.1%
Source of error | realistic | No SHSSS | Perfect detector
Photon counting | + | + | +
Shutter jitter | + | + | -
Beam flicker | + | + | -
Sample absorption | + | + | -
Radiation damage | + | + | -
Imperfect spindle | + | + | -
Corner correction | + | + | -
SHSSS | + | - | -
R-meas (∞-10 Å) | 2.8% | 0.7% | 0.7%
I/sigma asymptotic | 26.8 | 74.2 | 81.0
sqrt(1,000,000)
1,000,000
= Rmeas?
I/σ = 1000
Holton et al. (2014)."R-factor gap", FEBS J. 281, 4046-4060.

## Slide 8: S-SAD distribution of difficulty
Bijvoet ratio (%)
Number of proteins
I/σ = 33
I/σ = 50
I/σ = 100
I/σ = 200

## Slide 9
How do you measure SHSSS?

## Slide 10
Spatial Heterogeneity in Sharp Spot Sensitivity
How do you measure SHSSS?
Shuffle evaluation order:
Convert systematic → random

## Slide 11
How do you measure SHSSS?
0 1 2 3 4 5 6 7
Spot position (pixels)
Relative integrated spot intensity
1.10
1.05
1.00
0.95
0.90
0.85

## Slide 12: SHSSS for MX spots

## Slide 13: SHSSS for MX spots
down

## Slide 14: SHSSS for MX spots
down
up

## Slide 15: SHSSS for MX spots
down
up
Rseparate

## Slide 16: SHSSS for MX spots
odd
even
Rmixed

## Slide 17: SHSSS for MX spots
separate:
2.5%

## Slide 18: SHSSS for MX spots
separate:
mixed:
2.5%
0.9%

## Slide 19: SHSSS for MX spots
separate:
mixed:
2.5%
0.9%
2.5%2-0.9%2 = 2.3%2

## Slide 20
sqrt(5.92/mult + 2.12)
sqrt(5.92/mult + 0.22)
Rdiff (%) between half-sets
5
4
3
2
1
0
Image multiplicity
Q315r
10-pixel shift

## Slide 21
Rdiff (%) between half-sets
5
4
3
2
1
0
sqrt(4.12/mult + 3.22)
sqrt(4.12/mult + 0.22)
Image multiplicity
Pilatus
1-module shift

## Slide 22
Rdiff (%) between half-sets
5
4
3
2
1
0
sqrt(3.72/mult + 0.52)
sqrt(3.72/mult + 0.32)
Pilatus
1-pixel shift
Image multiplicity

## Slide 23: SHSSS: CCD vs PAD
3.5
3.0
2.5
2.0
1.5
1.0
0.5
Systematic component of Rshift (%)
distance between spots (mm)
0.1 1 10 100
Pilatus
Q315r
anomalous mates are
> 100 mm apart
different
modules

## Slide 24: Pilatus: 1-module shift

## Slide 25
Pilatus: 1-module shift

## Slide 26
Pilatus: 1-module shift

## Slide 27
Pilatus: 1-module shift - aligned

## Slide 28
Pilatus: 1-module shift - ratio
+3%
- 3%

## Slide 29
~3x105 photon/pixel ratio vs smooth
+1%
- 1%
Pilatus: flood field

## Slide 30: SHSSS: CCD vs PAD
3.5
3.0
2.5
2.0
1.5
1.0
0.5
Systematic component of Rshift (%)
distance between spots (mm)
0.1 1 10 100
Pilatus
Q315r
anomalous mates are
> 100 mm apart
different
modules
recent
model

## Slide 31
Gadolinium edge contrast map of fiber-coupled CCD detector. Black to white is a 20% range.
CCD: Gd edge contrast

## Slide 32: CCD calibration: 7235 eV

## Slide 33: CCD calibration: 7247 eV

## Slide 34: Detector calibration: 7247 eV
target:
oil
distance:
900 mm
2θ: 12°

## Slide 35: Detector calibration: 7235 eV
target:
oil
distance:
900 mm
2θ: 12°

## Slide 36: ADSC Q315r SN 926 (ALS 8.3.1)
-10%
+10%

## Slide 37
ADSC Q315r SN )

## Slide 38: ADSC Q315r SN )

## Slide 39: Detector calibration
calibration error (%)
megapixels

## Slide 40: Gadox calibration vs energy
photon energy (keV)
Relative absorption depth
same = good!
bad!
Mar

## Slide 41
7247 eV

## Slide 42
7235 eV

## Slide 43: Detector calibration
7223 eV

## Slide 44: Oh! SHSSS! how do we get rid of it?
Average over it
Model it out
go MAD

## Slide 45: Oh! SHSSS! how do we get rid of it?
Average over it
Model it out
go MAD

## Slide 46: Averaging doesn’t always reduce error

## Slide 47: Averaging over SHSSS
mult > (—)2
~3%
<ΔF/F>
= 3%
sqrt(1,000)
1,000
> 1000 is a waste!
photon
spot
“true” multiplicity = never use same pixel twice

## Slide 48: Dose slicing
N
photons
N
photons
unacceptable
damage
N
photons
unacceptable
read noise
100 um3 xtal = 1012 photons

## Slide 49: Riso vs dose
Riso (%)
change in dose (MGy)
data taken from Banumathi, et al. (2004) Acta Cryst. D 60, 1085-1093.
Riso ≈ 0.7 %/MGy

## Slide 50: 140-fold multiplicity
SUBSET OF INTENSITY DATA WITH SIGNAL/NOISE >= -3.0 AS FUNCTION OF RESOLUTION
RESOLUTION NUMBER OF REFLECTIONS COMPLETENESS R-FACTOR R-FACTOR COMPARED I/SIGMA R-meas CC(1/2) Anomal SigAno Nano
LIMIT OBSERVED UNIQUE POSSIBLE OF DATA observed expected Corr
9.17 188919 1061 1071 99.1% 3.9% 4.8% 188919 257.47 3.9% 100.0* 91* 5.024 450
6.49 356827 1933 1933 100.0% 5.2% 5.5% 356827 214.33 5.2% 100.0* 86* 3.836 882
5.30 466503 2515 2515 100.0% 7.2% 7.0% 466503 165.13 7.2% 100.0* 76* 3.257 1175
4.59 548405 2974 2974 100.0% 7.2% 6.6% 548405 175.42 7.3% 100.0* 67* 2.589 1403
4.10 615117 3353 3355 99.9% 7.7% 6.7% 615117 174.13 7.7% 100.0* 59* 2.264 1594
3.74 684947 3734 3737 99.9% 9.4% 8.3% 684947 143.09 9.4% 100.0* 49* 1.953 1783
3.47 743557 4051 4051 100.0% 11.2% 10.1% 743557 120.17 11.2% 100.0* 39* 1.696 1942
3.24 799722 4366 4366 100.0% 14.1% 13.9% 799722 91.14 14.1% 100.0* 30* 1.333 2103
3.06 839252 4598 4603 99.9% 19.5% 20.2% 839252 65.79 19.5% 100.0* 23* 1.117 2214
2.90 891398 4904 4908 99.9% 29.0% 31.7% 891398 44.85 29.1% 99.9* 17* 1.008 2369
2.77 931622 5148 5151 99.9% 40.5% 44.8% 931622 32.58 40.6% 99.8* 11* 0.901 2493
2.65 969629 5384 5387 99.9% 52.8% 58.8% 969629 25.16 52.9% 99.8* 10* 0.866 2605
2.54 997213 5574 5574 100.0% 67.4% 76.0% 997213 19.47 67.6% 99.6* 2 0.804 2705
2.45 1031222 5889 5889 100.0% 88.9% 101.2% 1031222 14.58 89.2% 99.2* 4 0.831 2859
2.37 788617 6008 6008 100.0% 109.3% 125.5% 788617 9.97 109.7% 98.1* 5 0.829 2925
2.29 614145 6252 6252 100.0% 138.2% 161.4% 614145 6.87 138.9% 96.1* 1 0.760 3037
2.22 470283 6481 6481 100.0% 197.1% 231.7% 470283 4.03 198.6% 83.5* -1 0.721 3159
2.16 259836 6629 6631 100.0% 227.3% 268.7% 259831 2.41 230.8% 46.9* -1 0.677 3224
2.10 36259 4169 6807 61.2% 154.4% 169.4% 36229 1.28 163.6% 47.0* -2 0.660 1999
2.05 13567 3372 7037 47.9% 170.1% 187.0% 13503 0.68 196.5% 25.7* 3 0.629 1578
total 12247040 88395 94730 93.3% 15.7% 16.4% 12246941 54.30 15.8% 100.0* 12* 1.217 42499
16 crystals, 360° each, inverse beam, 7235 eV AS MX1

## Slide 51
RESOLUTION COMPLETENESS R-FACTOR I/SIGMA R-meas CC(1/2) Anomal SigAno Nano
LIMIT OF DATA observed Corr
9.17 99.1% 3.9% 257.47 3.9% 100.0* 91* 5.024 450
6.49 100.0% 5.2% 214.33 5.2% 100.0* 86* 3.836 882
5.30 100.0% 7.2% 165.13 7.2% 100.0* 76* 3.257 1175
4.59 100.0% 7.2% 175.42 7.3% 100.0* 67* 2.589 1403
4.10 99.9% 7.7% 174.13 7.7% 100.0* 59* 2.264 1594
3.74 99.9% 9.4% 143.09 9.4% 100.0* 49* 1.953 1783
3.47 100.0% 11.2% 120.17 11.2% 100.0* 39* 1.696 1942
3.24 100.0% 14.1% 91.14 14.1% 100.0* 30* 1.333 2103
3.06 99.9% 19.5% 65.79 19.5% 100.0* 23* 1.117 2214
2.90 99.9% 29.0% 44.85 29.1% 99.9* 17* 1.008 2369
2.77 99.9% 40.5% 32.58 40.6% 99.8* 11* 0.901 2493
2.65 99.9% 52.8% 25.16 52.9% 99.8* 10* 0.866 2605
2.54 100.0% 67.4% 19.47 67.6% 99.6* 2 0.804 2705
2.45 100.0% 88.9% 14.58 89.2% 99.2* 4 0.831 2859
2.37 100.0% 109.3% 9.97 109.7% 98.1* 5 0.829 2925
2.29 100.0% 138.2% 6.87 138.9% 96.1* 1 0.760 3037
2.22 100.0% 197.1% 4.03 198.6% 83.5* -1 0.721 3159
2.16 100.0% 227.3% 2.41 230.8% 46.9* -1 0.677 3224
2.10 61.2% 154.4% 1.28 163.6% 47.0* -2 0.660 1999
2.05 47.9% 170.1% 0.68 196.5% 25.7* 3 0.629 1578
total 93.3% 15.7% 54.30 15.8% 100.0* 12* 1.217 42499
140-fold multiplicity: 16 crystals, 360° each, inverse beam, 7235 eV

## Slide 52: 140-fold multiplicity
18 σ
Phased anomalous difference Fourier
data Courtesy of Tom Peat & Janet Newman
16 σ

## Slide 53: 140-fold multiplicity
15σ = PO4
Phased anomalous difference Fourier
data Courtesy of Tom Peat & Janet Newman

## Slide 54: 140-fold multiplicity
~2σ = Mg?
Phased anomalous difference Fourier
data Courtesy of Tom Peat & Janet Newman

## Slide 55: 140-fold multiplicity
8.2σ = Mg?
DELFAN residual anomalous difference
data Courtesy of Tom Peat & Janet Newman
10 σ is enough for phasing!
Bunkoczi et al. (2015)."weak, single-wavelength anomalous", Nat. Meth. 12, 127-130.

## Slide 56: Discerning Na+ from Mg++
f’’ (electrons)
DELFAN peak height (σ)
Mg
Ne
Na
F
O
N

## Slide 57: Non-isomorphism in lysozyme
RH 84.2% vs 71.9%
Riso = 44.5%
RMSD = 0.18 Å
PDB:
3aw6
3aw7
Δa = 0.8%

## Slide 58: plunge cooling
foam insulation
warm, moist air
cold N2
liquid N2
ice

## Slide 59: plunge cooling
cold N2
foam insulation
warm, moist air
liquid N2
ice

## Slide 60: Warkentin method
foam insulation
liquid N2
Warkentin (2006) J. Appl. Crystallogr. 39, 805-811.

## Slide 61: Oh! SHSSS! how do we get rid of it?
Average over it
Model it out
go MAD

## Slide 62
~3x105 photon/pixel ratio vs smooth
+1%
- 1%
Pilatus: flood field

## Slide 63: Detector pixels are 3D
75 μm
150 μm
172 μm
51 μm
31 μm thick
26% loss
15%
loss
12% loss
17%
loss
320 μm thick
450 μm thick
500 μm thick
Pilatus
Eiger
ADSC HF
ADSC CCD
Arrival angle of 2 Å spot: 29°
Mean Penetration depth 12 keV:
237 μm Si 17.5 μm Gd2O2S
53%
side
loss
76%
side
loss
57%
side
loss
8%
side
loss
photon
photon
photon

## Slide 64: Detector pixels are 3D
θ
λ

## Slide 65: Detector pixels are 3D
θ
λ

## Slide 66: Detector pixels are 3D
5% thickness variation
θ
λ

## Slide 67: HPAD calibration
θ
θ
λ
λ
thick
width
N
N
N
trim
trim
trim
N
N
N
trim
trim
trim
N
trim

## Slide 68: HPAD vs mmPAD
N
N
N
trim
trim
trim
N
N
N
trim
trim
trim
Q
Q
Q
gain
gain
gain
= k
= k
= k
θ
θ
λ
λ
thick
width

## Slide 69: Oh! SHSSS! how do we get rid of it?
Average over it
Model it out
go MAD

## Slide 70: Two wavelengths are better than one

## Slide 71: S-MAD Sample size Dilemma

## Slide 72
wavelength dependence
crystal diameter (μm)
Å
Å
Damage limited crystal size (synchrotron)
Holton & Frankel, “minimum crystal size” Acta Cryst. D 66, 393 (2010).

## Slide 73: Darwin’s Formula
Ispot - photons/spot (fully-recorded)
Ibeam - incident (photons/s/m2 )
re - classical electron radius (2.818x10-15 m)
Vxtal - volume of crystal (in m3)
Vcell - volume of unit cell (in m3)
λ - x-ray wavelength (in meters!)
ω - rotation speed (radians/s)
L - Lorentz factor (speed/speed)
P - polarization factor
A - attenuation correction
F0 - structure factor at T = 0
B - Debye-Waller-Ott factor
s - 0.5/d-spacing
C. G. Darwin (1914)
P A | F0e-Bs2 |2
Ispot = Ibeam re2
Vxtal
Vcell
λ3 L
ωVcell

## Slide 74: The “partiality problem” at XFEL
0.006°
1 μm xtal
1 Å x-rays

## Slide 75
mult > (—)2
~100%
<ΔF/F>
Gd lyso: ΔF/F = 8.7% → mult = 132
The “partiality problem” at XFEL
Barends et al. (2014) Nature 505, 244-247.

## Slide 76: Osculating Ewald Spheres
Ewald sphere 2
diffracted ray
λ*
λ*
θ
1
Ewald sphere
λ*
(h,k,l)
diffracted ray
λ*
θ
d*
(-h,-k,-l)
d*

## Slide 77
h,k,l
-h,-k,-l
Detector
Detector
Detector
Detector
λ = 5 Å
sample
injector
Si(111)
52.87deg
Si(111)
Si(111)
Si(111)
Si(111)
Si(111)
2 Multilayer mirrors
d=2nm, W/B4C
KB Horiz focus
KB vertical focus
~ 1m
Notes: Proposed 5-crystal monochromator for colliding-beam crystallography. The generation of two opposing Ewald spheres means that for every h,k,l index reflected from the protein crystal, the Friedel mate is reflected in the opposite direction, allowing for very accurate anomalous difference measurements. The wavelength need not be set precisely at the sulfur or phosphorous edge, but merely at a wavelength where the f” of those elements are strong. For fs-class pulses, the distances above will have to be precise to withing ~3 microns, and for accurate reproduction of the partiality of the Friedel mates, the angular precision of the colliding beams must be within ~1% of the sample crystal’s rocking width, or ~100 microrads.

## Slide 78: Optimum sample size w/o rad dam
2 Å
2sin(90°)
3 Å
2sin(90°)
S
P
Mg
Na
C
F
N
O
Cl
Ar
K
Ca
Ne
Se
4 Å
2sin(90°)

## Slide 79: The Way Forward:
Average over SHSSS
  - non-isomorphism problem
Model it out
  - Impossible for CCDs
  - Need to constrain pixel model on PADs
go MAD
  - XFEL cures sample size dilemma
  - SINBAD cancels new errors
http://bl831.als.lbl.gov/~jamesh/powerpoint/SHSSS_UGA_2015.pptx
http://bl831.als.lbl.gov/xtallife.html

## Slide 80: Where do photons go?
beamstop
97%
transmitted
1 Å x-rays
attenuation correction
error cannot be > ~3%
100 μm
thick
protein

## Slide 81: Where do photons go?
beamstop
2%
transmitted
5 Å x-rays
attenuation correction
error can be ~98%
100 μm
thick
protein

## Slide 82: Where do photons go?
beamstop
96%
transmitted
5 Å x-rays
attenuation correction
error cannot be > ~4%
1 μm
thick
protein

## Slide 83: Scattering/absorption “rules”:
1 μm crystal ≈ 1 μm water
≈ 1 μm plastic
≈ 0.1 μm glass
≈ 1000 μm air
~ 50 mm He
≈ 370 mm water vapor @ 4C

## Slide 85
Diamond
($100, reusable)

## Slide 86
Diamond
($100, reusable)
oil ?

## Slide 87
Diamond
($100, reusable)
oil ?

## Slide 88
oil ?
λ=2d sinθ
2.5 Å data
with 5 Å X-rays
Only analytic absorption correction:
International Tables for Crystallography, Vol. C, 2nd ed., chapter 6.3

## Slide 89: attenuation correction
Bouguer, P. (1729). Essai d'optique sur la gradation de la lumière.
Lambert, J. H. (1760). Photometria: sive De mensura et gradibus luminis, colorum et umbrae. E. Klett.
Beer, A. (1852)."Bestimmung der Absorption des rothen Lichts in farbigen Flüssigkeiten", Ann. Phys. Chem 86, 78-90.
A = = exp(-μt)
IT
Ibeam
t
μ

## Slide 90: attenuation correction
Bouguer, P. (1729). Essai d'optique sur la gradation de la lumière.
Lambert, J. H. (1760). Photometria: sive De mensura et gradibus luminis, colorum et umbrae. E. Klett.
Beer, A. (1852)."Bestimmung der Absorption des rothen Lichts in farbigen Flüssigkeiten", Ann. Phys. Chem 86, 78-90.
A = = exp(-μ(tin+ tout))
IT
Ibeam
t
tin
tout
A = = exp(-μt)
IT
Ibeam
μ

## Slide 91: attenuation correction
Bouguer, P. (1729). Essai d'optique sur la gradation de la lumière.
Lambert, J. H. (1760). Photometria: sive De mensura et gradibus luminis, colorum et umbrae. E. Klett.
Beer, A. (1852)."Bestimmung der Absorption des rothen Lichts in farbigen Flüssigkeiten", Ann. Phys. Chem 86, 78-90.
A = = exp(-μ(tin+ tout))
IT
Ibeam
tin
tout
tin
tout
tin
tout
μ

## Slide 92: attenuation correction
Bouguer, P. (1729). Essai d'optique sur la gradation de la lumière.
Lambert, J. H. (1760). Photometria: sive De mensura et gradibus luminis, colorum et umbrae. E. Klett.
Beer, A. (1852)."Bestimmung der Absorption des rothen Lichts in farbigen Flüssigkeiten", Ann. Phys. Chem 86, 78-90.
A = = exp[-μxtal(txi+ txo)
-μsolvent(tsi + tso)]
IT
Ibeam
μxtal
txi
txo
tsi
tso
txi
txo
tsi
tso
txi
txo
tsi
tso
μsolvent

## Slide 93: 100% SeMet incorporation
Trivial to solve

## Slide 94: 100% S -Met incorporation
Impossible to solve

## Slide 95: Solved Structure Accuracy
fraction Sulfur
correlation coefficient
Perfect sites

## Slide 96: Solved Structure Accuracy
fraction Sulfur
correlation coefficient
Perfect sites

## Slide 97: Solved Structure Accuracy
fraction Sulfur
correlation coefficient
Perfect sites

## Slide 98: Solved Structure Accuracy
fraction Sulfur
correlation coefficient
Perfect sites

## Slide 99
Two-color experiment (1930)