## Slide 1: Acknowledgements
UCSF LBNL SLAC
ALS 8.3.1 creator: Tom Alber
Integrated Diffraction Analysis Technologies (IDAT)
Center for Structure of Membrane Proteins (CSMP)
Membrane Protein Expression Center II (MPEC)
Center for HIV Accessory and Regulatory Complexes (HARC)
UC Multicampus Research Programs and Initiatives (MRPI)
UCSF Program for Breakthrough Biomedical Research (PBBR) Plexxikon, Inc.
M D Anderson CRC
The Advanced Light Source is supported by the Director, Office of Science, Office of Basic Energy Sciences, Materials Sciences Division, of the US Department of Energy under contract No. DE-AC02-05CH11231 at Lawrence Berkeley National Laboratory.
Chris Neilson Aina Cohen Ana Gonzalez Mike Soltis
Meitian Wang Clemens Schulze-Briese Joe Ferarra

## Slide 2: SHSSS! The dominant source of error for anomalous difference measurements
Spatial
Heterogeneity in
Sharp
Spot
Sensitivity
Flat
Field

## Slide 3: SHSSS! The dominant source of error for anomalous difference measurements
Spatial
Heterogeneity in
Sharp
Spot
Sensitivity
Flat
Field
90%
110%
x 1.11 =
x 0.91 =
100%
100%

## Slide 4: SHSSS! The dominant source of error for anomalous difference measurements
Spatial
Heterogeneity in
Sharp
Spot
Sensitivity
Sharp
Spot
20%
110%
x 1.11 =
x 0.91 =
100%
22%

## Slide 5: SHSSS! The dominant source of error for anomalous difference measurements
Spatial
Heterogeneity in
Sharp
Spot
Sensitivity
Sharp
Spot
90%
30%
x 1.11 =
x 0.91 =
27.3%
100%
Integral: 127.3

## Slide 6: SHSSS! The dominant source of error for anomalous difference measurements
Spatial
Heterogeneity in
Sharp
Spot
Sensitivity
Sharp
Spot
20%
110%
x 1.11 =
x 0.91 =
100%
22.2%
Integral: 122.2

## Slide 7: SHSSS! The dominant source of error for anomalous difference measurements
Spatial
Heterogeneity in
Sharp
Spot
Sensitivity
Sharp
Spot
90%
30%
x 1.11 =
x 0.91 =
27.3%
100%
Integral: 127.3

## Slide 8: Can you count to 1,000,000 ?
= 0.1%
sqrt(1,000,000)
1,000,000
= 3%
sqrt(1,000)
1,000
> 1000 is a waste!
photon
spot
Theoretically:
In reality:
ISa ~ 33
Rmeas ≈ 0.1% ?
ISa = 1000
Rmeas = ≈ 3%

## Slide 9: S-SAD distribution of difficulty
Bijvoet ratio (%)
Number of proteins
I/σ = 33
I/σ = 50
I/σ = 100
I/σ = 200

## Slide 10: Dominant Source of Error
Source of error | realistic | No SHSSS | Perfect detector
Photon counting | + | + | +
Shutter jitter | + | + | -
Beam flicker | + | + | -
Sample absorption | + | + | -
Radiation damage | + | + | -
Imperfect spindle | + | + | -
Corner correction | + | + | -
SHSSS | + | - | -
R-meas (∞-10 Å) | 2.8% | 0.7% | 0.7%
I/sigma asymptotic | 26.8 | 74.2 | 81.0
Holton et al. (2014)."R-factor gap", FEBS J. 281, 4046-4060.

## Slide 11: SHSSS for MX spots

## Slide 12: SHSSS for MX spots
down

## Slide 13: SHSSS for MX spots
down
up

## Slide 14: SHSSS for MX spots
down
up
Rseparate

## Slide 15: SHSSS for MX spots
odd
even
Rmixed

## Slide 16: Residual SHSSS
 | CCD | Pilatus 001 | Pilatus 11-1
SHSSS (%) | 2.3 | 3.2 | 0.9
Max ISa | 44 | 31 | 111
% PDB mult = 3 | 7 | 1 | 96

## Slide 17: Pilatus: 1-module shift

## Slide 18
Pilatus: 1-module shift

## Slide 19
Pilatus: 1-module shift

## Slide 20
Pilatus: 1-module shift - aligned

## Slide 21
Pilatus: 1-module shift - ratio
+3%
- 3%

## Slide 22: Residual SHSSS
 | CCD | Pilatus 001 | Pilatus 11-1
SHSSS (%) | 2.3 | 3.2 | 0.9
Max ISa | 44 | 31 | 111
% PDB mult = 3 | 7 | 1 | 96

## Slide 23
~3x105 photon/pixel ratio vs smooth
+1%
- 1%
Pilatus: flood field

## Slide 24: Residual SHSSS
 | CCD | Pilatus 001 | Pilatus 11-1
SHSSS (%) | 2.3 | 3.2 | 0.9
Max ISa | 44 | 31 | 111
% PDB mult = 3 | 7 | 1 | 96

## Slide 25: Error from radiation damage
Riso (%)
change in dose (MGy)
data taken from Banumathi, et al. (2004) Acta Cryst. D 60, 1085-1093.
Riso ≈ 0.7 %/MGy

## Slide 26: what the is a MGy?
http://bl831.als.lbl.gov/
damage_rates.pdf
Holton J. M. (2009) J. Synchrotron Rad. 16 133-42

## Slide 29
Holton & Frankel (2010) Acta D 66 393-408.

## Slide 30: Oh! SHSSS! how do we get rid of it?
Average over it
Model it out
go MAD

## Slide 31: Oh! SHSSS! how do we get rid of it?
Average over it
Model it out
go MAD

## Slide 32: Dose slicing
N
photons
N
photons
unacceptable
damage
N
photons
unacceptable
read noise
1 um3 xtal = 107 photons

## Slide 33: Averaging over SHSSS
mult > (—)2
~3%
<ΔF/F>
“true” multiplicity = never use same pixel twice

## Slide 34: 140-fold multiplicity
SUBSET OF INTENSITY DATA WITH SIGNAL/NOISE >= -3.0 AS FUNCTION OF RESOLUTION
RESOLUTION NUMBER OF REFLECTIONS COMPLETENESS R-FACTOR R-FACTOR COMPARED I/SIGMA R-meas CC(1/2) Anomal SigAno Nano
LIMIT OBSERVED UNIQUE POSSIBLE OF DATA observed expected Corr
9.17 188919 1061 1071 99.1% 3.9% 4.8% 188919 257.47 3.9% 100.0* 91* 5.024 450
6.49 356827 1933 1933 100.0% 5.2% 5.5% 356827 214.33 5.2% 100.0* 86* 3.836 882
5.30 466503 2515 2515 100.0% 7.2% 7.0% 466503 165.13 7.2% 100.0* 76* 3.257 1175
4.59 548405 2974 2974 100.0% 7.2% 6.6% 548405 175.42 7.3% 100.0* 67* 2.589 1403
4.10 615117 3353 3355 99.9% 7.7% 6.7% 615117 174.13 7.7% 100.0* 59* 2.264 1594
3.74 684947 3734 3737 99.9% 9.4% 8.3% 684947 143.09 9.4% 100.0* 49* 1.953 1783
3.47 743557 4051 4051 100.0% 11.2% 10.1% 743557 120.17 11.2% 100.0* 39* 1.696 1942
3.24 799722 4366 4366 100.0% 14.1% 13.9% 799722 91.14 14.1% 100.0* 30* 1.333 2103
3.06 839252 4598 4603 99.9% 19.5% 20.2% 839252 65.79 19.5% 100.0* 23* 1.117 2214
2.90 891398 4904 4908 99.9% 29.0% 31.7% 891398 44.85 29.1% 99.9* 17* 1.008 2369
2.77 931622 5148 5151 99.9% 40.5% 44.8% 931622 32.58 40.6% 99.8* 11* 0.901 2493
2.65 969629 5384 5387 99.9% 52.8% 58.8% 969629 25.16 52.9% 99.8* 10* 0.866 2605
2.54 997213 5574 5574 100.0% 67.4% 76.0% 997213 19.47 67.6% 99.6* 2 0.804 2705
2.45 1031222 5889 5889 100.0% 88.9% 101.2% 1031222 14.58 89.2% 99.2* 4 0.831 2859
2.37 788617 6008 6008 100.0% 109.3% 125.5% 788617 9.97 109.7% 98.1* 5 0.829 2925
2.29 614145 6252 6252 100.0% 138.2% 161.4% 614145 6.87 138.9% 96.1* 1 0.760 3037
2.22 470283 6481 6481 100.0% 197.1% 231.7% 470283 4.03 198.6% 83.5* -1 0.721 3159
2.16 259836 6629 6631 100.0% 227.3% 268.7% 259831 2.41 230.8% 46.9* -1 0.677 3224
2.10 36259 4169 6807 61.2% 154.4% 169.4% 36229 1.28 163.6% 47.0* -2 0.660 1999
2.05 13567 3372 7037 47.9% 170.1% 187.0% 13503 0.68 196.5% 25.7* 3 0.629 1578
total 12247040 88395 94730 93.3% 15.7% 16.4% 12246941 54.30 15.8% 100.0* 12* 1.217 42499
16 crystals, 360° each, inverse beam, 7235 eV AS MX1

## Slide 35: Gadox calibration vs energy
photon energy (keV)
Relative absorption depth
same = good!
bad!
Mar

## Slide 36
7247 eV

## Slide 37
7235 eV

## Slide 38: Detector calibration
7223 eV

## Slide 39
RESOLUTION COMPLETENESS R-FACTOR I/SIGMA R-meas CC(1/2) Anomal SigAno Nano
LIMIT OF DATA observed Corr
9.17 99.1% 3.9% 257.47 3.9% 100.0* 91* 5.024 450
6.49 100.0% 5.2% 214.33 5.2% 100.0* 86* 3.836 882
5.30 100.0% 7.2% 165.13 7.2% 100.0* 76* 3.257 1175
4.59 100.0% 7.2% 175.42 7.3% 100.0* 67* 2.589 1403
4.10 99.9% 7.7% 174.13 7.7% 100.0* 59* 2.264 1594
3.74 99.9% 9.4% 143.09 9.4% 100.0* 49* 1.953 1783
3.47 100.0% 11.2% 120.17 11.2% 100.0* 39* 1.696 1942
3.24 100.0% 14.1% 91.14 14.1% 100.0* 30* 1.333 2103
3.06 99.9% 19.5% 65.79 19.5% 100.0* 23* 1.117 2214
2.90 99.9% 29.0% 44.85 29.1% 99.9* 17* 1.008 2369
2.77 99.9% 40.5% 32.58 40.6% 99.8* 11* 0.901 2493
2.65 99.9% 52.8% 25.16 52.9% 99.8* 10* 0.866 2605
2.54 100.0% 67.4% 19.47 67.6% 99.6* 2 0.804 2705
2.45 100.0% 88.9% 14.58 89.2% 99.2* 4 0.831 2859
2.37 100.0% 109.3% 9.97 109.7% 98.1* 5 0.829 2925
2.29 100.0% 138.2% 6.87 138.9% 96.1* 1 0.760 3037
2.22 100.0% 197.1% 4.03 198.6% 83.5* -1 0.721 3159
2.16 100.0% 227.3% 2.41 230.8% 46.9* -1 0.677 3224
2.10 61.2% 154.4% 1.28 163.6% 47.0* -2 0.660 1999
2.05 47.9% 170.1% 0.68 196.5% 25.7* 3 0.629 1578
total 93.3% 15.7% 54.30 15.8% 100.0* 12* 1.217 42499
140-fold multiplicity: 16 crystals, 360° each, inverse beam, 7235 eV

## Slide 40: 140-fold multiplicity
18 σ
Phased anomalous difference Fourier
data Courtesy of Tom Peat & Janet Newman
16 σ

## Slide 41: 140-fold multiplicity
15σ = PO4
Phased anomalous difference Fourier
data Courtesy of Tom Peat & Janet Newman

## Slide 42: 140-fold multiplicity
~2σ = Mg or Na?
Phased anomalous difference Fourier
data Courtesy of Tom Peat & Janet Newman

## Slide 43: 140-fold multiplicity
8.2σ = Na
DELFAN residual anomalous difference
data Courtesy of Tom Peat & Janet Newman
10 σ is enough for phasing!
Bunkoczi et al. (2015)."weak, single-wavelength anomalous", Nat. Meth. 12, 127-130.

## Slide 44: Non-isomorphism in lysozyme
RH 84.2% vs 71.9%
Riso = 44.5%
RMSD = 0.18 Å
PDB:
3aw6
3aw7
Δa = 0.8%

## Slide 45: plunge cooling
foam insulation
warm, moist air
cold N2
liquid N2
ice

## Slide 46: plunge cooling
cold N2
foam insulation
warm, moist air
liquid N2
ice

## Slide 47: Warkentin method
foam insulation
liquid N2
Warkentin (2006) J. Appl. Crystallogr. 39, 805-811.

## Slide 48: Oh! SHSSS! how do we get rid of it?
Average over it
Model it out
go MAD

## Slide 49
~3x105 photon/pixel ratio vs smooth
+1%
- 1%
Pilatus: flood field

## Slide 50: Detector pixels are 3D
75 μm
172 μm
51 μm
31 μm thick
26% loss
15%
loss
17%
loss
320 μm thick
450 μm thick
Pilatus
Eiger
ADSC
CCD
Arrival angle of 2 Å spot: 29°
Mean Penetration depth 12 keV:
237 μm Si 17.5 μm Gd2O2S
53%
side
loss
76%
side
loss
8%
side
loss
photon
photon

## Slide 51: Detector pixels are 3D
θ
λ

## Slide 52: Detector pixels are 3D
θ
λ

## Slide 53: Detector pixels are 3D
5% thickness variation
θ
λ

## Slide 54: Oh! SHSSS! how do we get rid of it?
Average over it
Model it out
go MAD

## Slide 55: Two wavelengths are better than one

## Slide 56: S-MAD Sample size Dilemma

## Slide 57
wavelength dependence
crystal diameter (μm)
Å
Å
Damage limited crystal size (synchrotron)
Holton & Frankel, “minimum crystal size” Acta Cryst. D 66, 393 (2010).

## Slide 58: Darwin’s Formula
Ispot - photons/spot (fully-recorded)
Ibeam - incident (photons/s/m2 )
re - classical electron radius (2.818x10-15 m)
Vxtal - volume of crystal (in m3)
Vcell - volume of unit cell (in m3)
λ - x-ray wavelength (in meters!)
ω - rotation speed (radians/s)
L - Lorentz factor (speed/speed)
P - polarization factor
A - attenuation correction
F0 - structure factor at T = 0
B - Debye-Waller-Ott factor
s - 0.5/d-spacing
C. G. Darwin (1914)
P A | F0e-Bs2 |2
Ispot = Ibeam re2
Vxtal
Vcell
λ3 L
ωVcell

## Slide 59: The “partiality problem” at XFEL
0.006°
1 μm xtal
1 Å x-rays

## Slide 60
mult > (—)2
~100%
<ΔF/F>
Gd lyso: ΔF/F = 8.7% → mult = 132
The “partiality problem” at XFEL
Barends et al. (2014) Nature 505, 244-247.

## Slide 61: Osculating Ewald Spheres
Ewald sphere 2
diffracted ray
λ*
λ*
θ
1
Ewald sphere
λ*
(h,k,l)
diffracted ray
λ*
θ
d*
(-h,-k,-l)
d*

## Slide 62
h,k,l
-h,-k,-l
Detector
Detector
Detector
Detector
λ = 5 Å
sample
injector
Si(111)
52.87deg
Si(111)
Si(111)
Si(111)
Si(111)
Si(111)
2 Multilayer mirrors
d=2nm, W/B4C
KB Horiz focus
KB vertical focus
~ 1m
Notes: Proposed 5-crystal monochromator for colliding-beam crystallography. The generation of two opposing Ewald spheres means that for every h,k,l index reflected from the protein crystal, the Friedel mate is reflected in the opposite direction, allowing for very accurate anomalous difference measurements. The wavelength need not be set precisely at the sulfur or phosphorous edge, but merely at a wavelength where the f” of those elements are strong. For fs-class pulses, the distances above will have to be precise to withing ~3 microns, and for accurate reproduction of the partiality of the Friedel mates, the angular precision of the colliding beams must be within ~1% of the sample crystal’s rocking width, or ~100 microrads.

## Slide 63: The Way Forward:
Average over SHSSS
  - non-isomorphism problem
Model it out
  - Impossible for CCDs
  - Need to constrain pixel model on PADs
go MAD
  - XFEL cures sample size dilemma
  - SINBAD cancels new errors
http://bl831.als.lbl.gov/~jamesh/powerpoint/SHSSS_WCPCW_2015.pptx
http://bl831.als.lbl.gov/xtallife.html
http://bl831.als.lbl.gov/xtalsize.html

## Slide 64: Averaging doesn’t always reduce error

## Slide 65: Optimum sample size w/o rad dam
2 Å
2sin(90°)
3 Å
2sin(90°)
S
P
Mg
Na
C
F
N
O
Cl
Ar
K
Ca
Ne
Se
4 Å
2sin(90°)

## Slide 66: Where do photons go?
beamstop
97%
transmitted
1 Å x-rays
attenuation correction
error cannot be > ~3%
100 μm
thick
protein

## Slide 67: Where do photons go?
beamstop
2%
transmitted
5 Å x-rays
attenuation correction
error can be ~98%
100 μm
thick
protein

## Slide 68: Where do photons go?
beamstop
96%
transmitted
5 Å x-rays
attenuation correction
error cannot be > ~4%
1 μm
thick
protein

## Slide 69: Scattering/absorption “rules”:
1 μm crystal ≈ 1 μm water
≈ 1 μm plastic
≈ 0.1 μm glass
≈ 1000 μm air
~ 50 mm He
≈ 370 mm water vapor @ 4C

## Slide 71
Diamond
($100, reusable)

## Slide 72
Diamond
($100, reusable)
oil ?

## Slide 73
Diamond
($100, reusable)
oil ?

## Slide 74
oil ?
λ=2d sinθ
2.5 Å data
with 5 Å X-rays
Only analytic absorption correction:
International Tables for Crystallography, Vol. C, 2nd ed., chapter 6.3

## Slide 75: attenuation correction
Bouguer, P. (1729). Essai d'optique sur la gradation de la lumière.
Lambert, J. H. (1760). Photometria: sive De mensura et gradibus luminis, colorum et umbrae. E. Klett.
Beer, A. (1852)."Bestimmung der Absorption des rothen Lichts in farbigen Flüssigkeiten", Ann. Phys. Chem 86, 78-90.
A = = exp(-μt)
IT
Ibeam
t
μ

## Slide 76: attenuation correction
Bouguer, P. (1729). Essai d'optique sur la gradation de la lumière.
Lambert, J. H. (1760). Photometria: sive De mensura et gradibus luminis, colorum et umbrae. E. Klett.
Beer, A. (1852)."Bestimmung der Absorption des rothen Lichts in farbigen Flüssigkeiten", Ann. Phys. Chem 86, 78-90.
A = = exp(-μ(tin+ tout))
IT
Ibeam
t
tin
tout
A = = exp(-μt)
IT
Ibeam
μ

## Slide 77: attenuation correction
Bouguer, P. (1729). Essai d'optique sur la gradation de la lumière.
Lambert, J. H. (1760). Photometria: sive De mensura et gradibus luminis, colorum et umbrae. E. Klett.
Beer, A. (1852)."Bestimmung der Absorption des rothen Lichts in farbigen Flüssigkeiten", Ann. Phys. Chem 86, 78-90.
A = = exp(-μ(tin+ tout))
IT
Ibeam
tin
tout
tin
tout
tin
tout
μ

## Slide 78: attenuation correction
Bouguer, P. (1729). Essai d'optique sur la gradation de la lumière.
Lambert, J. H. (1760). Photometria: sive De mensura et gradibus luminis, colorum et umbrae. E. Klett.
Beer, A. (1852)."Bestimmung der Absorption des rothen Lichts in farbigen Flüssigkeiten", Ann. Phys. Chem 86, 78-90.
A = = exp[-μxtal(txi+ txo)
-μsolvent(tsi + tso)]
IT
Ibeam
μxtal
txi
txo
tsi
tso
txi
txo
tsi
tso
txi
txo
tsi
tso
μsolvent

## Slide 79: 100% SeMet incorporation
Trivial to solve

## Slide 80: 100% S -Met incorporation
Impossible to solve

## Slide 81: Solved Structure Accuracy
fraction Sulfur
correlation coefficient
Perfect sites

## Slide 82: Solved Structure Accuracy
fraction Sulfur
correlation coefficient
Perfect sites

## Slide 83: Solved Structure Accuracy
fraction Sulfur
correlation coefficient
Perfect sites

## Slide 84: Solved Structure Accuracy
fraction Sulfur
correlation coefficient
Perfect sites

## Slide 85
Two-color experiment (1930)

## Slide 86: SHSSS: CCD vs PAD
3.5
3.0
2.5
2.0
1.5
1.0
0.5
Systematic component of Rshift (%)
distance between spots (mm)
0.1 1 10 100
Pilatus
Q315r
anomalous mates are
> 100 mm apart
different
modules
SSRL
11-1

## Slide 87
Gadolinium edge contrast map of fiber-coupled CCD detector. Black to white is a 20% range.
CCD: Gd edge contrast

## Slide 88: CCD calibration: 7235 eV

## Slide 89: CCD calibration: 7247 eV

## Slide 90: Detector calibration: 7247 eV
target:
oil
distance:
900 mm
2θ: 12°

## Slide 91: Detector calibration: 7235 eV
target:
oil
distance:
900 mm
2θ: 12°

## Slide 92: ADSC Q315r SN 926 (ALS 8.3.1)
-10%
+10%

## Slide 93
ADSC Q315r SN )

## Slide 94: ADSC Q315r SN )

## Slide 95: Detector calibration
calibration error (%)
megapixels

## Slide 96
How do you measure SHSSS?

## Slide 97
Spatial Heterogeneity in Sharp Spot Sensitivity
How do you measure SHSSS?
Shuffle evaluation order:
Convert systematic → random

## Slide 98
How do you measure SHSSS?
0 1 2 3 4 5 6 7
Spot position (pixels)
Relative integrated spot intensity
1.10
1.05
1.00
0.95
0.90
0.85

## Slide 99: SHSSS for MX spots
separate:
2.5%

## Slide 100: SHSSS for MX spots
separate:
mixed:
2.5%
0.9%

## Slide 101: SHSSS for MX spots
separate:
mixed:
2.5%
0.9%
2.5%2-0.9%2 = 2.3%2

## Slide 102: Discerning Na+ from Mg++
f’’ (electrons)
DELFAN peak height (σ)
Mg
Ne
Na
F
O
N

## Slide 103: SHSSS: CCD vs PAD
3.5
3.0
2.5
2.0
1.5
1.0
0.5
Systematic component of Rshift (%)
distance between spots (mm)
0.1 1 10 100
Pilatus
Q315r
anomalous mates are
> 100 mm apart
different
modules

## Slide 104: HPAD calibration
θ
θ
λ
λ
thick
width
N
N
N
trim
trim
trim
N
N
N
trim
trim
trim
N
trim

## Slide 105: HPAD vs mmPAD
N
N
N
trim
trim
trim
N
N
N
trim
trim
trim
Q
Q
Q
gain
gain
gain
= k
= k
= k
θ
θ
λ
λ
thick
width

## Slide 106
Rdiff (%) between half-sets
5
4
3
2
1
0
sqrt(3.72/mult + 0.52)
sqrt(3.72/mult + 0.32)
Pilatus
1-pixel shift
Image multiplicity

## Slide 107
Rdiff (%) between half-sets
5
4
3
2
1
0
sqrt(4.12/mult + 3.22)
sqrt(4.12/mult + 0.22)
Image multiplicity
Pilatus
1-module shift
anomalous mates are always on different modules!

## Slide 108
sqrt(5.92/mult + 2.12)
sqrt(5.92/mult + 0.22)
Rdiff (%) between half-sets
5
4
3
2
1
0
Image multiplicity
Q315r
10-pixel shift