## Slide 1: PowerPoint File available:
http://bl831.als.lbl.gov/
~jamesh/powerpoint/
SLS_insitu_2017.pptx

## Slide 2: Acknowledgements
UCSF LBNL SLAC
ALS 8.3.1 creator: Tom Alber
UC Multicampus Research Programs and Initiatives (MRPI)
UCSF Program for Breakthrough Biomedical Research (PBBR)
one-time NIH-DOE Inter-agency agreement (IAA)
R01 GM124149 to Holton
GM117126 to Sauter (DIALS)
Integrated Diffraction Analysis Technologies (IDAT)
Plexxikon, Inc.
Synchrotron Radiation Structural Biology Resource (SLAC)
The Advanced Light Source is supported by the Director, Office of Science, Office of Basic Energy Sciences, Materials Sciences Division, of the US Department of Energy under contract No. DE-AC02-05CH11231 at Lawrence Berkeley National Laboratory.
Robert Stroud James Fraser
Nick Sauter John Tainer Aina Cohen Mike Soltis
http://bl831.als.lbl.gov/~jamesh/powerpoint/SLS_insitu_2017.pptx

## Slide 3
Acknowledgments
Fluidigm
  Andy May
  Brian Fowler
  Hany Nassef
  Geoff Facer
LBNL
  James Holton
  Ken Frankel
  George Meigs
  John Tainer
Funded by STTR Phase II grant (NIGMS) 2R42GM071326-02A1

## Slide 4: Three Laws of in-situ MX
Trays must be Thin
  plastic = water = protein != glass
    Normalize for absorption!
  walls <= xtal size
Trays must be Tight
  evaporation = nonisomorphism
  Plastic = pile of worms
Trays must be Cheap
  $1-20/tray
  New tray = new experiment

## Slide 7: Advantages of in-situ diffraction
No evaporation
- dehydration
- pH change (acetate, NH3)
- oxidation in air
No mechanical disturbance
- mounting, fishing, cryo-cooling
Low thermal artifacts
- protein expansion, pH of tris

## Slide 9: opened drop: you have ~30 seconds!
equilibrated drop

## Slide 10: First diffraction from protein xtal
Bernal, J. & Crowfoot, D. (1934). "X-ray photographs of crystalline pepsin", Nature 133, 794-795.
1934

## Slide 11: Dehydration: 1934
100 μm
2 mm
= 1.5 μL
= 1.0 nL
10 μm
= 1.0 pL
and 2017?
1 μm
= 1.0 fL

## Slide 12
Dear James
The story of the two forms of lysozyme crystals goes back to about 1964 when it was found that the diffraction patterns from different crystals could be placed in one of two classes depending on their intensities. This discovery was a big set back at the time and I can remember a lecture title being changed from the 'The structure of lysozyme' to 'The structure of lysozyme two steps forward and one step back'. Thereafter the crystals were screened based on intensities of the (11,11,l) rows to distinguish them (e.g. 11,11,4 > 11,11,5 in one form and vice versa in another). Data were collected only for those that fulfilled the Type II criteria. (These reflections were easy to measure on the linear diffractometer because crystals were mounted to rotate about the diagonal axis). As I recall both Type I and Type II could be found in the same crystallisation batch . Although sometimes the external morphology allowed recognition this was not infallible.
The structure was based on Type II crystals. Later a graduate student Helen Handoll examined Type I. The work, which was in the early days and before refinement programmes, seemed to suggest that the differences lay in the arrangement of water or chloride molecules (Lysozyme was crystallised from NaCl). But the work was never written up. Keith Wilson at one stage was following this up as lysozyme was being used to test data collection strategies but I do not know the outcome.
An account of this is given in International Table Volume F (Rossmann and Arnold edited 2001) p760.
Tony North was much involved in sorting this out and if you wanted more info he would be the person to contact. I hope this is helpful. Do let me know if you need more.
Best wishes
Louise
Non-isomorphism in lysozyme

## Slide 13
Johnson’s ratio
Structure factor (e-)
Non-isomorphism in lysozyme

## Slide 14
Johnson’s ratio
Structure factor (e-)
Non-isomorphism in lysozyme

## Slide 16: Non-isomorphism in lysozyme
RH: 84.2%
vs 71.9%
Riso = 44.5%
3aw6
3aw7
Δcell = 0.7 %
RMSD = 1.70 Å
RMSD = 0.18 Å

## Slide 17
Magdoff & Crick (1955) “Ribonuclease II. Accuracy of measurement and shrinkage”, Acta Cryst. 8, 461-8.
Crick & Magdoff (1956) “The theory of the method of isomorphous replacement for protein crystals”, Acta Cryst. 9, 901-8.

## Slide 18
0.5% cell change → 15%
Crick & Magdoff (1956) “theory method isomorphous replacement”, Acta Cryst. 9, 901-8.
Same structure, different cell

## Slide 19
Same structure, different cell
0.5% cell change → 15%
Crick & Magdoff (1956) “theory method isomorphous replacement”, Acta Cryst. 9, 901-8.

## Slide 20
Same cell, rotated protein
0.5° rotation → 16%
Crick & Magdoff (1956) “theory method isomorphous replacement”, Acta Cryst. 9, 901-8.

## Slide 21
Same cell, rotated protein
0.5° rotation → 16%
Crick & Magdoff (1956) “theory method isomorphous replacement”, Acta Cryst. 9, 901-8.

## Slide 22
Uniform stretch
5% change → 0%

## Slide 23
5% change → 0%
Uniform stretch

## Slide 24: Elastic deformation = non-isomorphism

## Slide 25: Elastic deformation = non-isomorphism

## Slide 26: Plastic deformation = poor diffraction

## Slide 27: Plastic deformation = poor diffraction

## Slide 28: Xtal5_t1 graph
Spot intensity

## Slide 29
Spot intensity

## Slide 30: How much H2 are we making?
Meents et al. (2010) claimed:
"400 molecules of H2 for every 12.5 keV photon“
= 3.3x10-7 mol/J
200 kGy to 100 um3 = 0.24 J
= 1.2x10-8 mol H2
100 um3 of water = 6.7x10-8 mol
→ 18% of water has reacted!

## Slide 31: Dose-rate dependence of damage
dose rate (kGy/s)
maximum useful dose (MGy)

## Slide 32
Where:
IDL - average damage-limited intensity (photons/hkl) at a given resolution
105 - converting R from μm to m, re from m to Å, ρ from g/cm3 to kg/m3 and MGy to Gy
re - classical electron radius (2.818 x 10-15 m/electron)
h - Planck’s constant (6.626 x 10-34 J∙s)
c - speed of light (299792458 m/s)
fdecayed - fractional progress toward completely faded spots at end of data set
ρ - density of crystal (~1.2 g/cm3)
R - radius of the spherical crystal (μm)
λ - X-ray wavelength (Å)
fNH - the Nave & Hill (2005) dose capture fraction (1 for large crystals)
nASU - number of proteins in the asymmetric unit
Mr - molecular weight of the protein (Daltons or g/mol)
VM - Matthews’s coefficient (~2.4 Å3/Dalton)
H - Howells’s criterion (10 MGy/Å)
θ - Bragg angle
a2 - number-averaged squared structure factor per protein atom (electron2)
Ma - number-averaged atomic weight of a protein atom (~7.1 Daltons)
B - average (Wilson) temperature factor (Å2)
μ - attenuation coefficient of sphere material (m-1)
μen - mass energy-absorption coefficient of sphere material (m-1)
Self-calibrated damage limit
Holton & Frankel (2010) Acta D 66 393-408.
No flux
No symmetry
No exposure time

## Slide 33
The number of photons scattered
before crystal is dead
is independent
of flux & time
1 um3 = 105 photons (roomT)
1 um3 = 106 photons (synch)
1 um3 = 108 photons (XFEL)
Henderson, 1990; Gonzalez & Nave, 1994; Glaeser et al., 2000; Sliz et al., 2003; Leiros et al., 2006; Owen et al., 2006; Garman & McSweeney, 2006; Garman & Nave, 2009; Holton, 2009
(synch)

## Slide 35
Holton & Frankel (2010) Acta D 66 393-408.

## Slide 36: Multi-crystal strategies
Kendrew et al. (1960) "Structure of Myoglobin” Nature 185, 422-427.

## Slide 37: Cell clustering: 100 datasets
Reality:
Perfect isomorphism
(simulated data)

## Slide 38: Micro-focus challenge
http://bl831.als.lbl.gov/~jamesh/challenge/microfocus/

## Slide 39: Dose slicing
crystal’s useful life
N
photons
N
photons
N
photons
unacceptable
completeness
unacceptable
read noise

## Slide 40: Dose slicing: is that a spot?

## Slide 41: Dose slicing: is that a spot?

## Slide 42: Dose slicing: is that a spot?

## Slide 43: Dose slicing: is that a spot?

## Slide 44: Dose slicing: is that a spot?

## Slide 45: Dose slicing: is that a spot?

## Slide 46: What if?
You only have a few small crystals…
Should you:
  Collect 360° from each?
  Collect 10° from each at 36x exposure?
  Glue 36 xtals together, then collect 360° ?
  Glue, and do 12960° faint exposures?

## Slide 47
Chart: 
Resolution (Å)
CC vs right answer
“true” resolution limit

## Slide 48
Dose slice collection scenario |  |  | 
fine | coarse | glue | Glue & fine
3.17700 | 3.14057 | 3.17604 | 3.12086
“true” resolution vs strategy
301,640,334 photons/xtal

## Slide 49: Trick Question: What is the space group?
a = 63 b = 63 c = 63
α = 90 β = 90 γ = 90

## Slide 50: Pitfalls: sub-space groups
P622
What else?
P6122 P6522 P6222 P6422 P6322
P6 P61 P65 P62 P64 P63
P312 P321 P3112 P3121 P3212 P3221
P3 P31 P32
P2 P21 C2 C222 C2221
P1

## Slide 51: Multi-crystal merge preparation
Take each pair of datasets
“othercell” to find all possible re-indexings
Apply all re-indexing operators in turn
Evaluate best agreement (R, CC)
Cluster by acceptable distance (20% ?)
Re-index clusters to common lattice
Apply xscale_isocluster

## Slide 52: And the dominant source of error is…
Resolution
  Disorder (including Radiation Damage)
  Background
Phases
  Non-isomorphism (including Radiation Damage)
  Vibration (beam, xtal, spindle, etc.)
  Detector calibration

## Slide 59
~4Ǻ
from Lysozyme?

## Slide 61: X-ray vs tray

## Slide 63: attenuation factor
Bouguer, P. (1729). Essai d'optique sur la gradation de la lumière.
Lambert, J. H. (1760). Photometria: sive De mensura et gradibus luminis, colorum et umbrae. E. Klett.
Beer, A. (1852)."Bestimmung der Absorption des rothen Lichts in farbigen Flüssigkeiten", Ann. Phys. Chem 86, 78-90.
A = = exp[-μxtal(txi+ txo)
-μsolvent(tsi + tso)]
IT
Ibeam
μxtal
txi
txo
tsi
tso
txi
txo
tsi
tso
txi
txo
tsi
tso
μsolvent

## Slide 64: X-ray transmission of tray materials
Depth into material (mm)
transmission
at the selenium edge

## Slide 65: X-ray transmission of tray materials
Depth into material (mm)
transmission
protein
crystal
at the selenium edge

## Slide 66: X-ray transmission of tray materials
Depth into material (mm)
transmission
protein
crystal
water
at the selenium edge

## Slide 67: X-ray transmission of tray materials
Depth into material (mm)
transmission
PDMS
protein
crystal
water
at the selenium edge

## Slide 68: X-ray transmission of tray materials
Depth into material (mm)
transmission
glass
PDMS
protein
crystal
water
at the selenium edge

## Slide 69: X-ray transmission of tray materials
Depth into material (mm)
transmission
glass
PDMS
protein
crystal
water
air
at the selenium edge

## Slide 70: X-ray transmission of tray materials
Depth into material (μm)
transmission
glass
PDMS
protein
crystal
water
air

## Slide 71: X-ray transmission of tray materials
Depth into material (μm)
transmission
glass
PDMS
protein
crystal
water
nylon
air

## Slide 72: X-ray transmission of tray materials
Depth into material (μm)
transmission
glass
PDMS
protein
crystal
water
nylon
plastic
air

## Slide 73: X-ray transmission of tray materials
Depth into material (μm)
transmission
glass
PDMS
protein
crystal
water
nylon
plastic
beryllium
air

## Slide 74: X-ray transmission of tray materials
Depth into material (μm)
transmission
glass
PDMS
protein
crystal
water
nylon
plastic
beryllium
air
diamond

## Slide 75: X-ray transmission of tray materials
Depth into material (μm)
transmission
glass
protein
crystal
water
air

## Slide 76: X-ray scattering of tray materials
Differential cross section
(scattered photons/solid angle/incident photon/meter thickness)
 7.5 3.8 2.5 1.9 1.5
d-spacing (Ǻ)
water

## Slide 77: X-ray scattering of tray materials
Differential cross section
(scattered photons/solid angle/incident photon/meter thickness)
 7.5 3.8 2.5 1.9 1.5
d-spacing (Ǻ)
water
25% glycerol

## Slide 78: X-ray scattering of tray materials
Differential cross section
(scattered photons/solid angle/incident photon/meter thickness)
 7.5 3.8 2.5 1.9 1.5
d-spacing (Ǻ)
water
25% glycerol
Paratone-N oil

## Slide 79: X-ray scattering of tray materials
Differential cross section
(scattered photons/solid angle/incident photon/meter thickness)
 7.5 3.8 2.5 1.9 1.5
d-spacing (Ǻ)
water
25% glycerol
Paratone-N oil

## Slide 80: X-ray scattering of tray materials
Differential cross section
(scattered photons/solid angle/incident photon/meter thickness)
 7.5 3.8 2.5 1.9 1.5
d-spacing (Ǻ)
water
25% glycerol
paratone-N oil

## Slide 81: X-ray scattering of tray materials
Differential cross section
(scattered photons/solid angle/incident photon/meter thickness)
 7.5 3.8 2.5 1.9 1.5
d-spacing (Ǻ)
water
25% glycerol
paratone-N oil
PDMS

## Slide 82: X-ray scattering of tray materials
Differential cross section
(scattered photons/solid angle/incident photon/meter thickness)
 7.5 3.8 2.5 1.9 1.5
d-spacing (Ǻ)
water
25% glycerol
paratone-N oil
PDMS
plastic1

## Slide 83: X-ray scattering of tray materials
Differential cross section
(scattered photons/solid angle/incident photon/meter thickness)
 7.5 3.8 2.5 1.9 1.5
d-spacing (Ǻ)
water
25% glycerol
paratone-N oil
PDMS
plastic1
plastic2

## Slide 84: X-ray scattering of tray materials
Differential cross section
(scattered photons/solid angle/incident photon/meter thickness)
 7.5 3.8 2.5 1.9 1.5
d-spacing (Ǻ)
water
25% glycerol
paratone-N oil
PDMS
plastic1
plastic2
plastic3

## Slide 85: X-ray scattering of tray materials
Differential cross section
(scattered photons/solid angle/incident photon/meter thickness)
 7.5 3.8 2.5 1.9 1.5
d-spacing (Ǻ)
water
25% glycerol
paratone-N oil
PDMS
plastic1
plastic2
plastic3
paper

## Slide 86: X-ray scattering of tray materials
Differential cross section
(scattered photons/solid angle/incident photon/meter thickness)
 7.5 3.8 2.5 1.9 1.5
d-spacing (Ǻ)
water
25% glycerol
paratone-N oil
PDMS
plastic1
plastic2
plastic3
paper
glass

## Slide 87: X-ray scattering of tray materials
Differential cross section
(scattered photons/solid angle/incident photon/meter thickness)
 7.5 3.8 2.5 1.9 1.5
d-spacing (Ǻ)
water
25% glycerol
paratone-N oil
PDMS
plastic1
plastic2
plastic3
paper

## Slide 88: X-ray scattering of tray materials
Differential cross section
(scattered photons/solid angle/incident photon/meter thickness)
 7.5 3.8 2.5 1.9 1.5
d-spacing (Ǻ)
water
25% glycerol
paratone-N oil
PDMS
plastic1
plastic2
plastic3
paper
Plastic … n

## Slide 89: X-ray scattering of tray materials
Differential cross section
(scattered photons/solid angle/incident photon/meter thickness)
 7.5 3.8 2.5 1.9 1.5
d-spacing (Ǻ)
paratone-N oil
PDMS
Plastic … n
water

## Slide 90: Background scattering
Ibg - scattered photons/steradian
Ibeam - incident (photons/s/m2 )
t - exposure time (s)
re - classical electron radius (2.818x10-15 m)
NA - Avogadro number (6.02x1023)
ρ - density of material (g/m3)
Mr - molecular weight (g/mol)
V - volume of material (in m3)
P - polarization factor
(1+cos2(2θ) -Pfac∙cos(2Φ)sin2(2θ))/2
A - absorption factor
exp(-μxtal∙lpath)
f(s) - molecular structure amplitude (electrons)
s - scattering length (sin(θ)/λ)
P A | f(s) |2
Ibg = Ibeam t re2
NA ρV
Mr
R. W. James (1947)

## Slide 91: Background scattering
Resolution (Ǻ)
Electron equivalents
The form-factor of the cryostream
20 10 5 4 3 2.5 2 1.8 1.6 1.4 1.2 1
0 2 4 6 8 10 12 14 16
measured
theoretical

## Slide 92: Background scattering
Ibg - scattered photons/steradian
Ibeam - incident (photons/s/m2 )
t - exposure time (s)
re - classical electron radius (2.818x10-15 m)
NA - Avogadro number (6.02x1023)
ρ - density of material (g/m3)
Mr - molecular weight (g/mol)
V - volume of material (in m3)
P - polarization factor
(1+cos2(2θ) -Pfac∙cos(2Φ)sin2(2θ))/2
A - absorption factor
exp(-μxtal∙lpath)
f(s) - molecular structure amplitude (electrons)
s - scattering length (sin(θ)/λ)
P A | f(s) |2
Ibg = Ibeam t re2
NA ρV
Mr
R. W. James (1947)

## Slide 93: Diffuse scattering
Ids - scattered photons/steradian
Ibeam - incident (photons/s/m2 )
t - exposure time (s)
re - classical electron radius (2.818x10-15 m)
Vxtal - volume of crystal (in m3)
VASU - asymmetric unit (in m3)
P - polarization factor
A - absorption factor
a - particular atom in the ASU
fa(s) - atomic structure amplitude (electrons)
s - scattering length (sin(θ)/λ)
Ba - atomic B factor
|fa(s)|2 (1-exp(-2Ba∙s2))
Ids = Ibeam t re2
Vxtal
VASU
P A
Σ
a
R. W. James (1947)

## Slide 94: Diffuse scattering
Ids - scattered photons/steradian
Ibeam - incident (photons/s/m2 )
t - exposure time (s)
re - classical electron radius (2.818x10-15 m)
Vxtal - volume of crystal (in m3)
VASU - asymmetric unit (in m3)
P - polarization factor
A - absorption factor
a - particular atom in the ASU
fa(s) - atomic structure amplitude (electrons)
s - scattering length (sin(θ)/λ)
Ba - atomic B factor
|fa(s)|2 (1-exp(-2Ba∙s2))
Ids = Ibeam t re2
Vxtal
VASU
P A
Σ
a
R. W. James (1947)

## Slide 95: scattering/absorptoin “rules”:
1 μm crystal ≈ 1 μm water
≈ 1 μm plastic
≈ 0.1 μm glass
≈ 1000 μm air
Optimum: support thickness = crystal thickness

## Slide 96: What we really need
GRAPHENE
Thin
Tight
Cheap

## Slide 97: Crowfoot Cell
Bernal & Crowfoot, Nature (London) 133, 794 (1934).

## Slide 98: Water permeability of glass
Friedman & Smith (1960), American Antiquity, 476-493.
x2 = kt
x = thickness t = time
k = um/1000 year

## Slide 99: Water permeability of glass
Liritzis (2006), Archaeometry. 48, 533-547.

## Slide 100: Water permeability of glass
Thin
~ Tight
? Cheap ?

## Slide 101
Soliman, A. et al. (2011). Acta Cryst. D 67, 646-656.
250 μm
NatX-ray
CrystalQuick™ X
~50 μm tape
MiTeGen
MicroRT™
25 μm wall
25 μm wall
http://www.mitegen.com/products/micrort/micrort.shtml
Fluidigm
Topaz® DC
http://www.fluidigm.com/home/fluidigm/docs/Datasheet_1.96DC.pdf
200 μm
PDMS
PDMS
http://www.mitegen.com/mic_catalog.php?c=insituplates
100 μm
MiTeGen
in-situ-1
~50 μm tape
Ng, J. et al. (2008). Acta Cryst. D 64, 189-197.
1000 μm
Microlytic
Crystal Former™
in-situ diffraction systems
Notes: Next slide please

## Slide 102
Fluidigm
Topaz® DC
http://www.fluidigm.com/home/fluidigm/docs/Datasheet_1.96DC.pdf
200 μm
Soliman, A. et al. (2011). Acta Cryst. D 67, 646-656.
250 μm
NatX-ray
CrystalQuick™ X
~50 μm tape
MiTeGen
MicroRT™
25 μm wall
25 μm wall
http://www.mitegen.com/products/micrort/micrort.shtml
PDMS
PDMS
http://www.mitegen.com/mic_catalog.php?c=insituplates
100 μm
MiTeGen
in-situ-1
~50 μm tape
in-situ diffraction systems
Márquez, J. & Cipriani, F. (2014) Structural Genomics (Chen, Y. W., ed) pp. 197-203.
Crystal DirectTM
~50 μm tape
25 μm
Notes: Next slide please

## Slide 103: Zero-parallax microscope
pinhole
mirror
Styrofoam™ backlight
backstop
microscope
Notes: Next slide please

## Slide 104: Zero-parallax microscope
pinhole
mirror
Styrofoam™ backlight
backstop
microscope
Notes: Inspired by the ESRF microcrystal camera, the superbend endstations are all equipped with zero-paralax optics. The user sees their sample from the point of view of the x-ray beam. Sample alignment is not only more stable but more intuitive than an off-axis camera. Placing the crystal into the microscope crosshairs gaurantees that it will be hit by the beam.

## Slide 105: Zero-parallax microscope
pinhole
mirror
microscope
Notes: Inspired by the ESRF microcrystal camera, the superbend endstations are all equipped with zero-paralax optics. The user sees their sample from the point of view of the x-ray beam. Sample alignment is not only more stable but more intuitive than an off-axis camera. Placing the crystal into the microscope crosshairs gaurantees that it will be hit by the beam.

## Slide 106: Zero-parallax microscope
pinhole
mirror
microscope
Notes: Inspired by the ESRF microcrystal camera, the superbend endstations are all equipped with zero-paralax optics. The user sees their sample from the point of view of the x-ray beam. Sample alignment is not only more stable but more intuitive than an off-axis camera. Placing the crystal into the microscope crosshairs gaurantees that it will be hit by the beam.

## Slide 107: Zero-parallax microscope
pinhole
mirror
microscope
Notes: Inspired by the ESRF microcrystal camera, the superbend endstations are all equipped with zero-paralax optics. The user sees their sample from the point of view of the x-ray beam. Sample alignment is not only more stable but more intuitive than an off-axis camera. Placing the crystal into the microscope crosshairs gaurantees that it will be hit by the beam.

## Slide 108: Sample parallax
Image plane
mirror
x-rays
Microscope
light cone
Notes: Inspired by the ESRF microcrystal camera, the superbend endstations are all equipped with zero-paralax optics. The user sees their sample from the point of view of the x-ray beam. Sample alignment is not only more stable but more intuitive than an off-axis camera. Placing the crystal into the microscope crosshairs gaurantees that it will be hit by the beam.

## Slide 109: Sample parallax
Notes: Inspired by the ESRF microcrystal camera, the superbend endstations are all equipped with zero-paralax optics. The user sees their sample from the point of view of the x-ray beam. Sample alignment is not only more stable but more intuitive than an off-axis camera. Placing the crystal into the microscope crosshairs gaurantees that it will be hit by the beam.

## Slide 110: Sample parallax
n>1
Notes: Inspired by the ESRF microcrystal camera, the superbend endstations are all equipped with zero-paralax optics. The user sees their sample from the point of view of the x-ray beam. Sample alignment is not only more stable but more intuitive than an off-axis camera. Placing the crystal into the microscope crosshairs gaurantees that it will be hit by the beam.

## Slide 111: Sample parallax
n>1
Notes: Inspired by the ESRF microcrystal camera, the superbend endstations are all equipped with zero-paralax optics. The user sees their sample from the point of view of the x-ray beam. Sample alignment is not only more stable but more intuitive than an off-axis camera. Placing the crystal into the microscope crosshairs gaurantees that it will be hit by the beam.

## Slide 112: Sample parallax
n>1
Notes: Inspired by the ESRF microcrystal camera, the superbend endstations are all equipped with zero-paralax optics. The user sees their sample from the point of view of the x-ray beam. Sample alignment is not only more stable but more intuitive than an off-axis camera. Placing the crystal into the microscope crosshairs gaurantees that it will be hit by the beam.

## Slide 113: Sample parallax
n>1
Notes: Inspired by the ESRF microcrystal camera, the superbend endstations are all equipped with zero-paralax optics. The user sees their sample from the point of view of the x-ray beam. Sample alignment is not only more stable but more intuitive than an off-axis camera. Placing the crystal into the microscope crosshairs gaurantees that it will be hit by the beam.

## Slide 118
8000 Gy/s

## Slide 119: pseudo-Laue from Lysozyme

## Slide 120: Laue spot from lysozyme
10s exposure
after 30s
as in: Snell, et al. (1995), Acta Cryst D. 51, 1099-102.

## Slide 121: Rastering math
Damage & diffraction = flux / area
Gy/s = ph/s /2000 / um2
Example: flux= 1012 ph/s, 1 mm2 beam = 500 Gy/s
1 mm2 beam = 500 Gy/s → 30 MGy = 17 hours
100 µm beam = 50 kGy/s * 10x10 shots = 17 hours
10 µm beam = 5 MGy/s * 100x100 shots = 17 hours
1% sample life → 300 kGy / 500 Gy/s = 10 min
room temp → 200 kGy / 500 Gy/s = 7 min
Holton (2009)."beginner's guide rad dam", J. Synch. Rad. 16, 133-142.

## Slide 122: Multilayers: better for rastering
Si(111)
Existing mono
Si(111)
ML
ML
0.014 % (1x1012 ph/s
= 400 s/mm2)

## Slide 123
Si(111)
Existing mono
ML
ML
1 % (7x1013 ph/s
= 2 s/mm2)
Si(111)
Multilayers: better for rastering

## Slide 124: Wide bandpass reveals mosaicity from a still
Snell EH, Weisgerber S, Helliwell JR, Holzer K & Schroer K (1995)."Improvements in lysozyme protein crystal perfection through microgravity growth", Acta crystallographica. Section D, Biological crystallography 51, 1099-1102.
low mosaic
high mosaic
Δradial = 2ηD/cos22θ

## Slide 125: Si(111) vs multilayers
0.014%
Si(111)
1%
multilayer
4 Å
1.9 Å
resolution
spectral
dispersion

## Slide 126: Wide-bandpass MX?
Mono | Si(111) | multilayer | Difference (%) | good
Energy (eV) | 11111 | 12027.5 | +8.2 | Same
Scale | 1.0000 | 1.0964 | +9.6 | Yes
Rmerge (%) | 7.5 | 8.4 | +12 | No
Ranom (%) | 5.6 | 5.5 | -1.8 | Same
I/sd | 13.2 | 12.3 | -6.8 | Same
I/sd (1.5Å) | 2.3 | 1.8 | -21.7 | No
Ano Patt | 5.96 | 4.75 | -20.3 | Same
Ano Phased | 41 | 39.33 | -4.1 | Same
FOM | 0.238 | 0.192 | -19.3 | Same
CC | 0.555 | 0.647 | -6.1 | Yes
Rcryst (%) | 26.70 | 26.48 | -0.2 | Same
Ligand peak | 8.213 | 8.250 | +0.4 | Same

## Slide 127
ML good for raster
What about collection?

## Slide 128: High energy compresses pattern
2 Å data
2 Å data
Notes: Next slide please

## Slide 129: Example Data
70 μm lysozyme xtal
50 μm beam
37 Gy/s (0.775 Å)
30s exposures at 20C
90° of data, 97% complete
I/ = 1.5 at 1.9 Å
Rmerge 18% (overall) 5% (low)
ΔB same as 20 min at 100K
B factor
-2
-4
-6
scale vs batch
2
1
0

## Slide 130
Room Temperature Sample Mounting
at LCLS-MFX
10 Room Temperature
Sample Holder Locations
Crystallization Plate
Uni-Puck
LCLS-MFX June 2016
(Problem - samples could dry out)
=
sample base holder
solution vial
SAM Compatible
Crystallization Plate
SSRL

## Slide 131
Humidity Controlled Crystallization,
Storage and Delivery Plate
Ring magnet holds sample base inside
Sample delivery device in hydrated environment
Automated sample exchange with SAM
Compatible with Qiagen screw-caps
Capacity for 5 trays/50 sample bases
(Humidity controlled enclosure for shelf in progress)
SSRL

## Slide 132: Multi-Crystal Mounting Grid
Cohen, A. E. et al. (2014). Proc. Natl Acad. Sci. USA, 111, 17122–17127
Baxter E. L. et al. Acta Cryst. (2016). D72, 2–11.
Single thin window
After X-ray Exposure at LCLS
Raster Data Collection in Grids
Comparison of grid to copper-magnetic pin
Crystallization Drops on Window
SSRL

## Slide 133: Automation for Filling Grids
Adaptor in standard microplate form factor
Compatible with a variety of fluid handing robots
  Labcyte Echo (hanging drop)
  Art Robbins Gryphon and TTP Labtech Mosquito (LCP)
Gryphon
Mosquito
Baxter E. L. et al. Acta Cryst. (2016). D72, 2–11.
SSRL

## Slide 134: Three Laws of in-situ MX
Must be Thin
  plastic = water = protein != glass
    Normalize for absorption!
  walls <= xtal size
Must be Tight
  evaporation = nonisomorphism
  Plastic = pile of worms
Must be Cheap
  $1-20/tray
  New tray = new experiment

## Slide 135: In-situ diffraction
promises
Better resolution
Isomorphism
Control: “no hands!”
Dose rate advantage?
Mosaicity from still
Datasets possible
pitfalls
Radiation damage
Trays dry out
thick = background
Thin = evaporation
Graphene = $$$$
2D array = bulky
Microscope parallax
http://bl831.als.lbl.gov/~jamesh/powerpoint/SLS_insitu_2017.pptx

## Slide 136: In-situ diffraction
promises
Better resolution
Isomorphism
Control: “no hands!”
Dose rate advantage?
Mosaicity from still
Datasets possible
pitfalls
Radiation damage
Trays dry out
thick = background
Thin = evaporation
Graphene = $$$$
2D array = bulky
Microscope parallax
http://bl831.als.lbl.gov/~jamesh/powerpoint/SLS_insitu_2017.pptx

## Slide 137: … but does it work for membrane proteins?
Efremov, R. G. & Sazanov, L. A. (2011)."Structure of the membrane domain of respiratory complex I", Nature 476, 414-420.
“To improve diffraction properties, crystals were slowly cooled to 4 C, placed into micro-dialysis buttons, and mother liquor was exchanged by dialysis for 0.1 M sodium acetate (pH 4.8), 0.15 M sodium tartrate (pH 5.0), 30% PEG4000, 0.1% sodium cholate and 12.5% glycerol over a period of 5 days.”
CCP4BB:
“We used it for a large membrane protein complex and diffraction improved from ~7 to 2.7 A. The crystal is fished out and put into mother liquor solution in the button, sealed with dialysis membrane and the button is then placed into about 5 mls of mother liquor with slightly higher PEG concentration. Then you just exchange outside buffer every day or so for solutions containing higher concentrations of PEG. We went from ~9 to 30 % PEG4000 in about a week. You can easily observe crystal under microscope and if it cracks - you went too far/too quickly with PEG and need to use a bit less next time. Also, this method allows you to control all other components of the dehydrating solution - we needed to decrease salt concentration at the same time as increasing PEG. You can also introduce/increase cryo-protectant concentration at the same time. With these crystals, otherwise excellent dehydration machines already mentioned did not work, possibly because the process had to be really slow.”

## Slide 138
Experiments at Controlled Humidity
Test Case: Bovine Liver Catalase Crystals
Dehydration Protocol
starting RH=96%, final RH=90%, 1 % RH steps,
300 seconds equilibration time between steps
96 % RH
90 % RH
1.7 Å
2.7 Å
a=87.8 Å, b=140.8 Å, c=233.1 Å, P212121
a=84.6 Å, b=142.0 Å, c=233.8 Å, P212121
Arinax Humidity Control Device acquired for MFX and SSRL
Data collection at room temperature
Controlled dehydration to improve diffraction resolution
SSRL

## Slide 139: What you know when:
Autoindexing: Lattice (14)
Scaling: Point Group (43)
Refinement: Space Group (84)

## Slide 140: plunge cooling
foam insulation
warm, moist air
cold N2
liquid N2
ice

## Slide 141: ice = bad

## Slide 142: plunge cooling
cold N2
foam insulation
warm, moist air
liquid N2
ice

## Slide 143: Warkentin method
foam insulation
liquid N2
Warkentin (2006) J. Appl. Crystallogr. 39, 805-811.

## Slide 144: Tris buffer vs temperature
Temperature (C)
pH

## Slide 145: Tris buffer under cryo
Temperature (K)
pH*
Douzou (1977) Cryobiochemistry. Academic Press.

## Slide 146: Dielectric matching
Travers & Douzou (1970) J. of Phys. Chem. 74, 2243-2244.
Petsko (1975) J. Mol. Biol. 96, 381-388.
Douzou (1977) Cryobiochemistry. Academic Press.

## Slide 147
processing procedure | resolution limit |  |  | 
 | fine | coarse | glue | glue&fine
XDS/XSCALE | 3.17700 | 3.14057 | 3.17604 | 3.12086
XDS/aimless | 3.16251 | 3.15922 | 3.16675 | 3.12162
XDS/noscale | 3.27289 | 3.15617 | 3.38929 | 3.12162
MOSFLM-defaults | x | x | x | x
MOSFLM-cheat | 3.19755 | 3.18166 | 3.20354 | 3.19755
HKL2000-defaults | x | x | x | x
HKL2000-expert | 3.21803 | 3.12646 | 3.11551 | 3.21803
HKL2000-cheat | 3.16119 | 3.10338 | 3.10338 | 3.16300
DIALS-defaults | x | x | 3.42125 | x
DIALS-expert | 3.10296 | 3.09547 | 3.07653 | 3.08465
d*Trek | ? | ? | 4.34076 | ?
EVAL15 | 3.14352 | 3.10383 | 3.09936 | 3.14105
EVAL15-sadabs | 3.13920 | 3.10797 | 3.10171 | 3.13700

## Slide 148: Minimum exposure time:
~1 photon/pixel
  PAD w XDS, DIALS
~10 photon/pixel
PAD w HKL2k, MOSFLM
~30 photon/pixel
CCD detector

## Slide 149: Plate goniometers
Notes: The non-rotating “plate goniometer” (left) was developed last year. A new incarnation for taking rotation-camera data is now being developed (right). Either of these devices will accept most any crystallization tray that can be rotated 90 degrees to gravity and probe arbitrary points in each crystallization drop with a 50 um x-ray beam. Installing a plate goniometer takes ~30 minutes and the location ~300 mm downbeam from the “normal” protein crystallography sample mounting point makes this technology transferable to most any beamline.
Current status is that a 95% complete data set has been collected from a 50 um lysozyme crystal at room temperature, and trials are now underway to collect data sets from “real” crystals.