## Slide 2
8.3.1
C1
Failed
Jan 4 2023
Silicon
cracked

## Slide 3: Current best “spare”

## Slide 4: “French Glue” is no longer available

## Slide 5
COOLED ON ONE SIDE
COOLED ON CHANNELS

## Slide 6
HEAT LOAD ~ 40 W
PEAK POWER DENSITY
~0.1 W /mm^2
(ESTIMATED BY
DANIELE, JAN 2022)

## Slide 7
COOLED ON ONE SIDE
COOLED ON CHANNELS
STEADY-STATE TEMPERATURE DISTRIBUTION

## Slide 8
HEAT BUMPS
NOTE: RIPPLES ARE DUE TO VARIATION IN THERMAL CONDUCTIVITY DUE TO CHANNELS (VOIDS IN MATERIAL)
NOT DUE TO COOLING OR PRESSURE

## Slide 10
Worst-case Scenario: 150 W/m^2K

## Slide 11
Epotecny E707 glue is no longer available
Masterbond E33 failed 2x in 3 months
Solder not reproducible

## Slide 13
70 lb

## Slide 16: MasterBond E33

## Slide 20
Epotecny E707 glue is no longer available
Masterbond E33 failed 2x in 3 months
Solder not reproducible

## Slide 22
65.3 mm
52.7 mm
2.3 mm
2.3 mm
2.4 mm

## Slide 26
15
7.5
45
60
70
50
75

## Slide 28: Transient Heat Test
Chart: 
| 0 | 1 | 2 | 3 | 4 | 5 | 6
Txtal | 18.1 | 20.5 | 22.5 | 24 | 24.9 | 25.8 | 26.6
Tplenum | 18.1 | 18.4 | 19.4 | 20.6 | 21.6 | 22.5 | 23.3
Temperature (C)
Time (minutes)