## Slide 2
8.3.1
C1
Failed
Jan 4 2023
Silicon
cracked

## Slide 3: Current best “spare”

## Slide 4: “French Glue” is no longer available

## Slide 5
COOLED ON ONE SIDE
COOLED ON CHANNELS

## Slide 6
HEAT LOAD ~ 40 W
PEAK POWER DENSITY
~0.1 W /mm^2
(ESTIMATED BY
DANIELE, JAN 2022)

## Slide 7
COOLED ON ONE SIDE
COOLED ON CHANNELS
STEADY-STATE TEMPERATURE DISTRIBUTION

## Slide 8
HEAT BUMPS
NOTE: RIPPLES ARE DUE TO VARIATION IN THERMAL CONDUCTIVITY DUE TO CHANNELS (VOIDS IN MATERIAL)
NOT DUE TO COOLING OR PRESSURE

## Slide 10
Worst-case Scenario: 150 W/m^2K

## Slide 11
Epotecny E707 glue is no longer available
Masterbond E33 failed 2x in 3 months
Solder not reproducible

## Slide 13
70 lb

## Slide 15: MasterBond E33